## Slide 1
`
Quench Protection
Samuel Fraga
Electronics Group
On behalf of the Technology Unit
CNPEM’s Superconducting Wavelength Shifter (SWLS)
Technical Design Review
4–5th November 2024

## Slide 2
Design Premises
Protection System Overview
Hot-spot Temperature Analysis
Dump Resistor Specification
Quench Detection Requirements
Quench Detection Architecture
Prototypes, Tests and Results
Outline

## Slide 3
Adopt conservative requirements seeking robustness
Favor well-established solutions
Increase design complexity only as required by the project needs
Minimize the number of critical components (single points of failure)
Add redundancies whenever viable
Design Premises

## Slide 4
Active energy extraction with series resistor
3 current lead scheme allows treating extraction circuit as 2 independent RL circuits
2 redundant contactors per extraction circuit
Dump resistor bank specified to withstand partial failure
Grounding position to decrease maximum circuit voltage to ground
Quench detection on similar coils by voltage imbalance
Quench detection on current leads by overvoltage
Quench detection safety redundancy by overvoltage using higher threshold than imbalance detector
Overvoltage detection on current leads, current links and splices
Protection System Overview

## Slide 5
Current leads include ballast, so their resistance is negligible
Assume 2 decoupled RL circuits
Equivalent inductance obtained from stored energy at operating current
Worst-case scenario (highest inductance) is used in quench analysis
MIITS analysis script in Python calculates hot-spot temperature for a Copper and Nb-Ti composite wire considering:
  Cu and Nb-Ti specific heat as function of magnetic field density and temperature
  Cu resistivity as function of temperature, RRR and magnetic field density
  Cu and Nb-Ti density as function of temperature
  Detection time obtained from approximation of quench voltage evolution
Python script implementing a one-dimensional Wilson simulation was developed, but MIITS yielded more conservative results
Hot-spot Temperature Analysis
Low resistance
Highest inductance loop
First approximation for normal zone voltage in adiabatic case
Main coils: 2670 J
Side coils: 750 J

## Slide 6
Minimize dump resistance to minimize extraction voltage
Analysis uses worst-case scenario according to hot-spot temperature calculations for dump specification
Dump resistor matrix optimized to withstand the highest count of open-circuits to component count ratio
Minimum allowed dump resistance is 0.7 W
Voltage tap insulation is rated for 250 V (max extraction voltage)
Estimate of impact of resistor tolerance in design
Dump Resistor Specification
Parameter | Value
Maximum hot-spot temperature | 100 K
Maximum operating current | 300 A
Critical temperature | 9.2 K
Current-sharing temperature | 6.1 K
Operating temperature | 4.2 K
Copper area on wire | 2.76144e-7 m2
Nb-Ti area on wire | 2.99794e-7 m2
Copper RRR | 131
Validation time | 40 ms
Actuation time | 12 ms
Detection voltage | 100 mV
Maximum field on wire | 5.3 T
Equivalent constant inductance | 96 mH
Effect of 20 % tolerance in resistor units
Impact is manageable

## Slide 7
Detection threshold configurable to <= 100 mV
Validation time configurable to <= 40 ms
Any critical failure triggers the protection system
All triggers are active in the low voltage state (safety against power failure)
Voltage inputs rated to 1 kV for robustness (max voltage between taps is 500 V in the worst case)
Broken wire detection
Redundancy of voltage taps (2 per monitored point) for robustness against broken solder joints
Quench Detection Requirements

## Slide 8
Digital system based on FPGA
  Suits the large number of channels needed
  Eases the adjustment of detection parameters
  Supports data acquisition
Differential input channels, with ±20 V range, digitally isolated
18-bit resolution ADC gives LSB < 153 uV
20 kHz sampling rate, SPI data transfer
The chosen FPGA (Zynq 7000) is built in a SoC with an integrated ARM processor, used for data acquisition and setting of voltage threshold and validation time
Quench Detection Architecture
Modular hardware: separate boards for different circuit functional blocks;
Eurocard boards are assembled in a subrack and interconnected through a passive plane
Allows fast substitution of parts of the circuit, facilitating maintenance

## Slide 9
FPGA System Overview
User-Defined parameters (thresholds, validation time, channel disable) via MQTT
Channel-Separated configuration, all parallel.
Critical processing path for the interlock.
Higher Level (Petalinux) application for data pre-processing, compression, etc.
AXI-stream Compliant External Interface.
DMA implementation for efficient data transfer.
Modularized and Scalable design, all parametrized by Generics.

## Slide 10
FPGA Physical Layer
The system is composed of three macro-level subsystems:
ADC driver;
Detection critical path ​;
Acquisition path to processing system.

## Slide 11
Operation Flow
Reliable self-watching mechanisms for interlock.
Protection is triggered if communication with any ADC fails.
Failures in the acquisition pipeline are signaled by LEDs and logged for debugging.

## Slide 12
Prototypes, Tests and Results
Extraction circuit crate
Analog QDS prototype (at right) + DAQ (at left)
A prototype protection system was built for the test setup of the superconducting coils
The extraction circuit is very similar to the final application, differing only on the array of parallel and series resistors and the presence of a central GND
Detection by analog circuit, with 4 overvoltage channels
Voltage signals acquired by external DAQ (National Instruments NI-9239)
Protection method and extraction circuit were validated and due to limitations of the analog detection, we decided to move to a digital solution

## Slide 13
FPGA Validation Testbenches
Top level testbench
Timing analysis
Power analysis
Resources analysis
FPGA firmware was implemented and virtual testbenches validated the following functions:
  ADC driver
  Detection path
  Threshold and validation time configuration
  System watcher
Design reports indicate moderate resource utilization, low power consumption and no timing issues.

## Slide 14
FPGA Validation Testbenches
System Watcher Testbench
Threshold Comparison Testbench
ADC Driver Testbench
Validation Time Testbench

## Slide 15
cnpem.br
Thank you!