## Slide 1
11th BGC Collaboration Meeting
Summary of the 2024 Experimental Results at
Cockcroft Institute Presented by Hao Zhang Shakti Prasad Sethi Department of Physics, QUASAR Group, University of Liverpool

## Slide 2
11th BGC Collaboration Meeting
JEREMY modification test stand
  JEREMY modification with bellows, 4th skimmer and dump chamber
  Background pressure for nozzle distance at different bellow positions
HaloJet
Density scan with rectangular 2nd skimmer in comparison with circular 2nd skimmer
Summaries
Contents

## Slide 3
Dump chamber
Bellow
4th Skimmer
1st Skimmer
2nd Skimmer
3rd Skimmer
IP
JEREMY Test stand
Modified JEREMY Test stand with Bellow,
4th Skimmer and Dump chamber
JEREMY modification setup
11th BGC Collaboration Meeting
1st Skimmer
2nd Skimmer
3rd Skimmer
IP

## Slide 4
New large 1st skimmers (400um, 600um, 700um, 800um)
30um Flat nozzle used
 | Nozzle | 1st Skim | 2nd Skim | 3rd Skim
Size | 30um | 700um | 5.3mm | 0.1x30mm
Distance | --- | 4.5mm | 31.7mm | 183.3 mm
700um 1st skimmer
Skimmer mount
Alignment
JEREMY setup tested with new 1st skimmers planned
11th BGC Collaboration Meeting

## Slide 5
JEREMY modification with
Bellows, 4th Skimmer and Dump Chamber
 | 1st Skim | 2nd Skim | 3rd Skim | IP | 4th Skim | Bellow variation
Size | 700um | 5.3mm | 0.4x30mm | --- | 4x60mm | 
Distance | 4.5mm | 31.7mm | 302.03mm | 452.21mm | 624.89mm | 80x120mm
Bellow
4th skimmer
Dump chamber
11th BGC Collaboration Meeting

## Slide 6
Nozzle
1st Skimmer
2nd Skimmer
3rd Skimmer
IP
4th Skimmer
Bellows + Spacers
Spacers
Bellows + Mount plate
Bellows + Mount plate
Bellows + Mount plate
Varying bellows is easy, but only one set of bellows between 3rd, IP and 4th so flange changes needed.
Take a range of BG pressures at different chamber sizes to validate MOGA predictions
JEREMY modification New layout options
11th BGC Collaboration Meeting

## Slide 7
  Background pressure study with
  bellow positions at different nozzle skimmer distances
Nozzle – 1st skimmer distance variation (2-10mm)
JEREMY setup
Nozzle=0.03mm, 1st skimmer=0.6mm, 2nd skimmer=2mm, 3rd skimmer a=0.4,b=30
Modified JEREMY setup with Bellow distances 80mm, 100mm, 120mm between 2nd – 3rd skimmer Nozzle=0.03mm, 1st skimmer=0.6mm, 2nd skimmer=2mm, 3rd skimmer a=0.4,b=30,4th skimmer a=4,b=60
*Compare expected BG pressure per chamber, simulation vs measured
11th BGC Collaboration Meeting
Distances in mm |  |  | 
Nozzle-1st | 1st-2nd | 2nd-3rd | 3rd-IP
2-10 | 24.75 | 250.33-290.33 | 150.18
Distances in mm |  |  |  | 
Nozzle-1st | 1st-2nd | 2nd-3rd | 3rd-IP | IP-4th
2-10 | 24.75 | 250.33-290.33 | 150.18 | 172.68

## Slide 8
  Background pressure study with/without
  bellow positions at different nozzle distance
11th BGC Collaboration Meeting
JEREMY setup
Nozzle=0.03mm, 1st skimmer=0.6mm, 2nd skimmer=2mm, 3rd skimmer a=0.4,b=30
Distances in mm |  |  | 
Nozzle-1st | 1st-2nd | 2nd-3rd | 3rd-IP
2-10 | 24.75 | 250.33-290.33 | 150.18
Nozzle-skimmer distances | 1st skimmer | 2nd skimmer | 3rd skimmer | IP
2 (No gas) | 1.50E-08 | 7.30E-09 | 3.80E-09 | 4.29E-09
2 | 3.43E-03 | 1.10E-04 | 1.53E-05 | 1.92E-07
3 | 3.53E-03 | 1.00E-04 | 1.47E-05 | 1.91E-07
4 | 3.53E-03 | 9.87E-05 | 1.40E-05 | 1.85E-07
5 | 3.53E-03 | 9.77E-05 | 1.33E-05 | 1.79E-07
6 | 3.53E-03 | 9.63E-05 | 1.30E-05 | 1.72E-07
7 | 3.53E-03 | 9.60E-05 | 1.30E-05 | 1.66E-07
8 | 3.53E-03 | 9.57E-05 | 1.27E-05 | 1.61E-07
9 | 3.53E-03 | 9.57E-05 | 1.23E-05 | 1.57E-07
10 | 3.53E-03 | 9.50E-05 | 1.20E-05 | 1.52E-07

## Slide 9
11th BGC Collaboration Meeting
Modified JEREMY setup with Bellow distances 80mm, 100mm, 120mm between 2nd – 3rd skimmer
Nozzle=0.03mm
1st skimmer=0.6mm
2nd skimmer=2mm
3rd skimmer a=0.4,b=30
4th skimmer a=4,b=60
Nozzle-skimmer (mm) | Bellow distances (mm) | 1st | 2nd | 3rd | IP | Dump
2 | 80 | 3.30E-03 | 1.10E-04 | 1.40E-05 | 1.59E-07 | 1.20E-07
3 | 80 | 3.63E-03 | 1.10E-04 | 1.40E-05 | 1.67E-07 | 1.20E-07
4 | 80 | 3.63E-03 | 1.07E-04 | 1.33E-05 | 1.62E-07 | 1.20E-07
5 | 80 | 3.60E-03 | 1.00E-04 | 1.30E-05 | 1.57E-07 | 1.20E-07
6 | 80 | 3.63E-03 | 1.00E-04 | 1.30E-05 | 1.53E-07 | 1.17E-07
7 | 80 | 3.63E-03 | 1.00E-04 | 1.20E-05 | 1.48E-07 | 1.10E-07
8 | 80 | 3.57E-03 | 9.93E-05 | 1.20E-05 | 1.45E-07 | 1.10E-07
9 | 80 | 3.60E-03 | 9.93E-05 | 1.20E-05 | 1.41E-07 | 1.10E-07
10 | 80 | 3.57E-03 | 9.90E-05 | 1.20E-05 | 1.38E-07 | 1.10E-07
2 | 100 | 3.13E-03 | 1.03E-04 | 1.40E-05 | 1.52E-07 | 1.40E-07
3 | 100 | 3.47E-03 | 1.10E-04 | 1.40E-05 | 1.58E-07 | 1.50E-07
4 | 100 | 3.47E-03 | 9.87E-05 | 1.30E-05 | 1.50E-07 | 1.50E-07
5 | 100 | 3.47E-03 | 9.67E-05 | 1.27E-05 | 1.46E-07 | 1.50E-07
6 | 100 | 3.47E-03 | 9.47E-05 | 1.20E-05 | 1.41E-07 | 1.40E-07
7 | 100 | 3.47E-03 | 9.37E-05 | 1.20E-05 | 1.36E-07 | 1.40E-07
8 | 100 | 3.43E-03 | 9.30E-05 | 1.20E-05 | 1.33E-07 | 1.30E-07
9 | 100 | 3.43E-03 | 9.30E-05 | 1.13E-05 | 1.30E-07 | 1.20E-07
10 | 100 | 3.43E-03 | 9.23E-05 | 1.07E-05 | 1.27E-07 | 1.20E-07
2 | 120 | 3.03E-03 | 1.00E-04 | 1.37E-05 | 1.47E-07 | 1.30E-07
3 | 120 | 3.33E-03 | 1.03E-04 | 1.40E-05 | 1.50E-07 | 1.40E-07
4 | 120 | 3.33E-03 | 1.01E-04 | 1.30E-05 | 1.43E-07 | 1.40E-07
5 | 120 | 3.33E-03 | 9.47E-05 | 1.23E-05 | 1.38E-07 | 1.37E-07
6 | 120 | 3.33E-03 | 9.27E-05 | 1.20E-05 | 1.33E-07 | 1.30E-07
7 | 120 | 3.33E-03 | 9.17E-05 | 1.17E-05 | 1.28E-07 | 1.27E-07
8 | 120 | 3.33E-03 | 9.13E-05 | 1.13E-05 | 1.25E-07 | 1.20E-07
9 | 120 | 3.33E-03 | 9.07E-05 | 1.10E-05 | 1.23E-07 | 1.10E-07
10 | 120 | 3.33E-03 | 9.07E-05 | 1.10E-05 | 1.21E-07 | 1.10E-07
Nozzle skimmer distance 10mm |  |  | 
Nozzle-1st | 2nd | 3rd | IP
3.53E-03 | 9.50E-05 | 1.20E-05 | 1.52E-07

## Slide 10
HaloJet
11th BGC Collaboration Meeting
Fig. Principle of a gas-jet based beam profile monitoring
Fig. slotted 3rd skimmer
Fig. Number density distribution of the gas curtain,
(a) simulation and (b) experiment
(a)
(b)

## Slide 11
Density scan with Rectangular 2nd skimmer compared with Circular 2nd skimmer
11th BGC Collaboration Meeting
2D profile of circular 2nd skimmer
2D profile of rectangular 2nd skimmer
Fig. Density scan of circular and rectangular skimmer
Nozzle-1st skimmer distance | 1.5mm (initial distance)
Nozzle | 0.03
1st skimmer | 0.4
2nd skimmer | 2
3rd Skim | a=0.4, b=30
bellows 2-3 | 80mm
Nozzle-1st skimmer distance | 1.5mm (initial distance)
Nozzle | 0.03
1st skimmer | 0.4
2nd skimmer | a=0.5, b=1.8
3rd Skim | a=0.4, b=30
bellows 2-3 | 80mm

## Slide 12
Chamber pressure at nozzle skimmer distances
11th BGC Collaboration Meeting
  Fig. (a)
  Fig. (b)
  Fig. (c)
  Fig. (d)
  Fig. (e)
  Fig. (a)-(e), represents the chambers pressure of circular and rectangular 2nd skimmers of the JEREMY setup
  a=2
  a=0.5, b=1.8

## Slide 13
Summaries
11th BGC Collaboration Meeting
Gas jet calibrated with circular and rectangular 2nd skimmer.
Background pressure measured at different bellows distance for 0.4mm and 0.7mm 1st skimmer.
HaloJet experiment performed using slotted 3rd skimmer in the Version 2 setup, and the work also presented at IBIC 2024.
Pulse jet (Gated injection) experiment for density measurement is planned.

## Slide 14
THANK YOU
11th BGC Collaboration Meeting
Gas Jet group
Carsten Welsch
Hao Zhang
Oliver Stringer
Shakti Prasad Sethi

## Slide 15
Photon rate
11th BGC Collaboration Meeting