## Slide 2
Vacuum Systems
Slot 2
V. Baglin
CERN TE-VSC, Geneva

## Slide 3
Outline
Measure of vacuum
Production of vacuum
Summary

## Slide 4
1. Measure of vacuum

## Slide 5
1.1 Vacuum gauges

## Slide 6: Vacuum gauges pressure range
Pa
104
102
10
10-2
10-4
10-6
10-8
10-10
10-12
mbar
102
10
10-2
10-4
10-6
10-8
10-10
10-12
10-14
Capacitance
Spinning rotor
Piezo
Pirani
Penning
Bayard-Alpert
Helmer, extractor
16 orders of magnitude !

## Slide 7
Pirani Gauge
Pirani gauges are used in the range 1 atm -10-4 mbar to monitor the roughing of a vacuum system
Robust, gas dependent, accuracy 10-100 %
Operating principle based on the variation of the heat conductivity through a gas: balancing a Wheatstone bridge to keep power through filament constant
   pressure reading above 1 mbar are wrong
Inficon Pirani gauge
Pfeiffer Pirani gauge
Edwards Pirani gauge
wikipedia
M. Pirani 1906
Tungsten wire

## Slide 8
Penning gauge
Penning gauges are commonly used in the range 10-5 -10-10 mbar. They are use for interlocking purposes
Robust, gas dependent, accuracy ~ 20-50 %
It is a cold cathode ionisation gauge i.e. there are no hot filament: electron are produced by field emission
The operating principle is based on the measurement of a discharge current in a Penning cell which is a function of pressure : I+ = Pn, n is close to 1
F. Penning 1937
B
Pfeiffer Penning gauge

## Slide 9: A discharge in a Penning gauge
Penning gauge ON behind a glass window
In the dark, the discharge is seen: a plasma is created
Pictures courtesy B. Henrist, TE-VSC
Penning gauge: a cloud of electrons is kept in place by a solenoidal magnetic field created by a permanent magnet ring
Anode rod

## Slide 10
Bayard-Alpert Gauge
Bayard-Alpert gauges are used for vacuum measurement purposes in the range 10-6 -10-12 mbar.
It is a hot filament ionization gauge. Electrons emitted by the filament perform a few oscillations inside the grid and ionise the molecules of the residual gas. Ions are then collected by an electrode.
Ion collector = 0 V
Filament = + 50 V
Grid = + 150 V
Modulator = + 150 V
L.G. Pittaway. J. of Phys. D: Appl. Phys. 3, 1113-1121, 1970
Courtesy B. Jenninger, TE-VSC
Electron trajectories for three different angles of incidence at the grid boundary, in a cross section of a Bayard-Alpert gauge

## Slide 11: A burned filament
Pictures courtesy B. Henrist, TE-VSC
Obviously, even if there are 2 filaments, the gauge is polluted therefore the pressure measurement will not be correct !
It is wise to exchange the vacuum gauge

## Slide 12
Bayard-Alpert Gauge: Sensitivity
The ion current collection can be described by:
The gauge needs to be calibrated for several gases
SN2 ~ 40 mbar -1
The pressure reading is expressed in nitrogen equivalent
In UHV, typical collected currents are in the pA range
Where :
I+ is the ion current
I- is the filament current
σ is the ionization cross section
n the gas density
L the electron path length
The vacuum gauge sensitivity, S in mbar -1, is defined by:
Ie (mA) | P (mbar) | I+ (pA)
4 | 10-10 | 16
4 | 10-12 | 0.16
0.1 | 10-10 | 0.4
0.1 | 10-12 | 0.004

## Slide 13
Bayard-Alpert Gauge: Sensitivity
The sensitivity can be measured and also computed from simulations
B. Jenninger et al. Vacuum 138 (2017) 173-177
 | H2 | He | CH4 | Ne | N2 | CO | C2H6 | Ar | CO2 | Xe
Si (mbar-1) | 19.06 | 7.46 | 60.62 | 10.48 | 41.84 | 42.30 | 114.71 | 53.19 | 54.48 | 7.50
Srel,i (mbar,i/mbar,N2) | 2.20 | 5.61 | 0.69 | 3.41 | 1.0 | 0.99 | 0.36 | 0.79 | 0.77 | 4.83
The sensitivity relative error equals ~ 10 %
A pressure reading in N2 equivalent is: Pread =2 10-10 mbar, in the case the main molecular species is H2, the real pressure is: P = Srel,H2 x Pread = 4.4 10-10 mbar
Relative sensitivity:
Space charge
effect
4 mA

## Slide 14
1.2 Gas analysis

## Slide 15: Residual gas composition
Q: “Why do we need to know the gas composition in an accelerator?”
As already mentioned, the beam circulating in the accelerator has a probability to collide with the residual gas.
The beam can interact via elastic and inelastic processes (energy conserved or not), through many physical mechanisms.
Here an example for the BESSY-II light source in Berlin (1.7 GeV, e- accelerator).
Ref. https://accelconf.web.cern.ch/p99/PAPERS/THA104.PDF

## Slide 16
Residual Gas Analysers
Residual Gas Analysers are used in the range 10-4 -10-13 mbar. Their purpose is to do gas analysis to define the gas composition.
A filament produces electrons which ionise the residual gas inside a grid.
A mass filter is introduced between the grid and the ion collector.
The ion current can be measured in Faraday mode or in secondary electron multiplier mode.
G.J. Peter, N. Müller. CAS Vacuum in accelerators CERN 2007-003
Range: 10-14 till 10-5 A
Gain ~ 103 -104 electrons/ion
ΔM at FHWM = 0.5 AMU
1 mA
Q ~ 2 10-9 mbar.l/s
Picture Pfeiffer

## Slide 17
RGA: Cracking pattern
Ions produced inside the grid can be fragmented into sub-species by the collisions with electrons
Example Mass 28
Nitrogen is traced by mass:
  28 (N2+)
  14 (N+): I14 = 0.14 I28
Carbon monoxide is traced by mass:
  28 (CO+)
  12 (C+): I12 = 0.06 I28
Carbon dioxide is traced by mass:
  44 (CO2+)
  28 (CO+): I28 = 0.13 I44
  16 (O+): I16 = 0.16 I44
  12 (C+): I12 = 0.10 I44
The table gives the percentage of the fragments with respect to the main peak

## Slide 18
RGA: Spectrum
A typical spectrum of a baked* vacuum system: P = 10-10 mbar
28
44
40
2
15
12
16
19
M/e | Ion | Molecule
2 | H2+ | H2
15 | CH3+ | CH4
28 | CO+ | CO
40 | Ar+ | Ar
44 | CO2+ | CO2
12 | C+ | CO+CO2
16 | O+ | Filament artefacts
19 | F+ | 
H2
Di-Hydrogen
M/e=2;1
CH4
Methane
M/e=15;(16)
H2O
Water
M/e=18;17;16
CO
Carbon monoxide
M/e=28;12
CO2
Carbon dioxide
M/e=44;16;28;12
N2
Nitrogen
M/e=28;14
Ar
Argon
M/e=40;20
O2
Oxygen
M/e=32;16
* Baked vacuum system:
System which has received a thermal cycle above ~150 °C in order to remove H2O vapor

## Slide 19
RGA: Partial Pressure
The RGA needs to be calibrated against a total pressure gauge for standard gases.
To consider the RGA ageing, relative sensitives are used and a normalization factor, K, is introduced
 | H2 | CH4 | N2 | CO | Ar | CO2
Srel,i,RGA (Torr N2/Torri) | 1.09 | 0.93 | 0.99 | 1 | 1.18 | 1.96
According to Dalton’s law, the reading given by the total pressure gauge should be equal to the sum of the partial pressures, expressed in nitrogen equivalent:
So:
Therefore:

## Slide 20
Typical spectrum of a baked vacuum system
Baked system with P= 10-10 mbar N2-eq
“Baking” means heating under vacuum above 150 ºC for >24h, to remove water vapor; It is a pre-requisite to getting pressures below 10-8 mbar quickly)
 | H2 | CH4 | CO | Ar | CO2
I (A) | 1·10-9 | 2 ·10-12 | 7·10-11 | 5·10-13 | 2·10-11
P (mbar) | 2·10-10 | 1·10-13 | 6·10-12 | 4·10-14 | 3·10-12
% Pi | 96 | 0 | 3 | 0 | 1
28
44
2
Note: a simple estimation from the total pressure measurement would give P = 2.2·10-10 mbar!

## Slide 21
Typical spectrum in the presence of an air leak
RGA are also useful to identify, trace and estimate air leaks
Air leak
M/e | Ion | Molecule
14 | N+ | N2
28 | N2+ | N2
32 | O2+ | O2
40 | Ar+ | Ar
Assume an 8 cm diam. tube with specific conductance ~ 60 ℓ·m/s:
  a leak detected 10 m away from the gauge (at P= 8·10-7 mbar)
  Hence, a pumping speed of 6 ℓ/s at the level of the leak gives a leak rate of:
    6 x 8·10-7 = 4.8·10-6 mbar·ℓ/s
Beware, oxygen being highly chemically reactive, its mass is not always present in the spectrum!

## Slide 22
2. Production of vacuum

## Slide 23: Vacuum pumps pressure range
Pa
104
102
10
10-2
10-4
10-6
10-8
10-10
10-12
mbar
102
10
10-2
10-4
10-6
10-8
10-10
10-12
10-14
Oil sealed mech. pump
Turbo molecular pump
Sputter ion pump
Cryopumps
Ti sublimation pump
NEG & cryogenic pumps
16 orders of magnitude !

## Slide 24
Turbomolecular pumping group
Used to pump down from atmosphere and commission vacuum sectors down to 10-11 mbar
Mobile system based on rotary vane primary pump and turbomolecular pump
Valve
Turbomolecular pump
Primary pump
Primary pump
Turbomolecular pump
Primary pump
W. Becker, 1956
A. Pfeiffer GmbH

## Slide 25
Sputter Ion Pump
This pump operates in the range 10-5 -10-11 mbar. It is used to maintain the pressure in the vacuum chamber of an accelerator. It has NO exhaust, it traps the gas species.
Typical pumping speed range from 1 to 500 l/s (but 1000-2000 l/s custom models exist)
Titanium sputtered from the cathode bombarded by accelerated ions provides pumping.
The ion current is proportional to pressure, hence these pumps are often used for interlocks, in place of gauges
Pictures Agilent (Varian)
Hi-Voltage feed-thru
3-7 kV

## Slide 26: NEG cartridge
This pumps operates in the UHV regime in a range < 10-8 mbar
It is used as a complementary pumping system to e.g. ion pumps.
Zr based getter materials are sintered and form porous disks.
After activation at ~ 500ºC for 1h, the pumping speed can be large: 100 – 2 000 l/s
Large sorption capacity: >0.1 – 10 mbar.l for N2, CO and > 10 mbar.l for H2!
  Reminder: 1 mbar.l = 4.3 1019 molecules, ~ 1 monolayers of a 10m long,Ø10 tube
H2 diffuse into the bulk, CO is adsorbed in 1 active site whereas N2 requires 6.
Gas mixture: CO adsorption inhibit H2 and N2 pumping
Courtesy G. Bregliozzi, J. Callagher, 2012

## Slide 27
Non-Evaporable Getter (NEG)
P. Chiggiato and P. Costa Pinto, Thin Solid Films, 515 (2006) 382-388
Getters (eg Ti) are materials capable of chemically adsorbing gas molecules. To do so their surface must be clean.
For Non-Evaporable Getters (eg TiZrV films) a clean surface is obtained by heating to a temperature high enough to dissolve the native oxide layer into the bulk.
NEGs pump most of the gas except rare gases and methane at room temperature
1 µm thick film coated at 300˚C
Very large pumping speed : ~ 250 l/s/m for H2,
20 000 l/s/m for CO  distributed pumping
Very low outgassing rate (~200 CH4/s/cm2)
But : limited capacity and fragile coating, sensitive to pollutant (hydrocarbons, Fluor …)
C. Benvenuti, 1996

## Slide 28: Summary
In order to measure the total pressure one has to use a combination of vacuum gauges, adopting different technologies (although new “integrated” gauges are on the market, where one instrument does it all, from ambient pressure down to UHV)
Typical gauges are: Capacitance, Spinning Rotor, Piezoelectric, Pirani, Penning, Bayard-Alpert, Helmer, extractor;
In addition to the total pressure one needs to know also the residual gas composition, since the stored particle beams interact in different ways and intensities in presence of different gas species: the elastic and inelastic cross-sections of interaction (residual gas nuclei and electrons) have a strong dependence on the atomic number of the residual gas molecules/atoms.
Residual Gas Analyzers (RGAs) are typically used for the purpose of determining the gas composition; typically they are of the four-rod type, spanning 100-200 a.m.u.’s.
For a well-baked vacuum system with no leaks, the typical gas composition is H2, CH4, CO, CO2, with H2 being the biggest component (good, as it has lowest Z)
As for pumps, not a single pump can work in the full pressure range, from atmosphere (pumpdown phase) to UHV.
Typical pumps used on particle accelerators are: turbopumps, ion-sputtering pump, NEG, cryogenic; NEG-coating is also very much used on modern accelerators (0.2~2 mm-thick layer of NEG material, deposited via physical vapor deposition techniques, sputtering; NEG: ternary alloy of Ti, Zr, V.)

## Slide 29
Some References
Cern Accelerator School, Vacuum technology, CERN 99-05
Cern Accelerator School, Vacuum in accelerators, CERN 2007-03
Cern Accelerator School, Vacuum for particle accelerators, CERN-ACC-2020-0009
The physical basis of ultra-high vacuum, P.A. Redhead, J.P. Hobson, E.V. Kornelsen,
American Vacuum Soc.
Scientific foundations of vacuum technique, S. Dushman, J.M Lafferty. J. Wiley & sons.
Handbook of Vacuum Technology, K. Jousten (ed.), Wiley VCH
Vacuum Technology, A. Roth. Elsevier Science
Foundations of vacuum science and technology, Ed by J.M. Lafferty. J. Wiley & sons.
Some Journals Related to Vacuum Technology for Particle Accelerators
Journal of vacuum science and technology A and B
Vacuum
Nuclear Instruments and Technology (A and B)
Review Modern Instruments
Applied Physics
Fusion Engineering and Design

## Slide 30
Thank you for your attention !!!