## Slide 1
Measurement of transverse Beam Emittance
JUAS 2025, ESI-Archamps at CERN
Peter Forck (GSI and University Frankfurt)

## Slide 2
Measurement of transverse Emittance
The emittance characterizes the whole beam quality, assuming linear
behavior as described by second order differential equation.
It is defined within the phase space as:
The measurement is based on determination of:
either profile width σx and angular width σx′ at one location
or σx at different locations and linear transformations.
Synchrotron: lattice functions results in stability criterion
 beam width delivers emittance:
Different devices are used at transfer lines:
All beams: Quadrupole variation & ’three grid’ method using linear transformations
(not well suited in the presence of non-linear forces)
Lower energies Ekin < 100 MeV/u: slit-grid device, pepper-pot
(suited in case of non-linear forces).

## Slide 3
Outline:
Definition and some properties of transverse emittance
Quadrupole strength variation and position measurement
Slit-Grid device: scanning method
Summary

## Slide 4
Particle Trajectory and Characterization of many Particles
Single particle trajectories are forming a beam
They have a distribution of
start positions and angles
Characteristic quantity is
the beam envelope
 Goal:
Behavior of whole ensemble
Plot: Wille
Focus. quad.
drift
envelope of all particles
single particles
trajectory
Focus. quad.
drift

## Slide 5
Definition of Offset and Divergence
x
x’
x
s
Horizontal and vertical coordinates at s = 0 :
x : Offset from reference orbit in [mm]
x’: Angle of trajectory in unit [mrad]
x’ = dx / ds
Assumption: par-axial beams:
x is small compared to dipole radius ρ0
Small angle with px / ps << 1
Longitudinal coordinate:
Longitudinal orbit difference: l = - v0 ∙ ( t – t0 ) in unit [mm]
Momentum deviation: δ = ( p – p0 ) / p0 sometimes in unit [mrad] = [‰]
Reference particle: No offset x  y  l  0 & no angle x’  y’  δ  0 for all s
For continuous beam: l has no meaning  set l  0 !
l

## Slide 6
Definition of Coordinates and Taylor Expansion B-Field
Definition compared to reference orbit
x: horizontal
y: vertical
s: in beam direction
ρ0: radius of curvature
 on reference orbit
Taylor expansion of magnetic field at location s for field in y direction:
Drift (ρ0 ), dipoles and quadrupoles  linear beam optics
x
y
s
ρ0
beam

## Slide 7
Excurse: Definition of Coordinates and basic Equations
The basic vector is
6 dimensional:
The transformation of a single particle from a location s0 to s1 is given by the Transfer Matrix R:
Transfer Matrix is a 6x6 matrix 

## Slide 8
Some Properties of the Transfer Matrix
The transformation can be done successive:
for R1 = R( s0  s1 ) ,…, Rn = R( sn-1  sn )
It is R = Rn ∙ Rn-1∙ … ∙ R1 as it is a linear transformation
The matrix elements describe the coupling between the components
If all forces are symmetric along
the reference orbit
than the horizontal and
vertical plane are decoupled:
sub-matrix is sufficient
It is det(R) = 1 (Liouville’s Theorem)
i.e. R is invertible
Sub-matrix
here shown for drift L:

## Slide 9
Some Examples for linear Transformations
The 2-dim sub-space (x,x’) can be used in case there is no coupling like dispersion R16 =(x|δ)=0
Important examples are:
Drift with length L:
Horizontal focusing with quadrupole constant k end effective length l:
Horizontal de-focusing with quadrupole constant k end effective length l:
Ideal quad.: field gradient g = Bpole/a, Bpole field at poles, a aperture
 quadrupole constant k = | g | / (B)0
Thin lens approximation: l
 simple transfer matrix (math. proof by 1st order Taylor expansion)
Kick: Δx’ = - x/f
s
Horizontal focusing:
Δx’
f

## Slide 10
Imaging using Quadropules: Here Duplett Focusing
Transfer matrix:
Rchannel= Rdrift(b) ∙ Rde-focus(kL) ∙ Rdrift(l) ∙ Rfocus(kL) ∙ Rdrift(a)
y
Horizontal
focusing
Horizontal
de-focusing
L
L
l
b
a
trajectories
Plot: Rossbach, Schmüser
Single particle trajectories
forming a beam
Characteristic quantity is
the beam envelope
Goal: Description of
whole ensemble behavior!

## Slide 11
Definition of Beam Matrix
Transformation of a the envelope from s0 to s1 is given by the Beam Matrix :
6-dim Beam Matrix with decoupled hor. & vert. plane:
Horizontal
beam matrix:
Beam width for
the three
coordinates:
horizontal
vertical
longitudinal
hor.-long. coupling
10 values
all couplings
 21 values
Transformation of a single particle from s0 to s1 is given by the Transfer Matrix R:
The basic 6 dimensional vextor is:

## Slide 12
Geometrical Parameters of Phase Space Ellipsoid
The phase space distribution is
described by an ellipse
with the characteristics parameters
a
Area: A=  ab = 
b
x’
x
Corresponding parameters:
Beam Matrix | Statistic | Twiss Parameter
11 | <x2> | 
22 | <x’2> | 
12 | <xx’> | -
The determinate is preserved:
Proof:
 det() = det(R) ) det(RT)
= )
For quadratic matrices and :

## Slide 13
Geometrical Parameters of Phase Space Ellipsoid
The phase space distribution is
described by an ellipse
with the characteristics parameters
Corresponding parameters:
Beam Matrix | Statistic | Twiss Parameter
11 | <x2> | 
22 | <x’2> | 
12 | <xx’> | -
The determinate is preserved:
Gaussian distribution for an ensemble of particles

## Slide 14
The Emittance for Gaussian and non-Gaussian Beams
The beam distribution can be non-Gaussian, e.g. at:
 beams behind ion source
space charged dominated beams at LINAC & synchrotron
cooled beams in storage rings
General description of emittance
using terms of 2-dim distribution:
It describes the value for 1 standard derivation
Care:
No common definition
of emittance concerning
the fraction f
For Gaussian beams only: εrms  interpreted as area containing a fraction f of ions:
Emittance (f) | Fraction f
1  rms | 15 %
  rms | 39 %
2  rms | 63 %
4  rms | 86 %
8  rms | 98 %
Variances
Covariance
i.e. correlation

## Slide 15
Poll concerning linear Transformations
Poll 4.2:
What is wrong concerning linear transformations?
The trajectories can be calculated by a matrix.
The coordinate vector is
All elements within a typical accelerator can be described by linear transformations.
The emittance is preserved.
Poll 4.1:
Which optical elements are non-linear?
Dipole
Quadrupole
Sextupole
www.menti.com
code: 1208 6026
Notes: 15

## Slide 16
Outline:
Definition and some properties of transverse emittance
Quadrupole strength variation and position measurement
emittance from several profile measurement and beam optical calculation
Slit-Grid device: scanning method
Summary

## Slide 17
Emittance Measurement by Quadrupole Variation
From a profile determination, the emittance can be calculated via linear transformation,
if a well known and constant distribution (e.g. Gaussian) is assumed.
The basis is beam matrix transformation
Measurement of beam width
x2max = σ11(1, k)
matrix R(k) describes the focusing.
With the drift matrix the transfer is
R(ki ) = Rdrift  Rfocus(ki )
Transformation of the beam matrix
(1,ki) = R(ki )  (0)  RT (ki )
Task: Calculation of matrix (0)
at entrance s0 i.e. all three elements

## Slide 18
Measurement of transverse Emittance
Assumptions:  Constant emittance, in particular no space-charge broadening
 Only elliptical shaped beam distribution is considered
 No misalignment, i.e. beam center equals center of the quadrupoles
 If not valid: A self-consistent algorithm can be used.
The beam width xmax(s1) at s1 is measured  matrix element σ11(1, ki) = x2max (ki)
Different focusing of quadrupoles k1 , k2 ... kn are used  Rfocus(ki)
After the drift the transfer matrix is R(ki) = Rdrift  Rfocus(ki)
Task: Calculation of beam matrix (0) at entrance s0 (matrix elements give orientation)
The transformation of the beam matrix is: (1,ki) = R(ki )  (0)  RT (ki )
 Result: Redundant system of linear equations for matrix elements ij(0)
+ focusing k1

+ focusing kn
To have an error estimation at least three measurements must be done

## Slide 19
Measurement of transverse Emittance
Example: Square of the beam width at
ELETTRA 100 MeV e- Linac, YAG:Ce:
Using the ‘thin lens approximation’ i.e. the quadrupole has a focal length of f:
Measurement of the matrix-element (1,K) from (1,K) = R(K) ∙ (0) ∙ RT(K)
G. Penco (ELETTRA) et al., EPAC’08
For completeness: The relevant formulas
The three matrix elements at the quadrupole:
Focusing strength K [m-1]
beam width

## Slide 20
The ‘Three Grid Method’ for Emittance Measurement
Instead of quadrupole variation, the beam width is measured at different locations:
The procedure is:
Beam width x(i) measured
at the locations si
 beam matrix element
x2(i) = σ11(i).
The transfer matrix R(i) is known.
(without dipole a 3 × 3 matrix.)
The transformations are:
σ(i) = R(i)  σ(0)  RT(i)
 redundant equations:
 Result: at least equations for elements ij(0)
+ for

+ for

## Slide 21
Results of a ‘Three Grid Method’ Measurement
Solution: Solving the linear equations like for quadrupole variation
or fitting the profiles with linear optics code (e.g. MADX).
Example: The hor. and vert. beam envelope and the beam width at a transfer line:
Assumptions:  constant emittance, in particular no space-charge broadening
100 % transmission i.e. no loss due to vacuum pipe scraping
 no misalignment, i.e. beam center equals center of the quadrupoles.
Measured vertical width
Measured horizontal width
F DD F
D F
D F
D D
D F
Fitted beam vert. envelope
Path Length [m]
0
70
Vertical size [mm]
Hori. size [mm]
50
50
0
dipole
dipole
Fitted beam hori. envelope

## Slide 22
Poll concerning Emittance Measurement by linear Transformations
Poll 4.3:
What is correct concerning the emittance measurement by linear transformations?
Only the 2nd statistical moments are determined to characterize the beam distribution.
Any beam distribution can be reliably determined by this methods.
The transfer line can contain any optical element.
In all cases, three measurements are sufficient for an emittance determination.
www.menti.com
code: 1208 6026
Notes: 22

## Slide 23
Outline:
Definition and some properties of transverse emittance
Quadrupole strength variation and position measurement
emittance from several profile measurement and beam optical calculation
Slit-Grid device: scanning method
scanning slit  beam position & grid  angular distribution
Summary

## Slide 24
The Slit-Grid Measurement Device
Slit-Grid: Direct determination of position and angle distribution.
Used for protons with Ekin < 100 MeV/u  range R < 1 cm.
Slit: position P(x) with typical width: 0.1 to 0.5 mm
Distance: typ. 0.5 to 5 m (depending on beam energy 0.1 ... 100 MeV)
SEM-Grid: angle distribution P(x′)
LINAC
RFQ
SEM-Grid
Slit
0.5...5 m

## Slide 25
Slit & SEM-Grid
Beam surface interaction: e− emission
 measurement of current.
Example: 15 wire spaced by 1.5 mm:
Each wire is equipped with one I/U converter
different ranges settings by Ri
 very large dynamic range up to 106.
Example: Slit of width 0.1 mm (defect)
Moved by stepping motor, 0.1 mm resolution
Slit with e.g. 0.1 mm thickness
 Transmission only from Δx.

## Slide 26
Display of Measurement Results
The distribution of the ions is depicted as a function of
Position [mm]
Angle [mrad]
The distribution can be visualized by
Mountain plot
Contour plot
Calc. of 2nd moments <x2> , <x’2> & <xx’>
Emittance value εrms from
Beam: Ar4+, 60 KeV, 15 μA
at Spiral2 Phoenix ECR source.
Plot from P. Ausset, DIPAC 2009
Problems:
Finite binning results in limited resolution
Background  large influence on <x2>, <x’2> and <xx’>
Or fit of distribution i.e. ellipse to data
Effective emittance only
Modern alternative: Tomographic reconstruction

## Slide 27
The Resolution of a Slit-Grid Device
The width of the slit dslit gives the resolution in space Δx = dslit.
The angle resolution is Δ x′ = (dwire + 2rwire)/d
 discretization element Δx ・ Δx′.
By scanning the SEM-grid the angle resolution can be improved.
Problems for small beam sizes or parallel beams.
For pulsed LINACs: Only one measurement each pulse  long measuring time required.

## Slide 28
Poll concerning normalized Emittance
Poll 4.5:
Assuming an electron and proton beam of 10 MeV (starting from the same source). What do expect concerning the emittance of both beams?
The electron beam has a smaller emittance than the proton beam.
The electron beam has a larger emittance than the proton beam.
Both beams have the same emittance as it relates to the beam energy.
Poll 4.4:
A typical slit-grid emittance installation can be applied for …
all beams and all energies
electron beams for all energies
proton beam energies below 100 MeV only
proton beam energies above 100 MeV only
www.menti.com
code: 1208 6026
Notes: 28

## Slide 29
MEVVA
MUCIS
PIG
RFQ IH1 IH2
Alvarez DTL
HLI: (ECR,RFQ,IH)
Transfer to
Synchrotron
2.2 keV/u
β = 0.0022
120 keV/u
β = 0.016
11.4 MeV/u
β = 0.16
Gas Stripper
U4+
U28+
All ions, high current, 5 ms@50 Hz, 36&108 MHz
Foil Stripper
To SIS ↑
Constructed in the 70th, Upgrade 1999,
further upgrades in preparation
Slit Grid Emittance: Standard device, total 9 device
Pepper-pot Emittance: Special device, total 1 device
1.4 MeV/u ⇔ β = 0.054
 Conclusion
Appendix GSI Ion LINAC: Emittance Measurement Devices
Remark: For higher energies where ions cannot be stopped,
quadrupole variation is used.

## Slide 30
Poll concerning normalized Emittance
Poll 4.6:
In a synchrotron no special emittance measurement for a stable circulating beam is discussed.
What is the reason?
The emittance has no practical meaning.
The emittance of a circulating beam is always constant.
For emittance measurement, the beam must be extracted from the synchrotron and methods in transfer lines must be applied
For a stable circulating beam, the emittance is directly linked to the transverse profile as related to a relation between Twiss parameter ,  and  at one location.
www.menti.com
code: 1208 6026
Notes: 30

## Slide 31
Summary for transverse Emittance Measurement
Emittance measurements are very important for comparison to theory.
It includes size (value of ε) and orientation in phase space (ij or ,  and  )
i.e three independent values
Low energy beams  direct measurement of x- and x′-distribution
 Slit-grid: movable slit  x-profile, grid  x′-profile
 Variances exists: slit-slit, slit-kick, pepperpot .... method
All beams  profile measurement + linear transformation:
 Quadrupole variation: one location, different setting of a quadrupole
 ’Three grid method’: different locations
 Assumptions:  well aligned beam, no steering
 no emittance blow-up due to space charge.
Important remark: For a synchrotron with a stable beam storage,
width measurement is sufficient using

## Slide 32
Backup slides

## Slide 33
Quiz related to Beam Emittance and Slit-Grid Method
What is meant by beam quality = emittance, i.e.
what type of quantitative description is used?
What is the geometrical meaning of a phase space distribution for -  12  <xx’>  0 ?
Expectation: Is the emittance of a 10 MeV electron beam smaller or larger than a proton beam?
Describe shortly the measurement principle of a slit-grid method!
Do you know variances of the slit-grid method?
Why is a slit-grid measurement of 100 MeV electron and proton beam not meaningful?
 Back to quadrupole variation
Notes: 33

## Slide 34
Quiz related to Beam Emittance Measurement
The emittance is defined via statistical parameters <x2>, <xx’>, <x’2>. What does this mean?
What might causes the beam to be different from a Gaussian (i.e. simple) distribution?
What is the principle of an emittance measurement for quadrupole variation?
What is the principle for the 3 grid method?
Can you imaging how a quadrupole variation or 3-grid method must be modified if the beam is ‘non-Gaussian’?
Notes: 34

## Slide 35
Outline:
Definition and some properties of transverse emittance
Slit-Grid device: scanning method
scanning slit  beam position & grid  angular distribution
Pepper-pot device: single shot device
hole-plate  beam position & screen  angular distribution
Quadrupole strength variation and position measurement
Summary

## Slide 36
Ansatz:
Beam matrix at one location:
It describes a 2-dim probability distr.
The value of emittance is:
For the profile and angular measurement:
Definition of transverse Emittance
The emittance characterizes the whole beam quality:
Geometrical interpretation:
All points x fulfilling xt · σ -1 · x = 1
are located on a ellipse
σ22x2 – 2σ12xx‘ + σ11x‘2 = det σ = εx2

## Slide 37
The density function for a 2-dim Gaussian distribution is:
It describes an ellipse with the characteristics
profile and angle Gaussian distribution of width
and the correlation or covariance
The Emittance for Gaussian Beams

## Slide 38
The Emittance for Gaussian and non-Gaussian Beams
The beam distribution can be non-Gaussian, e.g. at:
 beams behind ion source
space charged dominated beams at LINAC & synchrotron
cooled beams in storage rings
General description of emittance
using terms of 2-dim distribution:
It describes the value for 1 standard derivation
For discrete distribution:
and correspondingly for all other moments
Variances
Covariance
i.e. correlation

## Slide 39
The Noise Influence for Emittance Determination
A real measurement of beamlets contains:
Noise i.e. fluctuation of the output
Bias i.e. electrical offset from amplifier
 Strong influence of noise reduction to
numerical values of <x>, <x’2> and <xx’> and on εrms
Algorithm & cut-level must be given for evaluation
General: Typical error Δε/ε > 10 %
Unbiased
vs.Unbiased
Unbiased,
Negatives
zeroed
Unbiased
vs.Unbiased
Unbiased,
Negatives
zeroed
Biased
Thresholded
Biased
Thresholded
Example: Dependence of εrms on threshold
From M. Stockli, BIW 2006 and
J. Pfister, Disseration Uni Frankfurt 2010
Threshold level of max. [%]
Emittance  [arb.u.]

## Slide 40
The Pepperpot Emittance Device
For pulsed LINAC: Measurement within one pulse is an advantage
If horizontal and vertical direction coupled  2-dim evaluation required
(e.g. for ECR ion source)
Good spatial resolution if many holes are illuminated.
Good angle resolution only if spots do not overlap.
Partly from H.R. Kremers et al., ECRIS 2010

## Slide 41
The Pepperpot Emittance Device at GSI UNILAC
Example GSI-LINAC 0.12 to 11 MeV/u:
Pepper-pot: 15 × 15 holes with Ø 0.1mm
on a 50 × 50 mm2 copper plate
Distance: pepper-pot-screen: 25 cm
Screen: Al2O3, Ø 50 mm
Data acquisition: high resolution camera
beam
drive for
pepper-pot
camera
laser
Good spatial resolution if many holes are illuminated.
Good angle resolution only if spots do not overlap.

## Slide 42
Result of a Pepperpot Emittance Measurement
Example: Ar 1+ ion beam at 1.4 MeV/u,
screen image from single shot at GSI:
Data analysis:
Projection on
horizontal and vertical plane
→ analog to slit-grid.

## Slide 43
The Artist View of a Pepperpot Emittance Device