## Slide 1
Pick-Ups for bunched Beams
JUAS 2025, ESI-Archamps at CERN
Peter Forck (GSI and University Frankfurt)
vertical tune versus time
Qy  0.02
time [ms]
vert. fractional tune qy
750
600
400
0
200
0.2
0.25
0.3
excitation
noise band
Q  0.05

## Slide 2
Outline:
Time and frequency domain treatment & Fourier Transformation
Signal generation  transfer impedance
Capacitive button BPM for high frequencies
Capacitive linear-cut BPM for low frequencies
Electronics for position evaluation
BPMs for measurement of closed orbit, tune and further lattice functions
Summary
Pick-Ups for bunched Beams

## Slide 3
Usage of BPMs
A Beam Position Monitor is an non-destructive device for bunched beams
It has a low cut-off frequency i.e. dc-beam behavior can not be monitored
The abbreviation BPM and pick-up PU are synonyms
1. It delivers information about the transverse center of the beam
Trajectory: Position of an individual bunch within a transfer line or synchrotron
Closed orbit: Central orbit averaged over a period much longer than a betatron oscillation
Single bunch position: Determination of parameters like tune, chromaticity, -function
Bunch position on a large time scale: bunch-by-bunch  turn-by-turn  averaged position
Time evolution of a single bunch can be compared to ‘macro-particle tracking’ calculations
Feedback: fast bunch-by-bunch damping or precise (and slow) closed orbit correction
2. Information on longitudinal bunch behavior (see next chapter)
Bunch shape and evolution during storage and acceleration
For proton LINACs: the beam velocity can be determined by two BPMs
For electron LINACs: Phase measurement by Bunch Arrival Monitor
Relative low current measurement down to 10 nA.

## Slide 4
Ibeam
Time Domain  Frequency Domain: Instrumentation
Time domain: Recording of a voltage as a function of time:
Frequency domain: Displaying of a voltage as a function of frequency:
Instrument:
Spectrum Analyzer
Instrument:
Oscilloscope
courtesy company Keysight
voltage
time
frequency
Photos and idea by Piotr Kowina
Fields seems unstructured
Other perspective:
Fields is structured

## Slide 5
Ibeam
Time Domain  Frequency Domain: Mathematics
Time domain: Recording of a voltage as a function of time:
Frequency domain: Displaying of a voltage as a function of frequency:
Mathematics for function :
Fourier Transformation:
Fourier Transformation:
Contains amplitude & phase
as the values are (complex in math. sense)
The same information is displayed differently
Care: Fourier Transformation of time domain data contains amplitude and phase

## Slide 6
Excurse: Properties of Fourier Transformation
a process can be described either with time domain’ or ‘frequency domain’
We assume but ; note: integral from to
FT:  amplitude or power
& phase tech: displayed for

A
Re(z)
Im(z)
z
Linearity: for
No loss of information: If exists, than
 the original function f(t) is ‘recovered’, or ‘you neither gain nor lose information’
Fourier Transform.:
tech. DFT(f)
Inv. F. T.:
tech. IDFT(f)

## Slide 7
Excurse: Properties of Fourier Transformation
Fourier Transform.:
tech. DFT(f)
Inv. F. T.:
tech. IDFT(f)
a process can be described either with time domain’ or ‘frequency domain’
Similarity Law: For a  0 it is for :
the properties can be scaled to any frequency range;  ‘shorter time signal has wider FT’ , e.g. Gaussian  with
Differentiation Law: For nth derivative f(n)(t) it is:
 differentiation in time domain corresponds to multiplication with i in frequency domain
Frequency shift: Modulation by 0 it is

## Slide 8
Excurse: Properties of Fourier Transformation
Condition for FT: If and convergent, than is Fourier transformable
 e.g. f(t) = sin(0t), polynomials f(t) =  ai ti and f(t)=et are not Fourier transformable !
Windowing: Def.
with a window function  finite integration
 as convolution
Example: 
Zero-crossing at:
Scaling of maxima:

## Slide 9
Excurse: Fourier Trans.  Convolution & technical Realization
Application: Chain of electrical elements calculated in frequency domain more easily
parameters are more easy in frequency domain (bandwidth, f-dependent amplification.....)
Engineering formulation for finite number of discrete samples:
Digital Fourier Transformation DFT(f): corresponds to math. FT for finite number time
Fast Fourier Transformation: FFT(f) dedicated algorithm for fast calc. with 2n increments
Transfer function H() and h(t) are used to describe electrical elements
Calculation with H() in frequency domain or h(t) time domain
 ‘Finite Impulse Response’ FIR filter or ‘Infinite Impulse Response’ IIR filter
Uin(t)
amplifier
high pass filter
rectifier
Uout(t)
low pass filter
Convolution Law: For the convolution
 FT can be calculated as
convolution in time domain can be expressed as multiplication of FT in frequency domain

## Slide 10
Poll concerning Fourier Transformation
Poll 5.1:
Which statement is wrong?
Fourier Transformation is a integral transformation
Any continuous function can be Fourier transformed
Experimentally, the Fourier Transformation amplitude is observed by a spectrum analyser
Poll 5.2:
What is correct concerning the Fourier Transformation (FT) in general?
Convolution in time domain can be calculated as a multiplication of the individual FTs
Convolution in time domain can be calculated by multiplication with the factor of the FT
The amplitude of the FT contains all required information
The mathematically calculated FT can not be displayed in an appropriate manner experientially
www.menti.com
code: 1208 6026
Notes: 10

## Slide 11
Outline:
Time and frequency domain treatment & Fourier Transformation
Signal generation  transfer impedance
Capacitive button BPM for high frequencies
Capacitive linear-cut BPM for low frequencies
Electronics for position evaluation
BPMs for measurement of closed orbit, tune and further lattice functions
Summary
Pick-Ups for bunched Beams

## Slide 12
Pick-Ups for bunched Beams
The image current at the beam pipe is monitored on a high frequency basis
i.e. the ac-part given by the bunched beam.
Beam Position Monitor BPM equals Pick-Up PU
Most frequent used instrument!
For relativistic velocities,
the electric field is transversal:

## Slide 13
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
V
-
-
-
-
-
-
-
Principle of Signal Generation of capacitive BPMs
Animation by Rhodri Jones (CERN)
The image current at the wall is monitored on a high frequency basis
i.e. ac-part given by the bunched beam.

## Slide 14
Model for Signal Treatment of capacitive BPMs
The wall current is monitored by a plate or ring inserted in the beam pipe:
The image current Iim at the plate is given by the beam current and geometry:
Using a relation for Fourier transformation:
Better mathematical deviation: See book section 5.3

## Slide 15
Transfer Impedance for a capacitive BPM
At a resistor R the voltage Uim from the image current is measured.
The transfer impedance Zt is the ratio between voltage Uim and beam current Ibeam
in frequency domain: .
This is a high-pass characteristic with ωcut= 1/RC:
Capacitive BPM:
The pick-up capacitance C:
plate ↔ vacuum-pipe and cable.
The amplifier with input resistor R.
The beam is a high-impedance current source:
Amplitude:
Phase:

## Slide 16
Example of Transfer Impedance for Proton Synchrotron
The high-pass characteristic for typical synchrotron BPM:
Large signal strength  high impedance
Smooth signal transmission  50 
Remark: For   0 it is Zt  0 i.e. no signal is transferred from dc-beams e.g.
 de-bunched beam inside a synchrotron
 for slow extraction in a transfer line
Parameter for linear-cut BPM:
C = 100pF, l = 10cm,  = 50%
fcut= ω/2π=(2πRC)-1
for R=50   fcut= 32 MHz
for R=1 M  fcut= 1.6 kHz
fcut
decrease by
fac. 1/ 2 = 3 dB

## Slide 17
Signal Shape for capacitive BPMs: differentiated  proportional
Depending on the frequency range and termination the signal looks different:
High frequency range ω >> ωcut :
 direct image of the bunch. Signal strength Zt  A/C i.e. nearly independent on length
Low frequency range ω << ωcut :
 derivative of bunch, single strength Zt  A, i.e. (nearly) independent on C
Intermediate frequency range ω  ωcut: Calculation using Fourier transformation
Example: Synchrotron BPM with 50  termination (reality at p-synchrotron : σ>>1 ns):
Ibeam(t)
Uim(t)

## Slide 18
Calculation of Signal Shape (here single bunch)
The transfer impedance is used in frequency domain! The following is performed:
1. Start: Time domain Gaussian function Ibeam(t) having a width of σt
2. FFT of Ibeam(t) leads to the frequency domain Gaussian Ibeam(f) with σf = ( 2πσt )-1
3. Multiplication with Zt(f) with fcut=32 MHz leads to Uim(f) = Zt(f) ∙ Ibeam(f)
4. Inverse FFT leads to Uim(t)
Fourier trans.
inverse Fourier trans.

## Slide 19
Calculation of Signal Shape: repetitive Bunch in a Synchrotron
Synchrotron filled with 8 bunches accelerated with facc=1 MHz
BPM terminated with R= 1 M  facc >> fcut :
Parameter: R=1 M  fcut=2 kHz, Zt=5  , all buckets filled
C = 100pF, l = 10cm,  = 50%, σt = 100 ns  σl = 15m
Fourier spectrum is composed of lines separated by acceleration frf
Envelope given by single bunch Fourier transformation
Baseline shift due to ac-coupling
Remark: 1 MHz< frf <10 MHz  Bandwidth 100 MHz=10∙frf for broadband observation

## Slide 20
Calculation of Signal Shape: repetitive Bunch in a Synchrotron
Fourier spectrum is concentrated at acceleration harmonics
with single bunch spectrum as an envelope.
Bandwidth up to typically 10*facc
Synchrotron filled with 8 bunches accelerated with facc=1 MHz
BPM terminated with R=50   facc<< fcut :
Parameter: R=50   fcut=32 MHz, all buckets filled
C = 100pF, l = 10cm,  = 50%, σt = 100ns  σl = 15m

## Slide 21
Calculation of Signal Shape: Bunch Train with empty Buckets
Fourier spectrum is more complex, harmonics are broader due to sidebands
Synchrotron during filling: Empty buckets, R=50 :
Parameter: R=50   fcut=32 MHz, 2 empty buckets
C = 100pF, l = 10cm,  = 50%, σt = 100ns  σl = 15m

## Slide 22
Calculation of Signal Shape: Filtering of Harmonics
Remark: For numerical calculations, time domain filters (FIR and IIR) are more appropriate
Generally: Ztot(ω) = Hcable(ω) ∙ Hfilter(ω) ∙ Hamp(ω) ∙ … ∙Zt(ω)
nth order Butterworth filter,
math. simple, but not well suited:
Effect of filters, here bandpass:
Ringing due to sharp cutoff
Other filter types more appropriate
Parameter: R=50 , 4th order Butterworth filter fcut = 2 MHz
C = 100pF, l = 10cm,  = 50%, σ = 100 ns

## Slide 23
Examples for differentiated & proportional Shape
Proton LINAC, e—LINAC & synchrotron:
100 MHz < frf < 1 GHz typically
R=50  processing to reach bandwidth
C5 pF  fcut =1/(2πRC) 700 MHz
Example: 36 MHz GSI ion LINAC
Proton synchrotron:
1 MHz < frf < 30 MHz typically
R=1 M for large signal i.e. large Zt
C100 pF  fcut =1/(2πRC) 10 kHz
Example: non-relativistic GSI synchrotron
frf : 0.8 MHz  5 MHz
Remark: During acceleration the bunching-factor is increased: ‘adiabatic damping’.

## Slide 24
-
-
-
-
-
+
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
V
-
-
-
-
-
-
-
-
+
+
+
+
+
-
-
+
-
+
-
+
-
+
-
-
-
-
+
+
+
-
+
+
-
+
-
-
+
-
+
-
-
-
-
-
+
+
+
-
+
+
-
-
+
-
-
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
Animation by Rhodri Jones (CERN)
Principle of Signal Generation of a BPMs: off-center Beam
The image current at the wall is
monitored on a high frequency basis
i.e. ac-part given by the bunched beam.

## Slide 25
Principle of Position Determination by a BPM
The difference voltage between plates gives the beam’s center-of-mass
most frequent application
S(ω,x) is called position sensitivity, sometimes the inverse is used k(ω,x)=1/S(ω,x)
S is a geometry dependent, non-linear function,
Units: S=[%/mm], sometimes S=[dB/mm] or k=[mm].
It is at least:
y from ΔU = Uup - Udown
Typical desired position resolution:
x  0.1 ... 0.3  x of beam width
Example: One turn = 4 bunches @ 35 MeV/u
Uright
Ureft

## Slide 26
The Artist’s View of a BPM

## Slide 27
Poll concerning Transfer Impedance
Poll 5.3:
What is correct concerning the transfer impedance ?
It can only be applied in frequency domain
It is the ratio between the measurable voltage and beam current
It is the ratio between the measurable voltage and wall current
Poll 5.4:
What is correct concerning BPM measurements in general?
Beams of all types of time structure can be monitored
Only positive beam particles (i.e. protons & ions) can be measured
The BPM circuit acts as a high pass filter
Poll 5.5:
Using a BPM, the position is determined by opposite plates via…
Difference divided by sum
Sum divided by difference
Poll 5.6:
Using a BPM: What is true concerning the absolute value of the sum & difference :
Both cases can appear
www.menti.com
code: 1208 6026
Notes: 27

## Slide 28
Outline:
Signal generation  transfer impedance
Capacitive button BPM for high frequencies
used at most proton LINACs and electron accelerators
Capacitive linear-cut BPM for low frequencies
Electronics for position evaluation
BPMs for measurement of closed orbit, tune and further lattice functions
Summary

## Slide 29
2-dim Model for a Button BPM
‘Proximity effect’: larger signal for closer plate
Ideal 2-dim model: Cylindrical pipe  image current density
via ‘image charge method’ for ‘pencil’ beam:
Image current: Integration of finite electrode size:
,  = 00


a
r
button
 =0

## Slide 30: 2-dim Model for a Button BPM
Ideal 2-dim model: Non-linear behavior and hor-vert coupling:
Sensitivity S is converts signal to position
with S [%/mm] or [dB/mm]
i.e. S is the derivative of the curve , here Sx = Sx(x, y) i.e. non-linear.
For this example: center part S=7.4%/mm  k=1/S=14mm
button
 real position
• measured posi.
‘Position Map’:
Horizontal plane:


a
r
button


## Slide 31
Button BPM Realization
LINACs, e--synchrotrons: 100 MHz < frf < 3 GHz  bunch length  BPM length
 50  signal path to prevent reflections
Button BPM with 50  
Example: LHC-type inside cryostat:
24 mm, half aperture a=25 mm, C=8 pF
 fcut=400 MHz, Zt = 1.3  above fcut
Ø24 mm
From C. Boccard (CERN)
50mm

## Slide 32
Button BPM at Synchrotron Light Sources
Example: Booster of ALS, Berkeley
The button BPM can be rotated by 450
to avoid exposure by synchrotron light:
Frequently used at boosters for light sources

## Slide 33
Button BPM at Synchrotron Light Sources
Due to synchrotron radiation, the button insulation might be destroyed
buttons only in vertical plane possible  increased non-linearity
PEP-realization
HERA-e
realization

## Slide 34
real beam position x [mm]
y = 20 mm
y = 10 mm
y = 0 mm
20
Sx(center) = 7.8 %/mm
for |x|<5mm & y=0mm
Sy(center) = 7.2 %/mm
for |y|< 5mm & x=0mm
button 2
position x [mm]
30
button 1
button 3
button 4
30
position x [mm]
button 1
button 2
button 4
button 3
28 mm
72 mm
 7 mm
button 1
button 2
button 3
button 4
18 mm
y = 20 mm
y = 10 mm
y = 0 mm
Simulations for Button BPM at Synchrotron Light Sources
Example: Simulation for ALBA light source for 72 x 28 mm2 chamber
Optimization: horizontal distance and size of buttons
Result: non-linearity and xy-coupling occur in dependence of button size and position
1 mm
steps
1 mm steps
y
x
from A.A. Nosych et al., IBIC’14

## Slide 35
Poll concerning Button BPMs and Signal Shape
Poll 5.7:
What is correct? For button BPMs the position sensitivity S(x) is …
independent with respect to the beam-pipe size
linear with respect to x (independent on the button geometry)
non-linear with respect to x (it depends on the button geometry)
The horizontal S(x) is independent on the vertical beam offset y
Poll 5.8:
Assume a small button BPM with a cut-off frequency of fcut = 3 GHz and a beam bunch of 1 ns length. The signal is proportional …
to the bunch current
to the derivative of the bunch current
to the button diameter
www.menti.com
code: 1208 6026
Notes: 35

## Slide 36
Outline:
Signal generation  transfer impedance
Capacitive button BPM for high frequencies
used at most proton LINACs and electron accelerators
Capacitive linear-cut or ‘shoe-box’ BPM for low frequencies
used at most proton synchrotrons due to linear position reading
Electronics for position evaluation
BPMs for measurement of closed orbit, tune and further lattice functions
Summary

## Slide 37
Linea-cut or ‘Shoe-box’ BPM for Proton Synchrotrons
Frequency range: 1 MHz < frf < 10 MHz  bunch-length >> BPM length.
Signal is proportional to actual plate length:
In ideal case: linear reading
Linear-cut BPM:
Advantage: Very linear, low frequency dependence
i.e. position sensitivity S is constant
Disadvantage: Large size, complex mechanics
high capacitance
Size: 200x70 mm2

## Slide 38
Technical Realization of a linear-cut BPM
Quadrupole
Linear Ampl.
Beam
70mm
180mm
Linear Ampl.
Horizontal
electrode
Right
channel
Technical realization at HIT synchrotron of 46 m length for 7 MeV/u 440 MeV/u
BPM clearance: 180x70 mm2, standard beam pipe diameter: 200 mm.

## Slide 39
70mm
180mm
Linear Ampl.
Horizontal
electrode
Right
channel
Technical Realization of a linear-cut BPM
Technical realization at HIT synchrotron of 46 m length for 7 MeV/u 440 MeV/u
BPM clearance: 180x70 mm2, standard beam pipe diameter: 200 mm.

## Slide 40
Comparison Linear-cut and Button BPM
 | Linear-cut BPM | Button BPM
Precaution | Bunches longer than BPM | Bunch length comparable to BPM
BPM length (typical) | 10 to 20 cm length per plane | 1 to 5 cm per button
Shape | Rectangular or cut cylinder | Orthogonal or planar orientation
Bandwidth (typical) | 0.1 to 100 MHz | 100 MHz to 5 GHz
Coupling | 1 M or 1 k (transformer) | 50 
Cutoff frequency (typical) | 0.01… 10 MHz (C=30…100pF) | 0.3… 1 GHz (C=2…10pF)
Linearity | Very good, no x-y coupling | Non-linear, x-y coupling
Sensitivity | Good, care: plate cross talk | Good, care: signal matching
Usage | At proton synchrotrons, frf < 10 MHz | All electron acc., proton Linacs, frf > 100 MHz

## Slide 41
Poll concerning linear-cut BPMs
Poll 5.9:
Linear-cut BPMs …
are used at proton synchrotrons with an acceleration frequency facc 10 MHz
can distinguish between counter-propagating beams within a collider
provide a linear position sensitivity over a wide range of beam offsets
www.menti.com
code: 1208 6026
Notes: 41

## Slide 42
Outline:
Signal generation  transfer impedance
Capacitive button BPM for high frequencies
used at most proton LINACs and electron accelerators
Capacitive linear-cut BPM for low frequencies
used at most proton synchrotrons due to linear position reading
Electronics for position evaluation
analog signal conditioning to achieve small signal processing
BPMs for measurement of closed orbit, tune and further lattice functions
Summary

## Slide 43
Excurse: General Noise Sources
Any electronics is accompanied with noise due to:
Thermal noise as given by the thermal movement of
electrons described by Maxwell-Boltzmann distribution
Within resistive matter
average movement cancels
but standard deviation remains:
this is white noise i.e. no frequency dependence
kB Boltzmann constant, T temperature, R resistivity, f bandwidth
Maxwell-Boltzmann distribution:
Example: Maxwell-Boltzmann distribution of a free electron gas
Shot noise as given by the fluctuations of finite amount of electrons
for most electronics not important due to large amount of electrons
Flicker noise or ‘1/f - noise’ as given by trapping of electrons in matter
pink noise due to a frequency , ‘corner-frequency’ fc  3 kHz i.e. 1/f-noise = thermal noise
Disturbance by Electro-Magnetic Interference EMI:
Not noise, but pick-up of electro-magnetic waves from the environment
leads unwanted signal deformation and depends on the surrounding
R
at T

## Slide 44
signal only
S/N=2
signal only
S/N=2
noise only
Excurse: Signal to Noise Consideration
The minimum noise is given thermal noise:
R is the input Ohmic resistor  input of a low-noise amplifier (realistic noise figure 2-3 x Ueff)
f is the bandwidth of the signal chain
Simulation:
Broadband ‘white’ noise
For Signal-to-Noise = 2
(Noise: std. deviation, signal: maximum)
 peak just visible
For Signal-to-Noise = 4
 peak well visible
Filtering improves S/N
here: n=50 taps moving average
 bandwidth restriction
sample
signal only
S/N=4
filtered
signal only
S/N=4
sample
Ueff

## Slide 45
General: Noise Consideration
1. Signal voltage given by:
2. Position information from voltage difference:
3. Thermal noise voltage given by:
Signal-to-noise ΔUim/Ueff is influenced by:
Input signal amplitude
 large or matched Zt
Thermal noise at R=50  for T=300 K
Bandwidth Δf
 Restriction of frequency width
because the power is concentrated
on the harmonics of frf
Example: GSI-LINAC with frf = 36 MHz
Remark: Additional contribution by non-perfect electronics typically a factor 2
Moreover, pick-up by electro-magnetic interference can contribute  good shielding required

## Slide 46
However: not only noise contributes but additionally beam movement by betatron oscillation
 broadband processing i.e. turn-by-turn readout for tune determination.
Comparison: Filtered Signal  Single Turn
Example: GSI Synchr.: U73+, Einj=11.5 MeV/u 250 MeV/u within 0.5 s, 109 ions
Position resolution < 30 μm (BPM half aperture a=90 mm)
average over 1000 turns corresponding to 1 ms
or 1 kHz bandwidth
Turn-by-turn data have much larger variation
1000 turn average for closed orbit
Variation < 10 μm (sufficient for application)
Single turn e.g. for tune
Variation  150 μm

## Slide 47
Broadband Signal Processing
Hybrid or transformer close to beam pipe for analog ΔU & ΣU generation or Uleft & Uright
Attenuator/amplifier
Filter to get the wanted harmonics and to suppress stray signals
ADC: digitalization  followed by calculation of of ΔU /ΣU
.
Advantage: Bunch-by-bunch possible, versatile post-processing possible
Disadvantage: Resolution down to  100 µm for linear-cut, i.e. 0.1% of aperture,
resolution is worse than narrowband processing

## Slide 48
Narrowband Processing for improved Signal-to-Noise
Narrowband processing equals heterodyne receiver (e.g. AM-radio or spectrum analyzer)
Attenuator/amplifier
Mixing with accelerating frequency frf  signal with sum and difference frequency
Bandpass filter of the mixed signal (e.g at 10.7 MHz)
Rectifier: synchronous detector
ADC: digitalization → followed calculation of ΔU/ΣU
Advantage: Spatial resolution about 100 time better than broadband processing
Disadvantage: No turn-by-turn diagnosis, due to mixing = ’long averaging time’
For non-relativistic p-synchrotron:  variable frf leads via mixing to constant intermediate freq.

## Slide 49
Excurse: Mixer and Synchronous Detector
Mixer: A passive rf device with
Input RF (radio frequency): Signal of investigation
Input LO (local oscillator): Fixed frequency
Output IF (intermediate frequency)
 Multiplication of both input signals, containing the sum and difference frequency.
Synchronous detector: A phase sensitive rectifier
RF
IF
LO
RF
IF
LO
limiting ampl.
input
Urf(t)=U0 sin t
low pass filter
output Udc U0
U0 sin2 t
+ Umax if sin t > 0
- Umax if sin t < 0

## Slide 50
Analog versus Digital Signal Processing
Digital receiver as modern successor of super heterodyne receiver
Basic functionality is preserved but implementation is very different
Digital transition just after the amplifier & filter or mixing unit
Signal conditioning (filter, decimation, averaging) on digital electronics e.g. FPGA
Advantage of DSP: Versatile operation, flexible adoption without hardware modification
Disadvantage of DSP: non, good engineering skill requires for development, expensive
Analog frequency
translator
Analog demodulator
and filter
ADC
BPM analog
signal
digital
output
Analog frequency
translator
ADC
Digital
filter
Digital
Signal Proc.
digital
output
BPM analog
signal
Digital processing (triggered by telecommunication development)
Traditional analog processing
Analog
Digital
Modern instrumentation uses digital techniques with extended functionality.

## Slide 51
Comparison of BPM Readout Electronics (simplified)
Type | Usage | Precaution | Advantage | Disadvantage
Broadband | p-sychr. | Long bunches | Bunch structure signal Post-processing possible Required for fast feedback Required for transfer lines with few bunches | Resolution limited by noise
Narrowband | all synchr. | Stable beams >100 rf-periods | High resolution | No turn-by-turn Complex electronics
Digital Signal Processing | all | Several bunches ADC 125 MS/s | Very flexible High resolution Trendsetting technology for future demands | Basically non! Limited time resolution by ADC  under-sampling Man-power intensive

## Slide 52
Poll concerning BPM Electronics
Poll 5.12:
Assume turn-by-turn position evaluation within a synchrotron. What type of electronics is appropriate?
Broadband signal processing
Narrowband signal processing
Digital signal processing as the only possible solution
Poll 5.11:
Within a transfer line between two synchtrotrons, the positon of a single bunch should be measured. What type of electronics is appropriate?
Broadband signal processing
Narrowband signal processing using the synchrotron bunching
The position of a single bunch can not be measured
Poll 5.10:
Which type of electronics delivers a better position resolution?
Broadband signal processing
Narrowband signal processing
www.menti.com
code: 1208 6026
Notes: 52

## Slide 53
Outline:
Signal generation  transfer impedance
Capacitive button BPM for high frequencies
used at most proton LINACs and electron accelerators
Capacitive linear-cut BPM for low frequencies
used at most proton synchrotrons due to linear position reading
Electronics for position evaluation
analog signal conditioning to achieve small signal processing
BPMs for measurement of closed orbit, tune and further lattice functions
frequent application of BPMs
Summary

## Slide 54
Trajectory Measurement with BPMs
Trajectory:
The position delivered by an individual bunch within a transfer line or a synchrotron.
Main task: Control of matching (center and angle), first-turn diagnostics
Example: LHC injection 10/09/08 i.e. first day of operation !
From R. Jones (CERN)
-10
-10
horizontal
vertical
Monitor number (530 BPMs on 27 km)
10
-10
0
x [mm]
10
-10
0
y [mm]
Tune values at LHC: Qh = 64.3, Qv = 59.3

## Slide 55
Close Orbit Measurement with BPMs
Single bunch position averaged over 1000 bunches  closed orbit with ms time steps.
It differs from ideal orbit by misalignments of the beam or components.
Example: GSI-synchrotron at two BPM locations, 1000 turn average during acceleration:
Closed orbit:
Beam position
averaged over many turns
(i.e. betatron oscillations).
The result is the basic tool
for alignment & stabilization
horizontal position at 2 BPMs
vertical position at 2 BPMs
time t = 0 ... 600 ms
0
5
-5
-10
-5
0
2
4
-2
time t = 0 ... 600 ms
0
6
-6
-12
position [mm]
position [mm]
position [mm]
along synchrotron
time t = 0 ... 600 ms
signal strength
horizontal
vertical

## Slide 56
Closed Orbit Feedback: Typical Noise Sources
From M. Böge, PSI, N. Hubert, Soleil
0.01
1
0.1
100
10
1000
frequency [Hz]
Power spectral density [mm2/Hz]
Beam movement:
Short term (min to 10 ms):
Traffic
Machine (crane) movements
Water & vacuum pumps
50 Hz main power net
Medium term (day to min):
Movement of chambers
due to heating by radiation
Day-night variation
tide, moon cycle
Lang term ( > days):
Ground settlement
Seasons, temperature variation

## Slide 57
Close Orbit Feedback: BPMs and magnetic Corrector Hardware
Orbit feedback: Synchrotron light source  spatial stability of light beam
Example from SLS-Synchrotron at Villigen, Swiss:
Procedure:
1. Position from all BPMs
2. Calculation of corrector setting via
Orbit Response Matrix
3. Digital feedback loop
 regulation time down to 10 ms
 Role od thumb:  4 BPMs per betatron wavelength
From M. Böge, PSI
Position from all BPMs
Calculation of corrector strength
Setting of correctors
feedback
Acc. optics
Uncorrected orbit: typ. x 2rms  1 mm
Corrected orbit: x 2rms  1 m up to 100 Hz bandwidth!

## Slide 58
Close Orbit Feedback: Results
Orbit feedback:
Example: 12 beam positions at GSI-SIS during ramping from 8.6 to 500 MeV/u for Ar18+
Procedure:
1. Position from all 12 BPMs
2. Calculation of corrector setting on fast (FPGA-based) electronics
3. Submission to corrector magnets
4. New position measurement
 regulation time down to 10 ms
feedback off
time [ms]
400
300
100
0
position [mm]
20
10
0
-10
200
feedback on
time [ms]
0
100
300
400
position [mm]
20
10
0
-10
200
Rule of thumb:
Movement related to tune i.e. ‘natural oscillations by periodic focusing’
To determine the ‘sine-like’ oscillation 4 BPMs per oscillation are required
 4 BPMs per tune value (but detailed investigation required to determine the # of BPMs)
position x(t) at 12 BPMs

## Slide 59
Graphics by R. Singh, GSI
Tune Measurement: General Considerations
Beam particle’s in-coherent motion  center-of-mass stays constant
Coherent excitations are required for the detection by a BPM
Excitation of all particles by rf  coherent motion
 center-of-mass variation turn-by-turn
Notes: Beam with several particles with randomly distributed phases and amplitude will give only very small fluctuations due to finite number of particles (schottky fluctuations) proportional to sqrt(N)
Beam excitation aligns the particle phases, and allows to measure center-of-mass before the beam decohores due to finite spread in frequency. Thus regular excitation needed. Types of excitation were studied as part of this work, Kick excitation (a wide band, or single time impulse excitation), requires high power, and significantly dilutes the phase space.
Chirp excitation, slow method, difficult to use in day to day operations, BTF excitation, to study the amplitude and phase response of the beam.
A compromise of the both the above methods, band limited (or pseudo random noise) for regular excitation without significant blow-up, usable along the ramp.

## Slide 60
Tune Measurement: General Considerations
The tune Q is the number of betatron oscillations per turn.
The betatron frequency is fβ = Q f0.
Measurement: excitation of coherent betatron oscillations + position from one BPM.
Example: Tune measurement for six turns with the three lowest frequency fits:
To distinguish
for q < 0.5 or q > 0.5:
Changing the tune slightly,
the direction of q shift differs.
From a measurement one gets only the non-integer part q of Q with Q = n  q.
Moreover, only 0 < q < 0.5 is the unique result.

## Slide 61
Tune Measurement: The Kick-Method in Time Domain
The beam is excited to
coherent betatron oscillation
→ the beam position measured
each revolution (’turn-by-turn’)
→ Fourier Trans. gives the non-integer tune q.
Short kick compared to revolution.
N = 200 turn  q > 0.003 as resolution
(tune spreads are typically q  0.001!)
The de-coherence time limits the resolution:
N non-zero samples
 General limit of discrete FFT:

## Slide 62
Tune Measurement: De-Coherence Time
The particles are excited to betatron oscillations, but due to the
spread in the betatron frequency, they getting out of phase (’Landau damping’):
Scheme of the individual trajectories of four particles after a kick (top)
and the resulting coherent signal as measured by a pick-up (bottom).
 Kick excitation leads to limited resolution
Remark: The tune spread is much lower for a real machine.

## Slide 63
Tune Measurement: Beam Transfer Function in Frequency Domain
Instead of one kick, the beam can be excited by a sweep of a sine wave, called ‘chirp’
 Beam Transfer Function (BTF) Measurement
as the velocity response to a kick
Principle:
Beam acts like a driven oscillator!
Using a network analyzer:
RF OUT is feed to the beam by a
kicker (reversed powered as a BPM)
The position is measured at
one BPM
Network analyzer: amplitude and
phase of the response
Sweep time up to seconds
due to de-coherence time per band
resolution in tune: up to 10−4
courtesy: F.Sannibale (LBNL)

## Slide 64
Tune Measurement: Result for BTF Measurement
BTF measurement at the GSI synchrotron, recorded at the 25th harmonics.
A wide scan with both sidebands at
h=25th-harmonics:
From the position of the sidebands q = 0.306 is determined.
Advantage: High resolution for tune and tune spread (also for de-bunched beams)
Disadvantage: Long sweep time (up to several seconds).
A detailed scan for the lower sideband
 beam acts like a driven oscillator:

## Slide 65
Tune Measurement: Gentle Excitation with Wideband Noise
Instead of a sine wave, noise with adequate bandwidth can be applied
 beam picks out its resonance frequency:
broadband excitation with white noise
of  10 kHz bandwidth
turn-by-turn position measurement by fast ADC
Fourier transformation of the recorded data
 Continues monitoring with low disturbance
Example: Vertical tune within 4096 turn
duration ≃ 15 ms
at GSI synchrotron 11  300 MeV/u in 0.7 s
Advantage:
Fast scan with good time resolution
Disadvantage: Lower precision
vertical tune versus time
Qy  0.02
time [ms]
vert. fractional tune qy
750
600
400
0
200
0.2
0.25
0.3
excitation
noise band
Q  0.05
vertical tune at fixed time  15ms
vertical fractional tune qy
0.2
0.25
0.35
0.3
amplitude [a.u.]

## Slide 66
Excurse: Example of Lattice Functions
The position of dipoles and quadrupoles
give the linear lattice functions
at injection point D = 0 is favored
chromatic correction with sextupoles,
Length C [m] |  | 1086
Energy Ekin [GeV/u] |  | 0.2  2
Tune Qh/v | h/v | 18.84 / 18.73
Max. dispersion ¦D¦ [m] |  | 1.73
Max. –function h/v [m] | h/v | 19.6 / 19.6
Natural chromaticity ξh/v | h/v | -1.19 / -1.20
Injected emittance εh/v [mm mrad] | h/v | 35 / 15
Injected mom. spread Δp/p0 [%] |  | 0.05
Definition of dispersion D(s):
xD(s) = D(s)  p/p0
Definition of chromaticity  per turn:
Q/Q0 =   p/p0
Example: GSI SIS100 ion synchrotron
x: 80 mm
y: 40 mm
1/6 of SIS100: Length 181 m
y

## Slide 67
-Function Measurement from Bunch-by-Bunch BPM Data
Excitation of coherent betatron oscillations: From the position deviation
xik at the BPM i and turn k the -function (si) can be evaluated.
The position reading is:
 a turn-by-turn position reading at many location (4 per unit of tune) is required.
The ratio of -functions at different location:
The phase advance is:
Without absolute calibration,
-function is more precise:

## Slide 68
From J. Wenninger: CAS on BD, CERN-2009-005 & J. Wenninger CERN-AB-2004-009
Dispersion Measurement from Closed Orbit Data
Dispersion D(si ): Change of momentum p by detuned rf-cavity
 Position reading at one location :
 Result from plot of xi as a function of Δp/p  slope is local dispersion D(si )
Example: Dispersion measurement D(s) at BPMs at CERN SPS
Theory-experiment correspondence
after correction of
BPM calibration
quadrupole calibration
: model expectation
green histogram: measurement
6.9 km

## Slide 69
Chromaticity Measurement
Plot of ΔQ/Q as a function of Δp/p
 slope is chromaticity ξ.
Measurement at LEP
Chromaticity ξ: Excitation of coherent betatron oscillations and
change of momentum p by detuned rf-cavity:
→ Tune measurement
(kick-method, BTF, noise excitation):

## Slide 70
Poll concerning Tune and Lattice Function Measurement
Poll 5.13:
A tune measurement is related to excitation of …
coherent transverse oscillations
incoherent transverse oscillations
coherent longitudinal oscillations
incoherent longitudinal oscillations
Poll 5.14:
The dispersion function along a transfer line is measured using a beam with momentum offset and measuring …
the arrival time at several locations
the beam position at several locations
the beam profile at several locations
nothing, as dispersion is only a theoretical quantity
Poll 5.15:
What is a closed orbit feedback system at a synchrotron?
The alignment of the trajectory of the injected beam during its first turn
The correction of individual bunches during storage
The correction of the beam position averaged over more than 1000 turns
The mechanical movement of the synchrotron magnets
www.menti.com
code: 1208 6026
Notes: 70

## Slide 71
MEVVA
MUCIS
PIG
RFQ IH1 IH2
Alvarez DTL
HLI: (ECR,RFQ,IH)
Transfer to
Synchrotron
2.2 keV/u
β = 0.0022
120 keV/u
β = 0.016
11.4 MeV/u
β = 0.16
Gas Stripper
U4+
U28+
All ions, high current, 5 ms@50 Hz, 36&108 MHz
Foil Stripper
To SIS ↑
Constructed in the 70th, Upgrade 1999,
further upgrades in preparation
BPM: Capacitive type, for position and time-of-flight
total 25 device
Appendix GSI Ion LINAC: Position and mean beam energy Measure
1.4 MeV/u ⇔β = 0.054
50 m

## Slide 72
Ion (Z) | 1 → 92 (p to U)
Circumference | 216 m
Inj. type | Multi-turn
Injection energy | 11 MeV/u
Max. final energy |  2 GeV/u
Ramp duration | 0.1 → 1.5 s
Acc. RF | 0.8 → 5 MHz
Harmonic | 4 (= # bunches)
Bunching factor | 0.4 → 0.08
Tune Qx & Qy | 4.3 & 3.3
Important parameters of SIS-18
injec-
tion
acceleration
extrac-
tion
BPM: for tune measurement
1 device
BPM: position 1 s to s
12 devices
Exciter: for tune excitation
Schottky pick-up:
p/p, Q,  ....
1 device
Appendix GSI Ion Synchrotron: Position, Tune etc. Measurement

## Slide 73
Meas. Stripline  SMES
Exc. Stripline  SEXC
Button BPMs  BPM
Abbreviation:
LTB:
3 BPM
BTS:
4 BPM
SEXC
STORAGE RING:
7 or 8 BPM/cell
(120 BPMs in total)
1 SEXC
BOOSTER:
11 BPM/quadrant
(44 BPMs in total)
1 SMES
1 SEXC
SEXC
SMES
Beam position:
Center of mass
Many locations!
Frequent operating tool
For position stabilization
i.e. closed orbit feedback
Tune: Qx = 18.18
Qy = 8.37
From U. Iriso, ALBA
Appendix: Synchrotron Light F. ALBA: Position, tune etc. Measure

## Slide 74
Summary Pick-Ups for bunched Beams
The electric field is monitored for bunched beams using rf-technologies
(’frequency domain’). Beside transformers they are the most often used instruments!
Differentiated or proportional signal: rf-bandwidth ↔ beam parameters
Proton synchrotron: 1 to 100 MHz, mostly 1 M  proportional shape
LINAC, e−-synchrotron: 0.1 to 3 GHz, 50   differentiated shape
Important quantity: transfer impedance Zt(ω, ).
Types of capacitive pick-ups:
Linear-cut (p-synch.), button (p-LINAC, e−-LINAC and synch.)
Remark: Stripline BPM as traveling wave devices are frequently used
Position reading: difference signal of four pick-up plates (BPM):
Non-intercepting reading of center-of-mass → online measurement and control
slow reading  closed orbit, fast bunch-by-bunch trajectory
 Excitation of coherent betatron oscillations and response measurement
excitation by short kick, white noise or sine-wave (BTF)
 tune q, chromaticity ξ, dispersion D etc.

## Slide 75
Backup slides

## Slide 76
Quiz related to BPM Signals
What is the cause of the wall current and wall current density?
What is its connection to the beam bunches?
What is the usage of the transfer impedance?
What is the frequency dependence of the transfer impedance and why is it so?
What is baseline shift?
Transfer impedance: Do you know other quantities of similar type?
How you determine the beam position?
Why it is clever to normalize the difference to the sum signal for position calculation?
What are the typical acceleration frequencies at electron and proton
LINACs, synchrotrons and cyclotrons?
 Back to BPM types
Notes: 76

## Slide 77
Quiz related to BPM Types
Assuming a Gaussian bunch of t =1 ns long bunch.
What frequency bandwidth is appropriate positon determination?
For what beam condition or accelerator type button BPMs can be arranged
horizontally & vertically? For what parameter it can’t be done?
Why is it an advantage to have a horizontally & vertically orientation of buttons?
Why is a linear positon behavior an advantage?
What can you do if the position map is non-linear? What are precautions?
 Back to position evaluaton
Notes: 77

## Slide 78
Principle: Network Analyzer provide an active frequency scan
 frequency response in amplitude and phase for a Device Under Test DUT
Scattering parameters Sij :
For same ZL at all ports & a2=0 (no output reflection):
Gain: Uthru / Uscan = b2 /a1 = S21
Input reflection: Urefl / Uscan = b1 /a1 = S11
concerning amplitude and phase (after calibration)
ZLUscan=a1
b2=ZL Uthru
ZLUrefl=b1
a2
S11
S22
S12
S21
DUT
Port 1
Port 2
Uscan(t): active frequency scan
DUT
Uscan
Urefl
Uthru
directional coupler
NWA
Port 1
Port 2
line impedance ZL
line impedance ZL
Excurse: Network Analyzer NWA

## Slide 79
Tune Measurement: Result for BTF Measurement
BTF measurement at the GSI synchrotron, recorded at the 25th harmonics.
A wide scan with both sidebands at
h=25th-harmonics:
From the position of the sidebands q = 0.306 is determined. From the width
Δf/f  5·10-4 the tune spread can be calculated via
Advantage: High resolution for tune and tune spread (also for de-bunched beams)
Disadvantage: Long sweep time (up to several seconds).
A detailed scan for the lower sideband
 beam acts like a driven oscillator:

## Slide 80
Phase Measurement from Bunch-by-Bunch BPM Data
Excitation of coherent betatron oscillations: From the position deviation
xik at the BPM i and turn k the betatron phase is measured.
Result:
Model does not describes the reality completely, corrections required
At interaction point IP (detector location) an additional phase shift is originated
Alignment by correction dipoles (steerer), quadrupoles or sextupoles.
Example: Phase advance (s) compared to the expected 0(s) by optics calculation e.g. MADX
at each BPM at CERN’s at LEP (e+ - e- collider with 27 km, previously in LHC tunnel)
From J. Borer et al, EPAC’92

## Slide 81
Phase Measurement from Bunch-by-Bunch BPM Data
Excitation of coherent betatron oscillations: From the position deviation
xik at the BPM i and turn k the beta-function can be determined.
Result:
Model does not describes
the reality completely
Corrections executed
Increase of the luminosity
Remark:
Measurement accuracy depends on
BPM accuracy
Numerical evaluation method
Example: Measured (s) compared to the expected 0(s) and normalized
for each BPM at BNL for RHIC (proton-proton or ions circular collider with 3.8 km length)
From X. Shen et al.,
Phys. Rev. Acc. Beams 16, 111001 (2013)
See e.g.:
R. Tomas et al., Phys. Rev. Acc. Beams 20, 054801 (2017)
A. Wegscheider et al., Phys. Rev. Acc. Beams 20, 111002 (2017)

## Slide 82
Closed Orbit Feedback: Results
Uncorrected orbit:
Beam offset and
oscillation
here x 2rms = 2.3 mm
Corrected orbit:
Beam
dynamically corrected
here x 2rms = 1 m !
path along synchrotron [betatron phase]
hor. offset x [mm]
5
-5
0
path along synchrotron [betatron-phase]
hor. offset [mm]
-0.2
-0.4
0.4
0
0.2
From M. Böge, PSI
Orbit feedback: Compensating variations of different kind, goal: Δx  0.1 ∙x
Synchrotron light source  spatial stability of light beam
Example from SLS-Synchrotron at Villigen, Swiss: