## Slide 1: Beam Gas Curtain – An Overview of 2024 on the LHC
Oliver Stringer
On behalf of the BGC collaboration
H. Zhang, M. Ady, D. Butti, A. R. Churchman, P. Forck, T. Lefevre, S. Mazzoni, C. Pasquino, A. Rossi, A. Salehilashkajani, M. Sameed, G. Schneider, O. Sedlacek, C. C. Sequeiro, S. Sethi, K. Sidorowski, S. Udrea, R. Veness, A. Webber-Date, and C. P. Welsch

## Slide 2: Outline
Motivation of work.
Beam Gas Curtain (BGC) working principles.
Accuracy validation against ATLAS and CMS.
2D proton beam profile and emittance measurements.
Highlights as a lead beam diagnostic.
Summary and outlook.

## Slide 3: Motivation
It is essential to be able to measure beam size within accelerators to determine emittance, and luminosity.
Current transverse profile diagnostics for the LHC include wire-scanners and the BSRT, both of which have limitations.
The High-Luminosity upgrade will introduce new challenges for existing diagnostics.
We present a non-invasive beam profile monitor that works for all stages of acceleration on the LHC.
  Only instrument able to acquire lead ion beam size measurements at injection energies.
  Only instrument able to acquire proton and ion beam size measurement through the acceleration cycle.

## Slide 4: Gas Curtain Generation
Accelerate gas to supersonic speeds.
Collimate the profile using skimmers to a 45 degrees thin curtain.
Version 1
Version 2
Version 3
18 years of research on gas jet technology.
Credit: Bethany Spear

## Slide 5: LHC Installation
Optical system
Gas curtain dump
Gas curtain injection
Notes: Jan 2023
Thanks - the support provided by the BI group, vacuum group and the operation team, technicians at CERN.
3.5x10^16 mol . m^-3 - simulation
2.6x10^16 mol . m^-3 - measurement

## Slide 6: Accuracy Validation
Noisy measurements (±10% peak-to-peak) but values consistent with emittance scans and reproducible over several fills
Visible neon transition and horizontal emittance
#9842
#9843
#9846
#9848
#9849
120 s integration per point
Courtesy of D.Butti
#9969
Consistent with both emittance scans
Proton Physics
Notes: (No space charge)
(No jet thickness)

## Slide 7: Visible or UV Fluorescence
Visible wavelength is a reduced signal, but a neutral emitter (independent of space charge).
wavelength [nm]
UV wavelength provides more signal but is a charged emitter.
VIS 120 s
±5% ~ 120s
beam
UV 40 s
±2.5% ~ 40s
beam
Visible resolution = ~65 µm
UV resolution = ~200 µm
6.8 TeV
6.8 TeV
Proton Physics
Courtesy of D.Butti
Notes: UV resolution is poor for stable beams. But good for Injection and Ramp (increased size and higher signal is needed).
Combination

## Slide 8: Combining UV and Visible Fluorescence
Combining UV and visible measurements allows accurate measurements every 40 seconds.
Fluctuations within ±5% in emittance.
#10084
>22h SB due to injectors unavailability
Absolute values still compatible with experiment emittance scans.
Proton Physics
Courtesy of D.Butti

## Slide 9: Combining UV and Visible Fluorescence
Combining UV and visible measurements allows accurate measurements every 40 seconds.
Fluctuations within ±5% in emittance.
#10084
>22h SB due to injectors unavailability
Absolute values still compatible with experiment emittance scans.
Proton Physics
Courtesy of D.Butti

## Slide 10: Highlights from 2024 Ion Run
BGC continuously running throughout ion run, close to operational instrument.
Online average emittance measurements.
Ion Physics
Courtesy of D.Butti
Notes: Larger cross-section and lower SR, even easier than Protons.

## Slide 11: Summary
BGC performance as emittance monitor thoroughly assessed in 2024.
Invited talks at IBIC23, IBIC24, and IPAC24.
Achievements with proton beams.
  Accurate and reproducible measurements for horizontal direction in stable beams.
  Combined UV-visible operation provides the best option for entire fill.
Achievements with ion beams.
  Stronger fluorescence and larger sizes simplifies BGC operation.
  Effective online emittance monitoring during ion physics production.

## Slide 12: Outlook
Excellent basis for healthcare applications (£340k JetDose award).
Install a EBTS dedicated BGC for potential active halo suppression methods post HL-LHC.
Install a second BGC device for Beam 2 during LS3.
Develop towards an operational instrument and use for machine optimisation.
Measure fluorescence cross-sections - new physics.
Investigating potential future with halo imaging and control.

## Slide 13: Thank you for your attention
Many thanks to the whole BGC collaboration for their fantastic support and work over the past 10 years.
This work was supported by the HL-LHC-UK phase II project funded by STFC under Grant Ref: ST/T001925/1 and the STFC Cockcroft Institute core grant No. ST/V001612/1.
Notes: https://www.youtube.com/watch?v=vBd-Loq6U2o&t=0s

## Slide 14: 2D Profile Measurements
Injection
Start of ramp
Injection
Start of Ramp
Stable Beams
Courtesy of D.Butti
Proton Physics

## Slide 15: BSRT Ramp Switch
BSRT Switching during Ramp (from Undulator to Dipole SR).
Offset from experiments.
collision
~30 points in ramp
Courtesy of D.Butti
Proton Physics

## Slide 16: Vertical Profile
#10127
deconvolution becomes necessary
collision
Stable Beams
V profile
Courtesy of D.Butti

## Slide 17: Performance with ion beams
Injection
Start of ramp
Stable beams
Protons VIS
Lead ions VIS
More favorable conditions with ion beams
stronger fluorescence signal
larger beam sizes
sufficient signal with visible configuration, no need to combine UV signal
Limited availability of BSRT with ions:
no light at injection
rarely calibrated at flat-top
there is already the case for extra diagnostics
Ion physics
Courtesy of D.Butti

## Slide 18: 2023 Proton & Ion Beams
Time: 1 min
Proton beam
Energy: 6.8 TeV
Intensity: 3.45e14
Time: 1 min
Lead ion beam
Energy: 6.8 TeV/Z
Intensity: 1.5e13
Time: 1 min
Lead ion beam
Energy: 450 GeV/Z
Intensity: 1.53e13
Proton beam
Energy: 450 GeV
Intensity: 2.2e14
Time: 1 min
σx =0.917mm σy =0.817 mm
σx =0.308 mm σy =0.432 mm
σx =1.653 mm σy =1.418 mm.
σx =0.543 mm σy =0.589 mm
y
x
PRELIMINARY
y
x
y
x
y
x
PRELIMINARY
PRELIMINARY
PRELIMINARY
Notes: 585nm filter +- 10mm
Blackened coating, ~10% reflection
Multi-layer sputtering coating 0.2% reflection
Proton injection beam: 20230524, subfolder(27), Sigma_x =0.9172, Sigma_y =0.8167.
Proton stable beam: 20230705, Sigma_x =0.3079, Sigma_y =0.4324.
Lead ion injection beam: 20231029, subfolder(33), Sigma_x =1.653, Sigma_y =1.418.
Lead ion stable beam: 20231011, subfolder(33), Sigma_x =, Sigma_y =0.563.

## Slide 19: Halo Monitor Concept
losses
BLM
jet
beam
photons
camera
jet
beam
BGC Profile Monitor
BGC Halo Monitor
Courtesy of D.Butti

## Slide 20: Pressure Profiles
Interaction chamber (VGP.4a.5L4.BGC) | Pressure [mbar]
Ne gas Jet OFF | 2.0e-10
Ne gas Jet ON | 4.00e-8
C. Castro Sequeiro, PRAB 27, 043201, 2024.
H. Zhang et. al, Vacuum 208, 111701, 2023.
Simulation
Measurement
Eq. Pressure: 1.08x10-6 mbar
20 mm
20 mm
Eq. Pressure: 1.45x10-6 mbar

## Slide 21: Injection Optimisation
Nozzle
1st Skimmer
2nd Skimmer
3rd Skimmer
IP
0.00
1.00 – 10.00 mm
6.58 – 48.49 mm
174.94 – 241.11 mm
294.94 mm
4th Skimmer
346.13 – 409.94 mm
d: 500 – 7,500 um
d: 30 – 10,000 um
a
b
a: 50 – 2,000 um
b: 2 – 60 mm
a: 50 – 10,000 um
b: 2 – 100 mm
d
d
a
b

## Slide 22: Gas Jet Optimisation Framework
Density / # m-3 | Curtain Length / mm | Curtain Width / mm | Variation in centre / % | BG Pressure / mbar
 | 31.14 | 0.25 | 9.61 | 
Density / # m-3 | Curtain Length / mm | Curtain Width / mm | Variation in centre / % | BG Pressure / mbar
 | 31.14 | 0.25 | 9.61 | 

## Slide 23: EBTS