## Slide 1: Compression BOX
Cable degradation under transverse stress
A. Brem, M. Duda, M. Daly, J. Van den Eijnden ,D. M. Araujo, C. Müller, H. Garcia, C. Hug, A. Milanese S. Sidorov, S. Otten, M. Dhallé, Z. Guo, A. Kario, A. Haziot, H. ten Kate and B. Auchmann
"This work was performed under the auspices of and with support from the Swiss Accelerator Research and Technology (CHART) program."

## Slide 2: Acknowledgment
Many thanks to the US Magnet Development Program and the Lawrence Berkeley National Laboratory for providing the Nb3Sn cable.

## Slide 3: Compression BOX setup at UTwente
Background Field: 7.5 T
Unpowered cable
Typical UTwente
Transverse Pressure Setup
Properties | Values
Material | 316 L
Overall dimensions L x W x H | 365 x 72 x 24 (mm)
Compressed cable length | 38 mm
Background field | 11 T
Expected Current | < 15 kA
Expected stress on cable | 200 MPa

## Slide 4: Last time
pure paraffin wax
filled paraffin wax
or epoxy
S. Otten et al., 2023, 10.1109/TASC.2023.3267051
tests with cable dimensions of approx. 1.5 mm x 10 mm with 11 T background field
CD1 cable

## Slide 5: Compression Box with new type of cable cross-check and validation experiments
Big Box experiment
- coil insert – 14 turns in and external field (10T)
P. Wanderer et al., 2007, 10.1109/TASC.2007.899984.
common-coil dipole magnet DCC017 at BNL
same cable in Compression BOX experiment to cross-check and validate the experiments
BigBox cable
D. M. Araujo et al., 2024, 10.1109/TASC.2024.3368995

## Slide 6: Latest Results
CD1 cable
filled paraffin wax
“standard”
Tm 57 °C
filled high melting point wax
Tm 105 °C
inconclusive results with the new cable
to be continued...

## Slide 7: Improvements on Compression BOX better strain monitoring
pusher
former
- average stress is known by a single strain gauge on the pusher block
- the current on the pushing solenoids is also known, correlelates with strain on pusher block
! friction, blockage or uneven loading is not measured/detected
better strain monitoring is required, ideally on both sides
pusher and former
courtesy of J. Van den Eijnden

## Slide 8: Improvements on Compression BOX better strain monitoring
Measure z-component of strain in pusher
Measure y-component of strain in hole
Possible cause: not all force transferred from pusher to the cable.
Idea: add holes and strain gauges to the former and pusher to measure force transferred
Correlation should show if all force is transferred

## Slide 9: Conclusion and Outlook
Compression BOX was adapted to a new type of cable
A filled high melting point wax was tested
- Compression BOX tests with cables which are going to be used in SMACC
- Adaptation of the compression BOX former to accommodate more strain gauges
- Continue the evaluation of filled high melting point wax as impregnation media

## Slide 10: Von-Mises Stress Contract Friction
CD1 cable 20 kA
11 T background field
Subscale cable
10 kA
6 T background field
Outlook - Box with Subscale cable
and also for compression BOX experiments

## Slide 11
Questions!