## Slide 1: Bunch Shape Monitor for Pedestrians
Jocelyn TAN
XEI section
BI Day, 13th December 2024

## Slide 2: Outline
A guided tour of the Linac4
Bunch Shape Monitor: how does it work ?
Diagnostics with a BSM
Conclusion

## Slide 3: Linac4 : 1st link in the LHC accelerator chain
source
Building 400
Linac2: 40 yo+, a bottleneck for LIU
2006: pre-studies, 2008: Approved, 2021: PSB injector
76 m long, normal-conducting Linac
180 m long transfer line to the Proton Synchrotron Booster (PSB)

## Slide 4: A machine for H- ions
Hydrogen gas bottle
H- ion source 45 keV

## Slide 5: Elements for transverse optics
 | Data
Magnets | PMQ 128 inside DTL tanks EMQ 24 intertanks drift tube Steerers 21 ‘’
Transverse plane | Beam size rms < 3 mm
DTL-CCDTL drift : 615 mm only !
1 dual plane steerer
1 EMQ
1 vacuum sector valve
4 Monitors: BPM, BLM, BCT, SEM
1 EMQ
1dual plane steerer

## Slide 6: The Lion’s share for RF cavities
RFQ + bunchers: 4 tanks
Ekin = 3 MeV
H- ion source 45 keV
CCDTL modules: 21 cells
Ekin = 102 MeV
PIMS modules: 14 cells
Ekin = 160 MeV
DTL modules: 111 cells
Ekin = 50 MeV
160 MeV, the beam velocity is at 50% of c
Total of 150 accelerating cavities
Bunching frequency of 352 MHz

## Slide 7: Close-up view of the Linac4 beam
7
T = 2.84 ns
A bunch train
Up to 600 s long
Bunches spaced every 2.84 ns (frf = 352 MHz)
Optimum duty cycle: 2 – 3 %
~ 163 ps
A non-relativistic point charge, moving in a metal pipe
Physical bunch length = 60 ps
Direct signal acq., with 10 samples:
System with F = 83 GHz

## Slide 8: Outline
A guided tour of the Linac4
Bunch Shape Monitor: how does it work ?
Diagnostics with a BSM
Conclusion

## Slide 9: Secondary electron emission
H- bunches
Bunching frequency fRF = 352 MHz
Time distribution
 = 100 m Tungsten wire
at -10kV
Secondary
electrons

## Slide 10: Collecting SE beam
H- bunches
Bunching frequency fRF = 352 MHz
Time distribution
Target
at -10kV
Secondary
electrons
Input collimator
Output
collimator
Electron multiplier
Secondary
electrons

## Slide 11: Steering SE beam
Electron multiplier
H- bunches
Bunching frequency fRF = 352 MHz
Time distribution
Target
at -10kV
Secondary
electrons
Output
collimator
Electrostatic plates *Steering scan
Input collimator

## Slide 12: Focusing SE
Electrostatic plates
*Steering scan
*Focusing
H- bunches
Bunching frequency fRF = 352 MHz
Time distribution
Input collimator
Target
at -10kV
Electron multiplier
Secondary
electrons
Output
collimator

## Slide 13: Phase scan
Electron multiplier
Linac4 bunch train [s]
Phase
Time
H- bunches
TRF = 2.84 ns  360
Secondary
electrons
Superpose 352 MHz signal
*phase scan
Time distribution
Target
at -10kV
Spectrometre effect
 = 1.8  bunch length = 56 ps !!
bunch length
Longitudinal bunch length
Linac4 bunch train [ms]
Phase []

## Slide 14: BSM aka Feschenko Monitor
Institute for Nuclear Research RAS (INR), Moscow
International experience on BSM manufacturing:
A. Feschenko: father of the BSM based on SE emission
Linac4 technical specifications (U. Raich)
INR built two BSMs, turn-key system
  NI aqn. chain & Labview SW
  Provide support for beam commissioning & training
  Fix some teething issues
Courtesy: S. Gavrilov, ARIES workshop, Jan. 2021
Electron multiplier
Wire actuator

## Slide 15: Outline
A guided tour of the Linac4
Bunch Shape Monitor: how does it work ?
Diagnostics with a BSM
Conclusion

## Slide 16: RF cavity issues
Lost RF settings on 1 CCDTL cell: Debunched beam
Wrong cavity phase: Distribution with tail
13th Dec 2024

## Slide 17: Pulse flatness
Linac4 pulse is chopped to feed 2 PSB rings
Beam loading effect: transition from drift tube to a cavity. Affects energy distribution of the pulse head
Feed-forward algorithm reduces the transients by a factor 5-6.
time [ns]
phase [deg of 352 MHz]
phase [deg of 352 MHz]
time [ns]

## Slide 18: BSM2: Longitudinal emittance reconstruction
BSM2: 64 m in the Linac transfer line
Data obtained by variation of the debuncher cavity amplitude
With inverse matrix processing, Long. Emittance is reconstructed

## Slide 19: BSM2: Energy Ramping
First successful step to commission longitudinal painting
  Energy ramping around 160 MeV
  Successfully performed with last 2 PIMS cavities ( modulation +/-1 MV)
  Required for high intensity beams in the PSB (e.g. Isolde)

## Slide 20: Conclusion
Principle: transfer the temporal bunch structure information in a local region of space
Versatile design: for a wide range of ion type, beam energy and intensities, RF freq.
Physics domains: Electrostatics, HV, RF, Beam dynamics, Solid State physics, Electronics
Phase resolution limit: 0.5 , i.e. ~4 ps (Equivalent BW of 125 GHz)
The Feschenko monitor has been an international success for Linac facilities
A must for longitudinal beam diagnostics

## Slide 21
Thank you
The BSM team
Jocelyn Tan
Mark McLean
James Storey
Stephane Bart Pedersen
Yu Wu (BE-OP)

## Slide 22: References
Energy and longitudinal emittance
  J. Tan, et al., Results from the CERN Linac4 longitudinal bunch shape monitor, Proc. of IBIC2018, paper WEPA19, Shanghai, China, 2018
  V. A. Dimov et al., Emittance reconstruction techniques in presence of space charge applied during the Linac4 beam commissioning, Proc. of HB2016, paper WEPM1Y0, Malmö, Sweden, 2016
  G. Bellodi et al. , Longitudinal beam profile measurements in Linac4 commissioning, Proc. of LINAC2014, paper MOPP025 Geneva, Switzerland, 2014
  A. Feschenko, Technique and instrumentation for bunch shape measurements, Proc. of RUPAC 2012, paper FRXOR01, Saint-Petersburg, Russia, 2012,
  S. Gavrilov, Bunch shape monitors for hadron LINAC commissioning, https://agenda.ciemat.es/event/1229/contributions/2278/, ARIES workshop, Jan. 2021
  J. Tan and W. Yu, The Linac4 Bunch Shape Monitor, Linac4 training, EDMS 2213651, Rev.2, 2020

## Slide 23: Past issues
Target interaction with beam: still operating with the genuine wire thanks to
  SW protection against overheating
  Operators awareness and training
2019: CV work nearby, spreading water on HV connectors
Electronics parts ageing: VMOD TTL boards, capacitors
Vacuum sector venting without N2 flux: reduction electron multiplier lifetime
2021: wire stuck in the beam due to frictions on the brass wheel gear

## Slide 24: Linac4 BSM tests in Moscow

## Slide 25: Focusing effect

## Slide 26: BSM vs BPM

## Slide 27: Temporary movable test bench for staged commissioning
Modular commissioning during machine installation : validation of machine parameters
Transverse : profiles, emittances by slit & grid, wire scanners, laser, halo, position
Longitudinal : transmission, energy, DE/E, bunch length
Low energy test bench at 3 and 12 MeV
High energy test bench at 50 and 107 MeV

## Slide 28: Beam instrumentation : operators’ eyes
Diagnostics for routine operation:
  Beam intensity & transmission efficiency : 16 Beam Current Transformers & 1 Faraday cup
  Losses and Machine Protection Monitors : 19 Beam Loss Monitors
  Beam position and Energy : 25 Beam Position Monitors
  Longitudinal emittance, energy spread : 2 Bunch Shape Monitors
  Transverse profile :
    SEM Grids: 16 hor. and 16 ver.
    Wire Scanners: 15
    Laser Emittance Meter: 2 stations around first horizontal bend (for long beam pulses)
    BTV Scintillating screen : 1 for the stripping foil test stand
Ensure safe and efficient during commissioning and operation
A wide variety of devices at convenient locations along the Linac and new TL defined in this document (2010)

## Slide 29: Introduction to new PSB injection region
BSW1 BSW2 BSW3 BSW4
H-
(=1 proton+2electrons)
H-
P+ Circulating beam
Stripping foil
(x% efficiency)
P+
Dump
H0
H-
To electronics
Need to measure H0 and H- charges !
On line monitoring of SF degradation
Protect dump in case of foil breakage
Courtesy Jean Tassan-Viol

## Slide 30: Linac4 Layout
Π-mode Structure
160 MeV
12 Modules
6 Klystrons
12 EMQ
12 steerers
23 m
Cell-Coupled Drift Tube Linac
100 MeV
7 Modules
6 Klystrons
7 EMQ
7 steerers
14 PMQ
25 m
Drift Tube Linac
50 MeV
3 Tanks
3 Klystrons
1 EMQ
2 steerers
114 PMQ
19 m
Pre-injector
3 MeV
Source(s) @ 45 kV
2 solenoids
RFQ (1 Klystron)
11 EMQ
3 Buncher Cavities
2 Chopper units
+ dump
LBE
B. Mikulec
LIU Workshop, 20-22 January 2020, https://indico.cern.ch/event/847980/

## Slide 31: Linac4 accelerating structures
DTL, 3 – 50 MeV
CCDTL, 50 – 100 MeV
PIMS, 100 – 160 MeV
7-cell cavities in p-mode (12 cavities)
  Prototype in construction
  Drift Tube Linac (3 tanks)
  Prototype built, under testing.
  Construction starts in 2009
Cell-Coupled Drift Tube Linac (7 modules)
  Modules of 3 DTL-type cavities (2 drift tubes), connected by coupling cells.
  Prototypes built and tested, construction starts in 2009
M. Vretenar, IFMIF commissioning workshop, Tokyo, Japan, 17.09.2009
https://edms.cern.ch/document/1028126/1