## Slide 1: INSIDE THE EXPERIMENTAL AREAS: WORK OF THE XEI SECTION
Wilfried DEVAUCHELLE SY-BI-XEI
13/12/2024

## Slide 2: OVERVIEW
Overview of the East & North Experimental Areas.
Our mission.
Overview of beam instrumentation in the Experimental Areas.
Some interventions.

## Slide 3: OUR MISSION
Provide diagnosis instruments and support for the secondary beamlines.
We are responsible for all instrumentation in the secondary beamlines.
Support :
Support the beam physicists and experiments during operation.
    Weekly experiment turnover.
Adapt instruments to the needs of the physics program.
Observables :
Profile / position.
Intensity / Trigger / Time Of Flight.
Beam composition.

## Slide 4: THE EXPERIMENTAL AREAS
East area Building 157
Intensity: 103 to 1010 part/second.
Slowly extracted: 400 ms -> 4,6 s.
2556 bunches per ring.
Primary beam
Secondary beam
Experimental areas with fixed target
Unbunched.
Higher intensities.

## Slide 5: WHERE ?
East area Building 157

## Slide 6: EAST AREA
Built in 60’s and renovated during LS2.
250 meters beamlines.
2 permanent EAs.
2 versatile EAs.
~50 Main Users.
6
T11
T10
T09
T08
13/12/2024
W. Devauchelle - BI Day 2024

## Slide 7: EAST AREA
DIVERSE ACTIVITIES
10 Profile monitors.
05 Intensity / trigger / Time Of , Flight monitors (T.O.F.).
04 Particle identification , , , monitors.
T11
T10
T09
T08
04 Users barrack.
01 BI barrack.
W. Devauchelle - BI Day 2024

## Slide 8: NORTH AREA
Built in the 70’s.
60,000 m2.
Three experimental halls, service buildings and underground tunnels.
Four fixed targets fired from SPS.
Six secondary beam lines (~7 km).
NA-CONS
Phase 1: 2019 - 2029
Phase 2: 2029 - 2035
~160 Experiments & 2000 users for 2023.
W. Devauchelle - BI Day 2024

## Slide 9: NORTH AREA
DIVERSE ACTIVITIES
114 Profile monitors.
049 Intensity / trigger / Time Of , Flight monitors.
021 Particle identification , , monitors.
Served by four secondary beam lines, running simultaneously
Houses two Neutrino Platforms : NP02, NP04 (ProtoDUNE)
029 User’s barrack.
009 BI barrack.
045 Test areas.
W. Devauchelle - BI Day 2024

## Slide 10: Overview of beam instrumentation in the Experimental Areas
PROFILE / POSITION MONITOR.
INTENSITY / TRIGGER.
FOR PARTICLES IDENTIFICATION.

## Slide 11
PROFILE / POSITION MONITOR
Finger scintillator scanner (FISC)
(1990’s)
beam
~56 units
Beam profile fiber monitor (XBPF)
(2018)
~20 units
W. Devauchelle - BI Day 2024

## Slide 12
PROFILE / POSITION MONITOR
Multi-wire proportional chamber (MWPC)
(1970’s)
~50 units
Delay wire chamber (DWC)
(1990’s)
~35 units
Gas electron multiplier (GEM)
(2010’s)
~6 units
W. Devauchelle - BI Day 2024

## Slide 13
PROFILE / POSITION MONITOR
FISC angle scan in the H6 beamline
Wire Chamber today
Very first beam profile in
East Area
W. Devauchelle - BI Day 2024

## Slide 14
INTENSITY / TRIGGER / TIME OF FLIGHT
Scintillator paddle (XSCI)
(1970’s)
~52 units
Ionisation chamber (XION)
(1960’s)
~5 units
Beam Trigger Fiber (XBTF)
(2018)
~5 units
W. Devauchelle - BI Day 2024

## Slide 15
INTENSITY / TRIGGER
Scintillator Data in CESAR GUI
High voltage Scan of photomultiplier
W. Devauchelle - BI Day 2024

## Slide 16
PARTICLE IDENTIFICATION
Electro-Magnetic Calorimeter (XEMC)
(1970’s)
~3 units
High of pulses
Pulses Count
Pulse High Analysis

## Slide 17
PARTICLE IDENTIFICATION
Cherenkov threshold counter (XCET)
(1970’s)
Cherenkov differential counter with achromatic ring focus (CEDAR)
(1970’s)
~16 units
~10 units
W. Devauchelle - BI Day 2024

## Slide 18
PARTICLE IDENTIFICATION
PRESSURE SCAN WITH CEDAR
CHERENKOV THRESHOLDS
(4)
W. Devauchelle - BI Day 2024

## Slide 19
PARTICLE IDENTIFICATION
Cherenkov threshold counter (XCET)
GAS CONTROL RENOVATION
~5 units
Older gas control panel
of East Area
Older gas control panel
of North Area
New gas control panel

## Slide 20: Examples of some interventions
CEDAR-Hydrogen beam test in 2022.
New request to fill a CEDAR with H2 gas at high pressure.
Critical for the NA-62 Experiment (ECN3).
Extremely challenging due to high risk of explosion.
Need to adapt 40-year old system.
SY-BI responsible for motors, diaphragm, PMTs, control, acquisition & commissioning.
Cross-collaboration with many CERN teams (HSE, gas systems, integration, beam physicists, controls, users…).
Successful experiment! – today is a key instrument in NA62.
Our team performing some tests in the CEDAR-H installed in the H6 beamline for testing.

## Slide 21: SOME INTERVENTIONS
HV lock down.
Disconnection.
Remove cable from the beam.
Update database.
FISC Removal
XCET PMT replacement
HV lock down.
Disconnection.
Dismount socket.
Replace PMT.
Remount socket.
Reconnection.
HV unlock & Scan.
Adjust Delay’s.

## Slide 22: SOME INTERVENTIONS
CEDAR INSTALLATION
Cabling Electrical & Gas
Mechanical installation (transport).
Electrical connections.
Gas connections (BE-EA-EC).
Setting.
Calibration.
PMT HV adjust.
Delay synchronization for trigger coincidence.
Movement limit.
Transport and placement
Delay’s scan

## Slide 23: SUMMARY
Secondary Beams from fixed Target, slowly extracted, not bunched, low intensity.
11 different type of instruments, 200+ devices – all necessary for operation!
XEI section: responsible of beam diagnostics for secondary beams.
40+ years old equipment in many cases – urgent consolidation needed.
To be maintained up until NA-CONS Phase 2.

## Slide 24
(1) From I. ORTEGA – SY-BI-XEI
(2) From D. BANERJEE – P Shutdown Lectures – BE-EA-LE
(3) From J. TAN – SY-BI-XEI
(4) From M. VAN DIJK – BE-EA-LE
AKNOWLEDGEMENT
(5) From B. HOLZER – EP-SME-62
(6) From J BUESA ORGAZ– BE-EA-EC

## Slide 25: Cherenkov threshold counter (XCET) Mirror refurbishment
W. Devauchelle - BI Day 2024

## Slide 26: TIME OF FLIGHT
(3)

## Slide 27: NORTH AREA BEAM CHARACTERISTICS
Proton beam characteristics | 
Primary beam momentum | 400 GeV/c
Spill intensity range | 2 x1012 – 4.2 x1013 ppp
Spill duration | typ 4.8 s (1 to 10 s)
Extracted intensity /year | ~1 x1019
Ion beam characteristics | 
Particle | Pb82+
Spill intensity range | >1 x107
Spill duration | <10 s
Secondary beams characteristics |  | 
EHN1 | 205 – 360 GeV/c | p, e-,e+, , 
EHN2 | 250 – 280 GeV/c | h, 
ECN3 | 75 GeV/c | K
Spill intensity range | 105 – 3 x108 ppp | 
(3)

## Slide 28: EAST AREA BEAM CHARACTERISTICS
Secondary Beam
Parameter | Value
Proton beam momentum [GeV/c] | 24 (requests to go lower for longer spills)
Maximum flux per PS spill | 3E11 for T09-T11 8E11 for T08
Maximum # protons per second (individual limits for RP) | 3.3E10 for T09-T11, 6.6E10 for T08, 4.4E10 for F6D
Primary Beam (T8)
Parameter | T09 Target | T10/T11 Target | 
Beam Line | T09 | T10 | T11
Max Momentum (GeV/c) | 15 | 11.5 | 3.5
Δp/p (%) | 0.7 to 15.0 | 0.7 to15.0 | 0.7 to15.0
Maximum intensity/spill | ~few x 106 | ~few x 106 | ~few x 106
Available particle types | Pure electrons (T09) or mixed/pure hadrons or pure muons |  | 
Parameter | T09 Target | T10/T11 Target | 
Beam Line | T09 | T10 | T11
Max Momentum (GeV/c) | 15 | 11.5 | 3.5
Δp/p (%) |  |  | 
Maximum intensity/spill | ~few x 106 | ~few x 106 | ~few x 106
Available particle types | Pure electrons (T09) or mixed/pure hadrons or pure muons |  | 
W. Devauchelle - BI Day 2024

## Slide 29: Multi-user lines PS East Area and SPS North Area beam time requests
Average overbooking of the SPS EHN1 lines about the same as in 2021 (and much more than in 2022!)
  2021 we had only 18 weeks protons in EHN1 (2023 22 weeks) and had the first beam since 2018
More calorimeter and exclusive main user requests now compared to 2021
 |  | Weeks requested by SPSC related user | Weeks requested for LHC experiments R&D | Weeks requested by other users | Total Requested | Weeks available
PS T9 & T10 | protons | 8 main 0 parasitic | 22 main 0 parasitic | 19 main 0 parasitic | 49 main 85% | 58
SPS H2, H4 | protons | 28 main 0 parasitic | 28 main 1 parasitic | 13 main 7 parasitic | 69 main 157% | 44
SPS H6, H8 | protons | 0 main 0 parasitic | 64 main 22 parasitic | 24 main 18 parasitic | 88 main 200% | 44