## Slide 1
An Introduction to Medical Imaging Devices
Martin L. Purschke

## Slide 2
Who am I? A nuclear physicist working on the sPHENIX experiment at BNL’s Relativistic Heavy Ion Collider (RHIC)
Long Island, NY
RHIC from space
Manhattan

## Slide 3
Wait, what?
So this guy is a collider phycisist?
What qualifies me to talk about Medical Imaging?

## Slide 4
Medical Imaging
Medical Imaging has a lot in common with nuclear physics (and accelerator/beam physics)
Yes, there is the part where an actual M.D. comes in and wants to see something in particular, but up to that point, it’s physics and engineering
Radiotracers, nuclei, detector technology, readout, analysis, …. All the natural habitats of physicists and engineers.
As a matter of fact, the by far most common diagnostics method using Te99 was developed at Brookhaven National Lab – you have accelerators, nuclear physicists, and the infrastructure
But also, I was deeply involved what I think is until today the coolest application of PET (Positron Emission Tomography) – the RatCAP
For the first time, the RatCAP allowed the imaging of the brain of an awake rat
I’ll talk about this later some more, but this is my foray into Medical Imaging, and why I know all this stuff.

## Slide 5
Medical Imaging Technologies
Without any claim to completeness (and some of what I’ll cover)
SPECT (a little)
PET
MRI
X-Ray
CT
Multi-modality imaging
(there are many more, often variants of a common theme)
You will see that nuclear physics, DAQs, and data processing plays a prominent role.
There is a place for you in this field if you can do TDAQ, or are a “data engineer”

## Slide 6
Medical Imaging Technologies grouped by other metrics
Not using ionizing radiation
MRI
Using ionizing radiation
X-Rays
CT
SPECT
PET
It’s ironic that the only technology here NOT using ionizing radiation, formerly known as Nuclear Magnetic Resonance Imaging (NMR), had to be renamed because people were freaking out over the word “nuclear” in the name

## Slide 7
What does one want, and how does one choose a technology?
Sensitivity - How well can you actually see what you are after?
Selectivity - How well does your method distinguish between, say, benign and malicious tissue?
Contrast - What is the dynamic range between the different features in your image?
For example, if you want to see if a bone is broken, X-rays (and by extension, CT) gives you a high value on all three – the contrast between bone and tissue is large
X-ray for a mammogram has a low value on all three

## Slide 8
Each modality has strengths and weaknesses
For example, MRI has excellent position resolution but is “blind” to metabolic processes
PET can show the metabolism but the position resolution is poor
Think of different modalities as looking at the same thing in different ways
MRI by and large images the density of protons (think water in tissue)
X-ray and CT image the electron density in tissue (that’s why bones show up so nicely)
SPECT and PET can be tailored to show metabolic processes in tissue

## Slide 9
Why PET is so cool
PET can be tailored to show metabolic processes in tissue
Remember, the chemical properties of a radioactive isotope are the same as the stable isotope – it will be used by the body indiscriminately
So you take a molecule that is used in your body for some process – sugar, dopamine, what have you, swap out an atom for a radioactive one
You get “hot spots” where lots of said molecules are used by the body
For example, a tumor that grows uses a lot of energy (sugar), so radioactively tagged sugar will accumulate a lot of activity there
Or tag molecules that the brain burns when it “thinks” – active areas show up as hot spots
Frequently used radiotracers are 11C and 18F (all beta+ = positrons)

## Slide 10
An Example - 18F
18F-FDG - Fludeoxyglucose is one of the commonly used radiotracers
It is a sugar, so it accumulates where the body burns a lot of energy (“uptake”)
Unusual “hot spots” can point to cancer
Whole-body PET scan using 18F-FDG to show liver metastases of a colorectal tumor

## Slide 11
11C Raclopride – Brain activity
Raclopride is a molecule that binds to dopamine receptors in the brain
Can show which part of the brain does what
Different activities stimulate different areas in the brain (singing, reading, listening, etc)
Image of a rat brain taken with the RatCAP using 11C-tagged raclopride
(I actually took those data!)

## Slide 12
SPECT (Single Photon Emission Computed Tomography)
Single Photons are detected
Low efficiency (next slide)
Can cherry-pick the photon-emitting isotope ( 125I, 60Co) (not that much of a menu though)
Ring of Photon
Detectors
Radionuclide decays, emitting γ rays.
Gamma lies on line defined by detector and collimator (known as a line of response or a LOR).
I have never worked with SPECT, but I see most groups abandoning this
Some Alzheimer’s-related research still shows up

## Slide 13
SPECT - collimation and low efficiency
Without a collimator, there’s no sense of “direction” of the photon – the photon could come from anywhere in the field of view
Only the Collimator selects a direction
Dramatic reduction in efficiency – most photons are lost in the collimator

## Slide 14
SPECT Image Examples
sense
S. Shokouhi, Vanderbilt U
Example of a phantom – the holes are filled with radiotracers (similar to what was used in these images)
3rd-largest bores can no longer be resolved
Different collimator configurations

## Slide 15
PET – Positron Emission Tomography
Back-to-back 511KeV photons are detected
No collimator needed
LOR defines the direction
Best I can tell, all the action today is in PET, not SPECT
Ring of Photon
Detectors
Radionuclide decays, emitting +.
+ annihilates with e– from tissue, forming back-to-back 511 keV photon pair.
511 keV photon pairs detected via time coincidence.
Positron lies on line defined by detector pair (known as a line of response or a LOR).

## Slide 16
The PET principle
PET measures coincident photons from an electron-positron annihilation
Each such coincidence establishes one “Line Of Response” (LOR)
It is mostly intuitive for folks working with accelerators
Commercial scanners often have sophisticated electronics or firmware for coincidence processing
Or you record all photon counts with detector element info and a time stamp (“list mode”)
When two photons are measured “simultaneously” within a certain time window, you consider them a coincident pair
You can either record all the data and do the analysis offline, or the PET scanner has that ability built in and reports only the coincidences (commercial scanners usually do)
511 keV
511 keV
e+
e-
511 KeV
511 KeV

## Slide 17
Offline Coincidence Processing
You cannot know which ones are true and which ones are random coincidences
But you can get an estimate of the random rate and their contribution by adding a delay to the window and destroying all true coincidences
time
Coincidence window
✔
✔
✔
✔
✔
✔
✕
✕
✕
✕
time
When all is done, the output from the coincidence processing is a sinogram
There are more corrections, such as scatter corrections, attenuation, etc…

## Slide 18
Photon Detection
511KeV is a tiny amount of energy when it comes to detection
High-Energy physics tends to ignore anything below a few 100 MeV
The resolution of a detector is primarily determined by the “count” of what you measure – the famous sqrt(N) error formula
511 KeV produce between 2000 and 20000 photons, the latter in the workhorse scintillator, NaI
NaI falls flat in basically every other metric – much too slow, it’s very hygroscopic, not very dense
Many new scintillators on the market, some are fantastically expensive – sometimes hundreds of dollars per cubic millimeter
Readout through PMTs, APDs, SiPMs…
We are looking for fast signals so the coincidence detection works.
For that a few hundred picoseconds timing resolution is enough
Keep that number in mind for later…

## Slide 19
A selection of scintillators

## Slide 20
Readout
Photomultipliers
Avalanche Photon Detectors
Heat Dissipation – needs cooling
Bulky
Needs High Voltage (~ 2KV)
Useless in a magnetic field
Excellent linearity and noise
Small footprint
Works in a magnetic field
Modest supply voltage ( ~400V)
Significant noise

## Slide 21
SIPMS
Silicon Photomultipliers are slowly taking over
Lower operating voltage (~ 70V)
Faster
Cheaper
Silicon base can allow for additional electronics features

## Slide 22
A Real-Life Pet scanner – MiniPET (KTH, Sweden)
206mm opening
Hamamatsu H9500 Position-sensitive PMT
LYSO Crystals 1.27 x 1.27 x 12 mm3
32 x 32 crystals per module, 12 modules
FPGA-based data processing
readout through Ethernet
Image of a rat heart

## Slide 23
Depth of Interaction – DOI measurement
DOI can resolve the LORs better
Usually requires some dual-ended readout for light-sharing
More complex readout
The parallax effect limits the position resolution
Problem in particular in smaller scanners
These LORs will give the same signal
The Parallax Problem
Miyaoka, R.S.; Lewellen, T.K. Nuclear Science Symposium, 1997. IEEE , vol.2, no., pp.939-943 vol.2, 9-15 Nov 1997

## Slide 24
0 degrees
Image Reconstruction - Sinogram
We are “looking” at the object from many angles instead of just one
This is what you do inadvertently with your eyes to see sharper

## Slide 25
0 degrees
22.5 degrees
Image Reconstruction - Sinogram

## Slide 26
0 degrees
22.5 degrees
45.0 degrees
Image Reconstruction - Sinogram

## Slide 27
0 degrees
22.5 degrees
45.0 degrees
67.5 degrees
Image Reconstruction - Sinogram

## Slide 28
0 degrees
22.5 degrees
45.0 degrees
67.5 degrees
90.0 degrees
Image Reconstruction - Sinogram

## Slide 29
0 degrees
22.5 degrees
45.0 degrees
67.5 degrees
90.0 degrees
112.5 degrees
Image Reconstruction - Sinogram

## Slide 30
0 degrees
22.5 degrees
45.0 degrees
67.5 degrees
90.0 degrees
112.5 degrees
135.0 degrees
Image Reconstruction - Sinogram

## Slide 31
0 degrees
22.5 degrees
45.0 degrees
67.5 degrees
90.0 degrees
112.5 degrees
135.0 degrees
157.5 degrees
Image Reconstruction - Sinogram

## Slide 32
Back projection
Filtered back projection
Image Reconstruction – Sinogram and reconstruction
It’s called a sinogram because any feature traces a sine wave

## Slide 33
Image Reconstruction: Back Projection
Remember each projection only gives you a line
You fill the different projections in a histogram (each line)
Sources in the same spot in different projections accumulate

## Slide 34
Image Reconstruction: Filtered Back Projection
A naïve back projection adds a lot of noise and artifacts to the image
The BP is the equivalent of a Fourier transformation
The filter clips “harmonics” that are beyond the Nyquist resolution and reduces noise
There a lot of different filters in use
A lot of papers written about that
You can get your PhD for finding a better one…
I’ll get to another (and for folks like us, more interesting, I think) image reconstruction method in a minute… but first:

## Slide 35
The RatCAP and Derivatives
One wants to use PET to study the neurophysiological activity and behavior in laboratory animals
Understand and treat illnesses in humans.
However, animals needed to be anesthetized during PET imaging.
Anesthesia can greatly depress the very brain functions and affect the neurochemistry that one is trying to study
Cannot study animal behavior while under anesthesia
The “Holy Grail” – study brain processes in the awake animal using PET

## Slide 36
The Quest
To convert something like this…
… into something like this…

## Slide 37
RatCAP: Rat Conscious Animal PET
A miniature, complete full-ring tomograph mounted to the head of an awake rat.
Compact, light weight (< 200 g), low power detector
Small field of view (38 mm dia. x 18 mm axial)
Attached to the head of the rat and supported by a tether which allows reasonable freedom of movement for the animal

## Slide 38
RatCAP: Rat Conscious Animal PET
Ring containing 12 block detectors of 32 2x2 mm2 x 5 mm deep LSO crystals with APDs and integrated readout electronics
APD
(Hamamastu
S8550)
4x8 array of
LSO crystals
(2x2x5 mm3)
Actual RatCAP Ring
All the magic lies in the RatCAP ASIC, which takes care of all the signal processing. Reports channel # with timestamp

## Slide 39
RatCAP Derivatives
Once you have a working ASIC, you try to re-use it for all kinds of other cool things…
Original RatCAP V1
RatCAP V2 MRI compatible
Breast Scanner Prototype
Breast Scanner UPenn Scanner
This “MRI compatible” was a major piece of work – no ferromagnetic material anywhere

## Slide 40
The Technology went commercial
BNL licensed the technology to an upstart company

## Slide 41
Other opportunistic applications – Plant Scanner
Here, the plant is fed 11C-labled CO2 with a cuvette
The large opening is needed mostly to allow the plant to fit into the scanner.
Most of the FOV is just empty
This supports research for better biofuels
You can track where the produced sugar is going in the plant

## Slide 42
Alternate Iterative MLEM-based Image Reconstruction
“Maximum Likelihood Expectation Maximization”
One divides the Field of View (the active area of the scanner) in small “voxels”. Ours are 1mm3 cubes
The System Matrix A describes the probabilities that the decay at a given voxel position results in the photons ending up in a given pair of detector elements (“LOR”)
Number of LORs go (N2) for N detector elements
FOV size = number of voxels goes (N3/2)
Matrix has dimensions number voxels x number of LORs
Each matrix is specific for a particular geometry
So: Activity Distribution x A = Detector Response
Great! Activity Distribution = A -1 x Detector Response
This is what you measure
This is what you want
Not so fast!
So all it takes is to invert the matrix, and we are set

## Slide 43
The System Matrix is HUGE
No way to invert such huge matrices
So we resort to an iterative process that “sculpts” an activity distribution until it matches the measured detector response
This method beats Filtered Back-Projection hands-down
But it is really computing intensive – we cannot yet do this for the full Breast Scanner
 | Ratcap | BS proto | UPenn/BS
Voxels | 15113 | 194427 | 1328360
LORs | 72192 | 342784 | 22444800
Matrix elements | 109 mill | 66646 mill | 29814774 mill

## Slide 44
Computing Challenges
As small as the detectors look, they produce a lot of data
The largest UPenn / Breast Scanner can produce 450MB/s ( 1.5 TB/hour )
The computing challenge is two-fold:

## Slide 45
Computing Challenges
The ~250,000 CPU hours for a system matrix is actually the smaller problem
The simulation of any voxel is independent of all others
You can distribute the workload as much as you like over as many CPUs as you like
We have been using an online cluster (of my day-job at the PHENIX experiment at the Relativistic Heavy Ion Collider) opportunistically
We have joined the Open Science Grid (OSG) to disentangle us from the seasonal availability of our local cluster
We can get about 1000 CPUs / day and can get a new matrix in about 2 weeks
We do this only a few times a year

## Slide 46
Reconstruction Times
The iterative, “activity-sculpting” process yields a usable image after about 100, a good one after 500...1000 iterations. Each iteration involves 2 matrix multiplications
At one point, in 2004 or so, the small RatCAP reconstruction was perceived to be a big problem
No more, we really have that optimized, 10-15 minutes per image.
Caveat: it's not just one image. We do dynamic images, time-sliced, 2-15 minutes depending on the experiment. They get correlated with external events. Like a movie of the activity distribution.
The plant scanner takes about 6-8 hours per image on a multi-core CPU. Multiple machines can work in parallel on independent images
The Breast Scanner? We don't know, we cannot presently do it “our way”. Estimated 400 hours, but we didn't bother to make a system matrix yet
If you think of clinical mammography applications, the patient expects to walk out of the facility with a preliminary result
No way this can be done – results come weeks later

## Slide 47
Switching gears…
Let’s talk about MRI for a few minutes.
I have been going on about MRI-compatible scanners now for a while…

## Slide 48
Protons have a magnetic moment “built in” [ not only protons, btw ]
Normally, that momentum vector can point whatever which way
Apply a magnetic field, and those momenta align their directions along the field
If you “push” them out of alignment, the a-momentum vectors start to precess
They precess with a characteristic Larmor frequency ω= g x B0
g is the gyromagnetic ratio, 42.58MHz/T for a proton
By setting the right field strength, you can dial in a particular Larmor frequency
For example, with 1.5T, you get 65.16 MHz
MRI in 5 minutes 

## Slide 49
If you apply a radio wave with the exact Larmor frequency, you “tip” the momentum vectors in a coherent way
As they relax, they emit RF at the same frequency
If there are many protons -> strong signal; not so many –> no or weak signal
you can use this to measure the proton density
MRI in 5 minutes 

## Slide 50
First of all, if you have a totally homogeneous field, all protons have the same Larmor frequency – not very useful
On top of your homogenous field, you apply a small gradient so the frequency becomes position-dependent – now we are getting somewhere
You could send a number of pulses which “scan” the frequency range and measure the response – was done in the early days, way too slow, but it would work
Rather, one sends a pulse with a narrow “white” frequency band, and measures the response for each frequency at the same time
One shot gives you the proton density along the field gradient (one axis of a picture)
MRI is hard
Why didn’t I think of that, and get the Nobel Price! Well…

## Slide 51
3 sets of coils let you make any gradient direction you want – loops for z, and “saddles” for x and y
Modern scanners use “fingerprint” designs with high efficiency
Making Gradients
Remember, that’s on top of a static high-strength field

## Slide 52
The points {A,B,C} and {D,E,F} radiate at different frequencies according to the field strength
However, A,B and C are indistinguishable! (and D,E,F, too)
Throw in a phase-encoding gradient:
But Wait! We are missing one dimension!
Gradient
By momentarily “speeding up the clock” for a region, you gain a phase shift for protons there
And you can assign a particular phase shift to a region

## Slide 53
Now you can get x by frequency and y by phase.
N pixels in y form an equation system with N variables (so a 100x200 image takes 200 measurements)
You need N measurements with different phase encodings to solve this
Remember: by the time you measure, A and B have the same frequency again (only a phase difference)
Phase and Frequency Encoding
2 pixels A and B take two “shots” S0 and S1
S0 = A sin ( ωt ) + B sin ( ωt ) = (A + B) sin ( ωt )
S1 = A sin ( ωt ) - B sin ( ωt ) = (A - B) sin ( ωt ) A = ½ ( S0 + S1)
B = ½ ( S0 – S1)
This is easy only for a 1800 phase shift

## Slide 54
Back to the MRI-compatible PET scanners
The trend in modern medical imaging goes to multi-modal imaging
Combine two different modalities to combine their strengths
PET-MRI – MRI gives you sub-millimeter resolution, PET gives you the metabolic information

## Slide 55
The 9.4T UPenn MRI for small animals
Our UPenn PET scanner was designed to fit into the bore of the Varian 9.4T (as strong as they come)
BTW – 9.4T -> 400MHz Larmor. You can buy them “by frequency” – 7T (300) etc
Our “UPenn” PET scanner mounted on the MRI tube
Going in
RF Coils and pickups
Magnet bore
Scanner

## Slide 56
The Brutally Hostile Environment inside a MRI Magnet
If you were to get too close to the bore with a steel screwdriver, you would not be able to hold on to it
A number of early versions of boards got ripped apart because some ferromagnetic material was present in unexpected places ( e.g. Ni in solder joints)
The RF pulses totally overwhelm the PET electronics.
The plot shows counts registered in the PET scanner as a function of time – the spikes are the interference from the RF pulses
Still:
It did work!

## Slide 57
Simultaneous PET-MRI Rat Brain Images
MRI
PET
Overlay
This is sort of the “money shot” – one of the all-out-everything-worked scans of a rat
What you want is simultaneous PET and MRI. If you do it sequentially, it’s much harder to combine the images

## Slide 58
Wrapping up with 3 quick items
Out of a stroke of good luck we were able to make an actual clinical test on human patients in Taiwan
Volunteering women who had already been diagnosed with breast cancer were scanned with our breast scanner prototype and MRI
Breast Scanner Prototype
The protocol and MRI magnet /gantry design didn’t allow for simultaneous scans though
We had to resort to sequential PET-then-MRI scans
Our about 3cm deep FOV in the prototype allows to only image a thin “slice” of tissue
FDG scans – sugar uptake by the tumor
Fantastic images despite some compromises

## Slide 59
Human trials
Transverse
Sagittal
Coronal
MRI
PET
PET-MRI (sequential)
Subject 3 –
8.67 mCi injection
Silicone Implant

## Slide 60
Human trials
Subject 3 –
8.67 mCi injection
Silicone Implant
The volunteering women already knew they had breast cancer
In one patient, our scanner detected a 2nd tumor that had been missed (same breast)
Once / if the full-sized scanner can get FDA approval, new tools for breast cancer detection will be available

## Slide 61
Time-Of-Flight PET – the Holy Grail of PET
Remember my “a few 100 ps” timing resolution to be able to find coincidences?
Well, there is another front. If we could use TOF to restrict the decay position along the LOR, the images would dramatically improve
A human scanner is ~30cm
10ps timing resolution – 3mm spatial resolution
Sub-100ps timing is hard! Especially if you have just 511KeV to work with.
Chronometer
PET Scanner
Time of Flight
+
-
Start
Stop
Patient
Δ time= Δ position

## Slide 62
TOF-PET – the Holy Grail of PET
If you manage to restrict the position to just ½ of the LOR, you gain a factor of 4 in image quality (or can use ¼ of the radiation dose)

## Slide 63
SPAD – Single Photon Avalanche Diodes
In my book, the most promising TOF-PET technology
At its core, a diode operated above its breakdown voltage
Avalanche == high current == healthy signal == good timing characteristics
You need to control the breakdown, need circuitry to prevent damage

## Slide 64
SiPMs and dSiPMs – digital Silicon Photomultiplier
dSiPMs – SiPMs with full digital readout and control
Control for individual pixels!
Google it..

## Slide 65
People who helped
Special thanks to
David Schlyer, BNL, USA
Paul Vaska, Univ. of Stony Brook, USA
Réjean Fontaine, Univ. of Sherbrooke, Canada
Steve Pickup, Univ. of Pennsylvania, USA
Sepideh Shokouhi, Vanderbilt University, USA
Fine Fiedler, Helmholtz-Zentrum Dresden-Rossendorf, Germany
Thank Y’all!

## Slide 66
In-Beam PET
Energy Deposition as function of tissue depth:
Photons or electrons
Protons or ions
Trend in cancer therapy: Proton beams (or even light-ion beams)
More complicated, expensive (hadron beams are much harder to make)
But: much superior to treatments with photon or electron beams
Why?
Depth
Energy Loss
“Bragg Curve”
Cell Damage
Using protons or Ions, you can target a particular depth in the tissue
(read: where the tumor is located)
Photons and electrons cause damage all along their path
Protons and ions cause damage at one narrow point

## Slide 67
Hadron Beams are hard
This is the Paul Scherrer Institute’s Proton Treatment Facility
Way beyond the capabilities of a typical hospital!
https://www.psi.ch/protontherapy/

## Slide 68
Brain Tumors
Here you can usually not afford to cause damage along the entire beam path
Proton beams ideally suited!
You “wiggle” the depth of the Bragg peak quickly by either adding more material in front, or changing the beam energy
Proton Beam
Absorber
Now you have a very powerful weapon!
And a new problem: “Friendly Fire”
If you miss your target area, that part of the brain (or the optical nerve, or whatever you hit) is dead

## Slide 69
Throw in in-beam PET
The idea: Use the positrons produced by the beam to image the “target area”
More energy loss = more positrons
Therapy Verification
Mispositioning
Correct for individual “body types” (tissue density)
Organ movement correction