Review of
Development Plan 2024
White Rabbit Switch
Calibration
4 December 2024
CERN
Adam Wujek
White Rabbit
Collaboration

---
Calibration overview
Do I need to calibrate my devices?
It depends…
WR devices should provide sub-ns accuracy out-of-the-box
But one could use not officially support equipment (e.g., SFP)
If better accuracy is needed then calibration may be required
WR switch
SFP modules
Fibers (also add/drop filters)
DelayAsymmetry (difference in ps) two fibers
delayCoefficient (difference as unit-less) one fiber
Coaxial cables

---
Link model
Timestamp point
Timestamp point
δMS and δSM deepens on fiber length and used wavelength
Difference in fiber length can be avoided with using single fiber
For LPDC ports εm and εs (bitslide) are equal 0

---
Mapping of parameters to dot-config
(δMS-δSM)/δSM is alpha (delayCoefficient) stored in
CONFIG_PORT01_FIBER=0
CONFIG_FIBER00_PARAMS="alpha_1310_1490=0.0002743"
ΔTXsfp, ΔRXsfp as tx=X, rx=Y in
CONFIG_SFP00_PARAMS="vn=Axcen Photonics,pn=AXGE-3454-0531,vs=AX12390009629, tx=0,rx=0,wl_txrx=1490+1310"
ΔRXM, ΔRXS – PORTXX_INST01_INGRESS_LATENCY
ΔTXM, ΔTXS – PORTXX_INST01_EGRESS_LATENCY
For LPDC ports εm and εs (bitslide) are equal 0, otherwise are read in runtime from HW

---
SFP database matching
Why needed? To get relative calibration values of SFPs RX and TX
For absolute calibration, values will be read from SFPs EEPROM (see SFF-8472, rev 12.4.2)
Entries in dot-config like:
CONFIG_SFP00_PARAMS="vn=Axcen Photonics,pn=AXGE-3454-0531,vs=AX12390009629,
tx=0,rx=0,wl_txrx=1490+1310"
vn – vendor name
pn – product name
vs – vendor serial
tx,rx – relative TX, RX delays (they may be negative!)
wl_txrx – TX and RX wavelengths in nm
VN, PN, VS, TX wavelength are read from SFP’s EEPROM
Matching order:
TX wavelength, PN, VN, VS
TX wavelength, PN, VN
TX wavelength, PN
After match, deltas are taken from tx= and rx=

---
Fiber database matching
Why needed?
To get delayCoefficient (alpha) of a fiber
Fiber has no memory that can store this value
Entries in dot-config like:
CONFIG_PORT01_FIBER=0
CONFIG_FIBER00_PARAMS="alpha_1310_1490=0.0002743"
Port has to define which fiber type is used (marked with gray)
For matching alpha_XX_YY, values of XX and YY are taken from wl_txrx parameter of SFP DB entry. Like:
CONFIG_SFP00_PARAMS="vn=Axcen Photonics,pn=AXGE-3454-0531,vs=AX12390009629,
tx=0,rx=0,wl_txrx=1490+1310"
If alpha_XX_YY is not matched, swapped XX and YY is tried (alpha_YY_XX)
Note alpha in the opposite direction is not -alpha, but: which is equivalent to:
- frac{alpha} {1+alpha}
{frac {1} {1+alpha}} - 1

---
Relative calibration
Unable to measure all delays separately
Use of the arbitrary value works pretty well

---
Golden calibrator
Created once
Kept as a reference for other devices
Possible to create another Golden calibrator from original one (local Golden calibrator
WR switch is not enough, needs pair of golden SFPs
Can use independent golden calibrator
Device calibrated with one will show offset when connected to another

---
Calibration

---
Getting measurements
Delays
wr_mon (try different options)
wrs_dump_shmem
Oscilloscope/Time Interval Counter
PSS level at 850mV
Make average over some time
Be aware about temperature variations
Before starting the calibration
Non-LPDC (13-18) ports can have up to ~70ps offsets between WRS reboots/linkup
Make sure used calibration values are for a proper WRS HW version
Make sure equal lengths of cables are used when necessary
Precise enough measurement devices are used (Oscilloscope, Time Interval Counter)

---
Equipment requirements
Fibers
One short (few meters)
One long (few kilometers)
Adapter to connect two fibers
WRS devices
Golden calibrator+device to be calibrated
Or two identical devices to make golden calibrator
Two oscilloscope cables with the same length (or known lengths)
Oscilloscope or Time Interval Counter
Stable 1PPS and 10MHz source

---
Reference fiber latency
Set all delays and alpha to 0
PORT01_INST01_EGRESS_LATENCY=0
PORT01_INST01_INGRESS_LATENCY=0
CONFIG_FIBER00_PARAMS=””
Measure delayMM of (e.g., read from wr_mon)
Short fiber (delayMM1)
Long fiber (delayMM2)
Short+long connected with adapter (delayMM3)
For LPDC, bitslide can be skipped (is 0)
δ1 = delayMM3 - delayMM1
δ2 = delayMM3 - delayMM1

---
Fiber asymmetry
Set all delays and alpha as 0
Measure PPS skew of
Short fiber (skewPPS1)
Long fiber (skewPPS2)
Skew is positive if
PPSslave follows (right side) PPSmaster
Calculate α (delayCoefficient) with the formula:
Take δ2 from Reference fiber latency calculations

---
Calibrator pre-calibration
Assume that both devices are identical
(We already know that SFPs used in both directions are different, but let’s ignore it)
∆TXM + ∆RXM = ∆TXS + ∆RXS
Then: delayMM1 − δ1 = 2(∆TX + ∆RX)
When we assume that golden calibrator has no asymmetry
∆RX = ∆TX
This assumption is why we have „relative calibration”
delayMM1 − δ1 = 4(∆TX) = 4(∆RX)
In other words: ∆TX = (delayMM1 − δ1)/4

---
WR Device calibration
Set delays on golden from “Calibrator pre-calibration”
Set correct alpha measured in “Fiber asymmetry”
Restart devices to apply changes
Connect devices
Read delayMM from slave
Calculate: ∆TXS = ∆RXS
∆’TXS = ∆’RXS = ½ * (delayMM - ∆TXM - ∆RXM – δ1)
Save deltas on slave, restart it
Measure skewPPS on oscilloscope and correct ∆’TXS = ∆’RXS
Calculate:
∆TXS = ∆’TXS - skewPPS
∆RXS = ∆’RXS + skewPPS
Save new values on slave, restart WRS
Repeat measurement of skewPPS and last calculation if needed
(use ∆TXS as ∆’TXS and ∆RXS as ∆’RXS)
Repeat this slide for all ports

---
SFP calibration
Keep the setup from the previous slide
Measure skewPPS on oscilloscope
Calculate:
∆TXSfp = - skewPPS
∆RXSfp = skewPPS
Save new values on slave, restart WRS
Repeat measurement of skewPPS and last calculation if needed
(subtract or add the new value of skewPPS to ∆TXSfp and ∆RXSfp)

---
Calibration automation
As today, only WR switch is supported (no node)
Some steps can be automated
Data collection from measurement device
Supported Tektronix MSO64b and
LeCroy WaveRunner 8104
More can be easily added
Possible manual input
Re-configuration of WR Switch
Calculations
Needs some manual interactions, e.g., fiber changing

---
Calibration automation, requirements
2x WR Switch
One local calibrator
Or two the same to make golden calibrator
Optional Oscilloscope (manual input supported)
Tektronix MSO64b
LeCroy WaveRunner 8104
1-2 fibers depends if golden calibrator is available
Pair of SFPs
PC with installed python to run script
Optional time/freq reference to improve
calibration

---
Calibration automation
Supported Calibration Scenarios
Full calibration
Create new calibrator
Calibrate against the new calibrator
Measure delays of reference fibers
Calculate delay asymmetry coefficient (alpha)
Partial calibration
Create new calibrator
Calibrate against the new calibrator
Provide delays of reference fibers
Normal calibration
Calibrate against already existing calibrator
Calibration of SFPs is included
See https://ohwr.org/project/wr-switch-sw/-/tree/master/userspace/host_tools/calibration
wrs-user-manual Appendix 5
https://ohwr.org/project/wr-switch-sw/-/wikis/uploads/d1f78666704fb292982453e1429b9f10/wrs-user-manual-v7.0.pdf

---
Absolute calibration (future)
Electrical
Electrical to Optical (SFP)
In-situ Fiber calibration (as alternative to lab measurements)

---
In-situ fiber calibration
Swapping of one-way delays
Step 1: Host A transmits with wavelength X, Host B with Y
Measure:
Clock offset against stable source d1
Step 2: Host A transmits with wavelength Y, Host B with X
Measure:
Clock offset against stable source d2
(d2-d1)/2 is the delayAsymmetry
Host B
Host A
Host B
Host A

---
In-situ fiber calibration
Three one-way delays (with tunable SFP)
Host B transmits with wavelength X all the time, Host A with a help of tunable SFP transmits two different wavelengths
Step 1: Host A transmits with wavelength Y1
Measure:
Round_Trip_Time1
Step 2: Host A transmits with wavelength Y2
Measure:
Round_Trip_Time2
Based on wavelength X, Y1, Y2, Round_Trip_Time1, Round_Trip_Time2 it is possible to calculate delayCoefficient (alpha) for X, Y1
Host B
Host A
Host B
Host A

---
Future work
Characterize changes with temperature changes
delayCoefficient
Port delays
SFP delays
Apply corrections in runtime

---
More info
WR Calibration procedure v1.1
https://ohwr.org/project/white-rabbit/uploads/76cdbdbadccc9d6c54d5caf246550fbf/WR_Calibration-v1.1-20151109.pdf
WR Calibration Application Note1:
https://ohwr.org/project/white-rabbit/-/wikis/uploads/11f29675590d6d62b296064bd4e87921/WR_Calibration-Application_Note_1_v0.4.pdf
WRS: User Manual, Appendix 5:
https://ohwr.org/project/wr-switch-sw/-/wikis/uploads/d1f78666704fb292982453e1429b9f10/wrs-user-manual-v7.0.pdf