## Slide 1: BSRH: Establishing performance at the LHC
Jan Pucek on behalf of BI-PM
HL-LHC collaboration week in Genoa, October 2024

## Slide 2: Important parameters (prepare, carry out, analyze MD)
Wire scanners BWS:
  ßy = 418.23m
  ßx = 185.15m
Coronagraph BSRH:
  ßy = 366.63m
  ßx = 193.8m
Primary collimators TCP (V):
  ßy = 71.49m
  ßx = 148.43m
For εN = 2.5um rad
σy = 386 um
σx = 257 um
σy = 362 um
σx = 263 um
σy = 160 um
σx = 230 um

## Slide 3: Outline
Coronagraph – principle
Calibration measurements
Data analysis
Performance overview
System status
Summary table

## Slide 4: Coronagraph – principle

## Slide 5: Coronagraph – principle
Telescope provides an image of an object
In case of measuring low intensity tails, the occulter is not sufficient – diffraction created at the aperture will disturb the image
Coronagraph is an “enhanced” telescope that uses “field” lens to image the aperture on the Lyot’s stop – blocking the diffraction of the aperture
1. Telescope – two lenses
Aperture – creates diffraction
Aperture
Occulter to block the ”core”
2. Telescope with occulter
Field lens to image the aperture onto the “eyepiece” lens
Lyot’s stop
3. Coronagraph

## Slide 6: Introduction – the source
Distance between D3 and BSRTM (extraction mirror) ≈ 21m
Synchrotron radiation emitted in a narrow cone (due to Lorenz transformation) ≈ 140 µrad
D3
D3
D3
D3
Beam 1
Beam 2
RF cavities
D4
D4
SR radiation
Extraction mirror
*not to scale

## Slide 7: Introduction – the source
Sharp edge of the extraction mirror within the light spot -> source of diffraction in horizontal plane
Simulation and figure produced by D. Butti
When imaged
Diffraction tail

## Slide 8: Concept of diffraction
Vertical plane
Horizontal plane
Measured profile ≠ True beam profile
Coronagraph can lower the diffraction contribution of the sharp mirror cut

## Slide 9: Calibration measurements - imaging
Pixel size calibration:
  Bunches with 3 different emittances
  Comparing the BWS and the BSRH profiles in both planes independently
Charge scaling calibration:
  Requires the full measured profile from the BSRH and the corresponding FBCT measurement
Calibrating this factor

## Slide 10: Calibration measurements
Results: σ = (277 ± 8)μm
m = (0.64 ± 0.006)
pixel size = (18.3 ± 0.2) μm/px
Results (L-BFGS-B):σ = (320 ± 30)μm
m = (0.68 ± 0.03)
pixel size = (17.0 ± 0.8) μm/px
BWS ⮾ PSF
BSRH
σy = 362 um
σx = 263 um

## Slide 11: Calibration measurements
The charge scaling factor = (42.7±1.7)p+/count

## Slide 12: Illustration of the analysis procedure
BWS
BSRH
Raw data
Gaussian Fit
Integration region
Scaling factor
Raw data
Projection + Fit
Integration region  Charge in halo

## Slide 13: Machine development session(s) - overview
How to asses the performance of the instrument?
Scrape the beam to known size  beam halo = 0
Blow up a bunch (ADT) to populate the halo region
Measure charge in halo by two instruments (BWS, BSRH)
MD1:
Nominals and pilots (- charge resolution of the BWS)
Tried to measure changes in the profile
MD4:
Pilots only (+ charge resolution of the BWS)
Online analysis ready
Lessons learnt

## Slide 14: Data handling
Subtract background (known from collimator position) makes the measurement only relative
1.6x109

## Slide 15: Simulation proof of linearity
Diffraction creates offset, yet linearity unchanged (for the “Gaussian” blow up)
Courtesy of D. Butti
Constant offset

## Slide 16: Performance overview

## Slide 17: Vertical plane
Halo above 3.5σ
Halo above 4.7σ
Total population = 1011
Total population = 9x1010
Blow up
Steady state
Blow up
Steady state
Scraping
Green curve = 15s moving average of the BSRH data, gated on 10 pilot bunches
Blue curve = BWS (average of IN/OUT) of the single pilot
Red curve = FBCT reading of the single pilot

## Slide 18: Horizontal plane
1
Halo above 3.5σ
Halo above 4.7σ
Reason behind the difference is unclear
Factor 2…
Single bunch measurement, 2s moving average, limit of “static” core
Measurement of 10 pilots, one blown up, 15s moving average

## Slide 19: System status
YETS underway
  Solving the image intensifier ripples
  Splitting the optical line BSRT to BSRH to make the systems work simultaneously
  Apodizer to be tested
Demonstrate potential profile/image deconvolution
Prepare procedures for automatized data taking during physics fills
Only way to increase the resolution is to increase the beam size
Imaging
Coronagraph

## Slide 20: Summary Table
Performance Metric | Required Range | Use Case(s) | BSRH
Contrast | 10^-4 to 10^-6 | Collimation: 10^-4 at upper bound MP: 10^-6 at 6.7σ | Currently unknown
Relative Integral | 0.2% to 5% | Collimation: 5% to 0.5% MP: 1.4% to 0.2% | <0.1% (relative)
Absolute Integral | 10^10 to 10^12 | Collimation: 1.5×10^12 MP: (1-4)×10^10 | Currently unknown (~1.5% during MD)
1D Profile Capability | Yes/No | Required for Beam-Beam | Yes
2D Image Capability | Yes/No | Required for Beam-Beam | Yes
Max. Acquisition Rate | 10-60 seconds | MP: ~10s Others: ~60s | ~20Hz (time for full beam depends on the gated sample - >100s for full machine bunch per bunch, around 3s for 48b., limited by data processing)
Bunch-by-Bunch Gating | Yes/No | Required for Beam-Beam | Possible
Number of Turns Needed | - | All cases accept multi-turn | Order of 100 turns (depends on gated sample)
Interlock Capability | Yes/No | Required for MP | Warning possible, Software interlock?

## Slide 21: Final words
Important measurements during 2 MDs
A lot of SW work in the future
Hopefully in 2025 first halo data of physics run
  Currently only relative
  Through additional MDs we might lower the uncertainty and prove possible deconvolution of the measured profile/image and the PSF

## Slide 22: Thank you for your attention

## Slide 23: Backup

## Slide 24: BSRH intensity problem
±10% intensity oscillation due to finite capacitance of the intensifier – change of intensifier model

## Slide 25: Lyot’s stop - data
Lyot’s stop closed
Lyot’s stop opened