## Slide 1: BGC: development in LHC and implementation study for LS3
D. Butti on behalf of the BGC collaboration

## Slide 2
beam
gas jet
B1
B2
RF section
UND
D3
D3
UND
Point 4
Beam Gas Curtain principle

## Slide 3: Beam Gas Curtain as beam imaging system
proton beam
gas jet
detector
fluorescence photons
BGC animation on YouTube
Conceived as overlap monitor for the Hollow Electron Lens, reproposed as beam size monitor:
minimally-invasive
simple beam imaging
weak fluorescence signal, only suitable for avg measurements
Full cycle with protons (H only)
Online measurements with ions (both H/V)
Interesting emittance measurements in run 2024…
…can it be useful for halo measurements?

## Slide 4: Tail information in BGC images
log scale
HALO
single frame - 50 ms
integrated frame - 30 s
Contrast driven by signal-to-noise ratio
values ~2% peak-to-tail in present BGC configuration
lin scale
~2% peak-to-tail
Spatial resolution driven by system’s optical resolution
measured consistent throughout the year
Dynamic range not an issue
image build as (slow) sum of single photon frames

## Slide 5
Improving image contrast
lin scale
log scale
Longer exposure times increase beam and background signals
less noisy profiles but no exceptional contrast gain
~1%
Some options available to improve contrast
increase light collection efficiency → under study, may gain factor 4
increase jet density/active volume → under study
reduce dark counts and light background → already pushed, little margin
reduce losses background → requires dedicated shielding
advanced denoising techniques → to be studied
use alternative gas species (e.g. N2) → gain/feasibility to be tested
Margin to reach stable ~1% contrast with exposure ~1 min, feasibility of exceeding this requires further investigation and testing

## Slide 6
Summary of BGC imaging
Contrast: ~1% likely feasible in halo, margin for improvement but feasibility to be studied
Halo integral: requires experiments to assess sensitivity to actual halo population change (e.g. EoF scraping)
1D/2D information: likely 1D (jet much thinner than beam size needed for 2D beam image)
Acquisition rate: slow ~1 min, average measurements only
Interlock: only software (image processing required)
Full cycle: mildly affected by beam energy
Ad-hoc optics: larger size improves spatial resolution, not too much (i.e. < 2σ) to avoid signal dilution
Take-home message for BGC imaging:
limited performance in acquisition rate and contrast
robust and reproducible thanks to its simplicity
→ insufficient as single halo monitor, might provide a solid cross-reference for other monitors

## Slide 7
Fluorescence is not the only signal produced by the BGC, jet-induced losses clearly detectable downstream
BSRTM
MBRS
MBRB
BGC
FLUKA simulation of BGC induced dose by D. Prelipcean
Losses baseline subtracted and normalized by standard deviation
Beyond fluorescence: losses-based diagnostics
jet activated

## Slide 8: Concept of BGC as beam profiler
System can be converted into a beam profiler
use jet as “gas collimator” and detect losses
losses
BLM
jet
beam
camera
jet
beam
Present BGC configuration as beam imaging system
jet is “gas screen” and detect fluorescence photons
info about 2D beam distribution
low fluorescence cross-section, limited halo performance
intrinsically a 1D instrument
measures integrated population within jet, profile obtained by scanning jet position
photons

## Slide 9: Concept of BGC as beam profiler
Measurement principle
full transverse scan of the beam. Beam distribution obtained by deconvolution of losses signal and jet profile
take background reference with jet away from beam
approach jet to beam to monitor evolution of tail population
Scan repeated whenever full-profile is requested and to adjust jet position
Jet used as a probe to scan the beam
Background subtraction can be regularly adjusted by displacing jet but very high sensitivity required to resolve halo signal

## Slide 10: Gas curtain as “interceptive” device
Matter particles intercepted by the beam (mol.m-2)
Chart: 
| 100000000000000 | 250000000000000 | 3.5000000000000002E+24 | 2.2000000000000001E+25
Atomic number | 4 | 4 | 12 | 102
Molecular mass of interceptor
Quench, material strength and handling limits
Stolen from R. Veness through G. Aliana Cervera

## Slide 11: Gas distribution: jet
Real jet profile not an ideal rectangle but has finite edge sharpness and pressure gap
Pressure gap can be estimated from fluorescence emission
Pb @ INJ
Finite background pressure → parasitic losses from core
Edge sharpness measured by movable gauge
Finite edge sharpness → ultimate limit for spatial resolution

## Slide 12: Gas distribution: background
Gas injected by BGC not fully confined within to jet but spans over several meters
At present, a fraction of the gas load belongs to the jet while the majority remains as background
A fraction of the losses is the useful signal produced by the jet
beam
Fraction depends on detection location (e.g. higher jet sensitivity closer to jet)

## Slide 13: Contrast of BGC as beam profiler
How to quantify the contrast of the BGC profiler?
Basic model where signal depends only on geometric overlap between jet and halo
Proportionality between contrast to integrated population and sensitivity to losses, with jet fraction as modulating function
losses
no jet = no signal
only jet = no bkg
region of interest

## Slide 14: Improving contrast
BLM
BLM
BLM
BGC
Credit: D. Prelipcean
Options available to improve losses detection (vs background)
optimize detection location → dedicated IC BLMs installed in YETS2425
maximize detection sensitivity → best detector technology to be studied
BWS
BLM
BLM
Optimizing contrast requires maximizing production and detection of losses created by the jet
Options available to improve jet losses (vs background)
optimize jet aspect ratio for maximum jet-beam overlap → under study
increase gas pressure in jet → to be studied, limited margin
increase pumping to reduce background → under study (helpful also to reduce gas load)

## Slide 15
Summary of BGC profiler
Take-home message for BGC profiler:
using BGC as beam profiler is a very recent concept
BGC profiler configuration very similar to existing one, already validated in the machine
→ instrument has potential, but further investigations needed (before LS3) to quantify performance
Contrast (integral): to be quantified depending on achievable sensitivity to jet losses
1D/2D information: intrinsically 1D
Acquisition rate: potentially bunch-by-bunch, integration time depends on achieved sensitivity to jet losses
Interlock: potentially HW, if interlock on simple loss threshold is reliable
Full cycle: not affected by beam energy
Ad-hoc optics: larger size improves spatial resolution, not too much (i.e. < 2σ) to avoid signal dilution

## Slide 16: Timeline for BGC development
BGC imaging
BGC profiler
B1 prototype commissioned and preliminary measurements
study of emittance monitoring performance
reduce jet thickness to increase resolution
regular operation and study of final HL design
change gas species
move towards operational instrument
upgrade of B1 prototype
installation of B2 instrument
(details in Gerhard’s talk)
today
concept proposed and preliminary investigations
dedicated BLMs installation
studies on losses and HL design
rotate jet to profiler configuration
proof of principle in dedicated MD
installation of required 1 instrument / plane
(details in Gerhard’s talk)
done
planned
tentative / optional

## Slide 17: Conclusion
Beam Gas Curtain is an R&D instrument with demonstrated potential as average emittance monitor
Two candidate configurations for halo monitoring
Beam imaging based on fluorescence
principle and implementation validated
requires optimisation for halo measurements
→ known limitations but little uncertainty on implementation by HL-era
Beam profiler based on losses
new concept, proof of principle possible before LS3
performance and limitation still to be assessed
→ principle more suitable for halo measurement but exploitation is more challenging
Tests and studies planned for next year will provide estimate of ultimate performance of both configurations
Primarily work as emittance monitors, halo measurements would come as additional feature

## Slide 18: Asked AI for ideas about this review…
Looks like there are simpler solutions…

## Slide 19: Thank you for your attention!