## Slide 1: EM stripline developments status and plans
Michal Krupa SY-BI-BP / WP13
HF BPM review, 15/01/2025

## Slide 2: Stripline BPMs
Long electrodes with a connector on each end
  For head-tail (HT): one end for signal generation, the other end loaded with 50 Ω
Current HT – elongated version of the standard LHC stripline BPM
50 Ω
load
Signal output
SiO2 cables
RF hybrid
Long coax to a scope

## Slide 3: Stripline BPM for intra-bunch measurements
Perfect stripline generates an electric signal which is a replica of the bunch current (+ reflection) which is modulated by the intra-bunch position
  Commonly used at CERN and other facilities as an HT monitor
  Gating on the positive signal only for HT measurements
  “Standard” electrode length to resolve LHC bunches – 400 mm
In real (imperfect) world the signal is distorted due to the electrode impedance variations, and other reflections and BW limitations in cables and electronics
The same longitudinal profile
as the bunch (for an on-axis bunch)
Mirrored reflection
Time delay 2 x electrode length / speed of light

## Slide 4: Can we reuse the existing LHC HT BPM design?
End flanges, aperture, body length, electrode length are appropriate for the HL-LHC BPW
Three issues:
  Currently used LHC HT BPMs are single plane only – design update would be needed to accommodate the request for a double-plane BPW
  Non-optimal electrode design: very complex mechanical connection to the BPM body and poor impedance control
  Limited availability of spare FTs
Electrode
FT
Further improved
TDR up to 20 GHz
Change of N connector grade and pin dimensions to compensate for decrease of ε during thermal treatment
LHC stripline
HL-LHC stripline

## Slide 5: Elongated HL-like stripline
Same approach as the LHC HT pick-ups: use the latest stripline BPM design but make the electrode longer (400 mm)
  Electrode and feedthroughs already optimized for 6 GHz BW, measured up to 20 - 26.5 GHz
  Machining techniques already developed and demonstrated
  Good reproducibility achieved after a few design iterations
  The same approach is used for the BPTQR

## Slide 6: Signal levels
Stripline BPMs generate high-amplitude signals exceeding 100 Vpp in the LHC
~ 15 W per BPM output with full beam
  Can get much worse for off-centred beams
This power must be taken by the front-end electronics (e.g. hybrid, attenuators)
2.3e11 ppb

## Slide 7: Expected performance
BPM thermal noise (10 GHz BW) with a pilot bunch: ~ 1 μm
  Resolution will be (most likely) limited by the front-end (e.g. hybrid) and acquisition electronics (e.g. scope)
H-9 hybrid, used in the existing LHC HT, is the best available on the market
  BW: 2 GHz; Power: 25 W; Isolation: min. 30 dB
  Collaboration with Oxford University to develop a better hybrid – many prototypes not exceeding the H-9 performance
  Most likely candidate for “classical” stripline BPW – no performance improvement
    Alternative solution in the next talk
Better digitizers might be available in 5 years to improve on the existing HT resolution / sampling
200 μm
20 μm
65 μm

## Slide 8: Lab RF measurements
RF measurements performed for the cryogenic HL-LHC striplines:
  Electrode impedance
  Electrode symmetry
  Electrode cross-coupling
  Electrical center w.r.t. the mechanical fiducials
  Beam position sensitivity
  Beam position nonlinearity
The same measurements are possible for a stripline BPW – all measurement equipment and tooling already exists or can be easily adapted

## Slide 9: Compatibility with environmental constraints
Mechanics:
  Preliminary design compatible with the available space done
  No known alignment constraints
Bakeout:
  BPM, feedthroughs and SiO2 cables compatible with the LHC bakeout requirements
  50 Ω loads will need to be removed before and reinstalled after each bakeout (5 min per BPM)
Radiation:
  BPM, feedthroughs and SiO2 cables are rad-hard up to 10’s MGy (needed for the IP1/IP5 IT)
  50 Ω loads and hybrids might eventually suffer from radiation damage – replacement parts must be available
Accessibility:
  Preliminary design giving access to all BPM connectors done

## Slide 10: Proposed strategy and milestones
Study:
  RF measurements of the existing LHC HT BPMs, hybrids and cables [Q1’25]
  Estimation of performance limitation due to each component [Q2’25]
Design:
  Decision on the necessity of the intermediate flange [Q2’25]
  Final BPTQR design + stripline BPW design [Q2-Q4’25]
  Final alignment and support system design [Q3-Q4’25]
  Production of manufacturing drawings [Q1’26]
Production:
  Procurement of RF FTs and SiO2 cables [done]
  Procurement of raw material for BPTQR and stripline BPW [Q1’25]
  Production of two prototype elongated HL-like electrodes [Q2-Q3’25]
  Production of a pre-series alignment platform [Q2’26]
  Production of two pre-series BPTQR + BPW assemblies [Q2’26-Q4’26]
  Series production of alignment platforms, BPTQR’s and BPW’s [Q2’27-Q4’27]
Testing:
  Metrology and lab RF measurements of prototype electrodes [Q3-Q4’25]
  Alignment tests with the pre-series alignment platform and pick-ups [Q4’26-Q1’27]
  Lab RF measurements (and at CLEAR?) of pre-series BPTQR’s and BPW’s [Q1’27]
Giusy Passarelli [DOCT]
Loann Havart [TECH]
EN-MME job already opened
EN-MME or external, TBC
SY-BI

## Slide 11: Thank you for your attention