## Slide 1: ATLAS Gas System Status 2024/25
Mattia Busato
10/02/25

## Slide 2: Outlook
Piquet intervention statistics
General gas system maintenance
2024 activities
2025 activities

## Slide 3: Piquet intervention statistics
Chart: 71 interventions in 2024
| RPC | TGC | TFC | TRT | MDT | MMG | PGS
 | 29 | 6 | 7 | 2 | 4 | 16 | 7
RPC many interventions due to humidifier malfunction and/or humidity level out of range
MMG many interventions due to Profibus problem
TFC few interventions due to frequency converter of the pump malfunctioning
PGS interventions will be on BE-EA

## Slide 4: General gas systems maintenance
Filters replacement and cleaning
Pump membranes replacement
Mass flow controller recalibration
Safety valves recalibration
Analysis devices calibration
Analysis pumps replaced
Humidifier water tanks refilled

## Slide 5: Activities 2024
RPC flowcell replacement campaign  400+ flowcells recalibrated for the 9 distribution racks.
TGC flowcell recalibration  100+ flowcells in rack 63 and 67.
MDT motherboards and ELMBs replacement  10 distribution racks
RPC purifiers maintenance  material replacement and columns clogging investigation
TFC pump replacement

## Slide 6: Activities 2025
MDT distribution flow regulation
RPC distribution high flow tests
ANUBIS gas flow increase
Requests for 2025 EYETS ?