## Slide 1: BOC RF Measurement
16.04.2025
Pablo Martinez Reviriego; Paz Alonso Arias; Alexej Grudiev

## Slide 2: Measurement Setup – VNA settings
Measurement 2:
Central freq: 11.994 GHz
Span 10 MHz
Spectral resolution: 500 Hz
No. of points: 20001
IF Bandwidth: 1 kHz
Measurement 1:
Central freq: 11.994 GHz
Span 5 MHz
Spectral resolution: 250 Hz
No. of points: 20001
IF Bandwidth: 1 kHz

## Slide 3
S-parameters
Measurement 1
Measurement 2

## Slide 4
1 cavity model
Parameter | Value
 | 11.994 GHz
 | 2.35e5
 | 6.6
 | 35606
Parameter | Value
 | 11.994 GHz
 | 2.35e5
 | 6.6
 | 35606

## Slide 5
Measurement fit
 | Fit |  | Smith chart |  | Designed
 | Offset | Amp | Offset | Amp | 
 | 11.9943 | 11.9943 | 11.9943 | 11.9943 | 11.994
 | 180655 | 192447 | 159923 | 175099 | 2.36e5
 | 43431 | 47874 | 45091 | 43975 | 3.58e4
 | 4.16 | 4.02 | 3.55 | 3.98 | 6.6
 | Fit |  | Smith chart |  | Designed
 | Offset | Amp | Offset | Amp | 
 | 11.9943 | 11.9943 | 11.9943 | 11.9943 | 11.994
 | 180655 | 192447 | 159923 | 175099 | 2.36e5
 | 43431 | 47874 | 45091 | 43975 | 3.58e4
 | 4.16 | 4.02 | 3.55 | 3.98 | 6.6
Offset model
Amplitude model
S21
S21

## Slide 6
Measurement comparison
 | 1 | 2 | Designed
 | 11.9943 | 11.9940 | 11.994
 | 180655 | 195177 | 2.36e5
 | 43431 | 44487 | 3.58e4
 | 4.16 | 4.39 | 6.6
 | 5.4e-4 | 2.9e-4 | 
 | 1 | 2 | Designed
 | 11.9943 | 11.9940 | 11.994
 | 180655 | 195177 | 2.36e5
 | 43431 | 44487 | 3.58e4
 | 4.16 | 4.39 | 6.6
 | 5.4e-4 | 2.9e-4 | 
1
2
S21
S21

## Slide 7
General model
Parameter | Value
 | 11.994 GHz
 | 2.35e5
 | 6.6
 | 35606
Parameter | Value
 | 11.994 GHz
 | 2.35e5
 | 6.6
 | 35606

## Slide 8
Simulation

## Slide 9
Simulation

## Slide 10
4-params fit
 | Fit | Designed
 | 163989 | 2.36e5
 | 44426 | 3.58e4
 | 3.69 | 6.6
[GHz] | 11.9943 | 11.994
[kHz] | 14.8 | 0
 | 7.7e-4 | 0
 | Fit | Designed
 | 163989 | 2.36e5
 | 44426 | 3.58e4
 | 3.69 | 6.6
 | 11.9943 | 11.994
 | 14.8 | 0
 | 7.7e-4 | 0

## Slide 11
5-params fit
 | Fit | Designed
 | 166211 | 2.36e5
 | 41231 | 3.58e4
 | 4.03 | 6.6
 | 48352 | 3.58e4
 | 3.44 | 6.6
[GHz] | 11.9943 | 11.994
[kHz] | 16.7 | 0
 | 1.86e-4 | 0
 | Fit | Designed
 | 166211 | 2.36e5
 | 41231 | 3.58e4
 | 4.03 | 6.6
 | 48352 | 3.58e4
 | 3.44 | 6.6
 | 11.9943 | 11.994
 | 16.7 | 0
 | 1.86e-4 | 0

## Slide 12
6-params fit
 | Fit | Designed
 | 209795 | 2.36e5
 | 136702 | 2.36e5
 | 41335 | 3.58e4
 | 5.07 | 6.6
 | 48213 | 3.58e4
 | 2.83 | 6.6
[GHz] | 11.9943 | 11.994
[kHz] | 9.6 | 0
 | 3.9e-6 | 0
 | Fit | Designed
 | 209795 | 2.36e5
 | 136702 | 2.36e5
 | 41335 | 3.58e4
 | 5.07 | 6.6
 | 48213 | 3.58e4
 | 2.83 | 6.6
 | 11.9943 | 11.994
 | 9.6 | 0
 | 3.9e-6 | 0

## Slide 13
6-params fit S22
 | Fit | Designed
 | 171994 | 2.36e5
 | 159267 | 2.36e5
 | 43448 | 3.58e4
 | 3.96 | 6.6
 | 45654 | 3.58e4
 | 3.49 | 6.6
[GHz] | 11.9943 | 11.994
[kHz] | 13.4 | 0
 | 4.9e-6 | 0
 | Fit | Designed
 | 171994 | 2.36e5
 | 159267 | 2.36e5
 | 43448 | 3.58e4
 | 3.96 | 6.6
 | 45654 | 3.58e4
 | 3.49 | 6.6
 | 11.9943 | 11.994
 | 13.4 | 0
 | 4.9e-6 | 0

## Slide 14
Measurement 1: Summary
 | 6-params fit (S22) | 6-params fit | 5-params fit | 4-params fit | 3-params fit | Designed
 | 171994 | 209795 | 166211 | 163989 | 175099 | 2.36e5
 | 159267 | 136702 |  |  |  | 
 | 43448 | 41335 | 41231 | 44426 | 43975 | 3.58e4
 | 45654 | 48213 | 48352 |  |  | 
 | 3.96 | 5.07 | 4.03 | 3.69 | 3.98 | 6.6
 | 3.49 | 2.83 | 3.44 |  |  | 
[GHz] | 11.9943 | 11.9943 | 11.9943 | 11.9943 | 11.9943 | 11.994
[kHz] | 13.4 | 9.6 | 16.7 | 14.8 | 0 | 0
 | 4.9e-6 | 3.9e-6 | 1.86e-4 | 7.7e-4 | 2.2e-3 | 
 | 6-params fit (S22) | 6-params fit | 5-params fit | 4-params fit | 3-params fit | Designed
 | 171994 | 209795 | 166211 | 163989 | 175099 | 2.36e5
 | 159267 | 136702 |  |  |  | 
 | 43448 | 41335 | 41231 | 44426 | 43975 | 3.58e4
 | 45654 | 48213 | 48352 |  |  | 
 | 3.96 | 5.07 | 4.03 | 3.69 | 3.98 | 6.6
 | 3.49 | 2.83 | 3.44 |  |  | 
 | 11.9943 | 11.9943 | 11.9943 | 11.9943 | 11.9943 | 11.994
 | 13.4 | 9.6 | 16.7 | 14.8 | 0 | 0
 | 4.9e-6 | 3.9e-6 | 1.86e-4 | 7.7e-4 | 2.2e-3 | 

## Slide 15
4-params fit
 | Fit | Designed
 | 182190 | 2.36e5
 | 45133 | 3.58e4
 | 4.0 | 6.6
[GHz] | 11.9940 | 11.994
[kHz] | 9.3 | 0
 | 3.9e-4 | 0
 | Fit | Designed
 | 182190 | 2.36e5
 | 45133 | 3.58e4
 | 4.0 | 6.6
 | 11.9940 | 11.994
 | 9.3 | 0
 | 3.9e-4 | 0

## Slide 16
5-params fit
 | Fit | Designed
 | 184693 | 2.36e5
 | 41933 | 3.58e4
 | 4.4 | 6.6
 | 48940 | 3.58e4
 | 3.8 | 6.6
[GHz] | 11.9940 | 11.994
[kHz] | 8.1 | 0
 | 8e-5 | 0
 | Fit | Designed
 | 184693 | 2.36e5
 | 41933 | 3.58e4
 | 4.4 | 6.6
 | 48940 | 3.58e4
 | 3.8 | 6.6
 | 11.9940 | 11.994
 | 8.1 | 0
 | 8e-5 | 0

## Slide 17
6-params fit
 | Fit | Designed
 | 199522 | 2.36e5
 | 172207 | 2.36e5
 | 42040 | 3.58e4
 | 4.7 | 6.6
 | 48791 | 3.58e4
 | 3.5 | 6.6
[GHz] | 11.9940 | 11.994
[kHz] | 7.5 | 0
 | 6.8e-5 | 0
 | Fit | Designed
 | 199522 | 2.36e5
 | 172207 | 2.36e5
 | 42040 | 3.58e4
 | 4.7 | 6.6
 | 48791 | 3.58e4
 | 3.5 | 6.6
 | 11.9940 | 11.994
 | 7.5 | 0
 | 6.8e-5 | 0

## Slide 18
6-params fit S22
 | Fit | Designed
 | 194156 | 2.36e5
 | 172911 | 2.36e5
 | 44210 | 3.58e4
 | 4.39 | 6.6
 | 46200 | 3.58e4
 | 3.74 | 6.6
[GHz] | 11.9940 | 11.994
[kHz] | 1.8 | 0
 | 6.6e-5 | 0
 | Fit | Designed
 | 194156 | 2.36e5
 | 172911 | 2.36e5
 | 44210 | 3.58e4
 | 4.39 | 6.6
 | 46200 | 3.58e4
 | 3.74 | 6.6
 | 11.9940 | 11.994
 | 1.8 | 0
 | 6.6e-5 | 0

## Slide 19
Measurement 2: Summary
 | 6-params fit S22 | 6-params fit | 5-params fit | 4-params fit | 3-params fit | Designed
 | 194156 | 199522 | 184693 | 182190 | 195177 | 2.36e5
 | 172911 | 172207 |  |  |  | 
 | 44210 | 42040 | 41933 | 45133 | 44487 | 3.58e4
 | 46200 | 48791 | 48940 |  |  | 
 | 4.39 | 4.7 | 4.4 | 4.04 | 4.39 | 6.6
 | 3.74 | 3.5 | 3.8 |  |  | 
[GHz] | 11.9940 | 11.9940 | 11.9940 | 11.9940 | 11.9940 | 11.994
[kHz] | 1.8 | 7.5 | 8.1 | 9.3 | 0 | 0
 | 6.6e-5 | 6.8e-5 | 8e-5 | 3.9e-4 | 1.4e-3 | 
 | 6-params fit S22 | 6-params fit | 5-params fit | 4-params fit | 3-params fit | Designed
 | 194156 | 199522 | 184693 | 182190 | 195177 | 2.36e5
 | 172911 | 172207 |  |  |  | 
 | 44210 | 42040 | 41933 | 45133 | 44487 | 3.58e4
 | 46200 | 48791 | 48940 |  |  | 
 | 4.39 | 4.7 | 4.4 | 4.04 | 4.39 | 6.6
 | 3.74 | 3.5 | 3.8 |  |  | 
 | 11.9940 | 11.9940 | 11.9940 | 11.9940 | 11.9940 | 11.994
 | 1.8 | 7.5 | 8.1 | 9.3 | 0 | 0
 | 6.6e-5 | 6.8e-5 | 8e-5 | 3.9e-4 | 1.4e-3 | 

## Slide 20
Summary
 | Measurement 1 |  | Measurement 2 | 
 | 6-params fit S22 | 6-params fit | 6-params fit S22 | 6-params fit
 | 171994 | 209795 | 194156 | 199522
 | 159267 | 136702 | 172911 | 172207
 | 43448 | 41335 | 44210 | 42040
 | 45654 | 48213 | 46200 | 48791
 | 3.96 | 5.07 | 4.39 | 4.7
 | 3.49 | 2.83 | 3.74 | 3.5
[GHz] | 11.9943 | 11.9943 | 11.9940 | 11.9940
[kHz] | 13.4 | 9.6 | 1.8 | 7.5
 | 4.9e-6 | 3.9e-6 | 6.6e-5 | 6.8e-5
 | Measurement 1 |  | Measurement 2 | 
 | 6-params fit S22 | 6-params fit | 6-params fit S22 | 6-params fit
 | 171994 | 209795 | 194156 | 199522
 | 159267 | 136702 | 172911 | 172207
 | 43448 | 41335 | 44210 | 42040
 | 45654 | 48213 | 46200 | 48791
 | 3.96 | 5.07 | 4.39 | 4.7
 | 3.49 | 2.83 | 3.74 | 3.5
 | 11.9943 | 11.9943 | 11.9940 | 11.9940
 | 13.4 | 9.6 | 1.8 | 7.5
 | 4.9e-6 | 3.9e-6 | 6.6e-5 | 6.8e-5
1
2