## Slide 1: X-Boxes Update
01.10.2025

## Slide 2: iFast structure – Low Power Measurements
S parameters measurement
300 MHz span
11.994 GHz center frequency
No. points 10001
IF Bandwidth 1 kHz
4 ports full-calibration
No vacuum
Port 1, Output
Port 2, Input

## Slide 3: iFast structure – Low Power Measurements
S parameters measurement
300 MHz span
11.994 GHz center frequency
No. points 10001
IF Bandwidth 1 kHz
4 ports full-calibration
Including air correction

## Slide 4: iFast structure – Low Power Measurements
Bead pull
CW mode
No. points 8001
IF Bandwidth 100 Hz
Time sweep 65 s

## Slide 5: iFast structure – Low Power Measurements
Bead pull @11.994GHz
11.994 GHz CW mode
No. points 8001
IF Bandwidth 100 Hz
Time sweep 65 s
Plateau probably due to motor stuck
Number of peaks: 113= 111 cells + 2 coupler cells?

## Slide 6: XBOXES Summary
No superconductivity (faulty cryocooler)
Changed cryocooler filters, still not working
S. Calatroni provided the contact of the provider. Asked for repair options
Waiting for klystron (in Italy)
Cryocooler sent for repair
Started pulsing 13th June
150 Hz in each line
50ns RF pulse, no compression
Setpoint DUT 8MW in both lines
Compressed pulse 50ns
RF pulse 2us
Setpoint DUT 2MW (limited by Pulse compressor)
XBOX2 (Not operative)
XBOX3 (Operative)

## Slide 7: XBOX3 Configuration
SSA A
SSA B
KLY A
KLY B
PC 1
HYB
PC 2
DUT 2
DUT 1
Load 1
Load 2
KLYINA
KLYINB
PKIB
PKIA
PKRA
PKRB
PLIA
PLRA
PSIA
PSRA
PEIA
PERA
PLIB
PLRB
PSIB
PSRB
PEIB
PERB
DCUP
DCDOWN
Bunker
HPDC
HPDC
DC
DC
HPDC
HPDC
HPDC
HPDC
HPDC
HPDC
Faraday Cups

## Slide 8: XBOX3 Configuration – Line 2
SSA A
SSA B
KLY A
KLY B
PC 1
HYB
Old PC
TD31N1
DUT 1
Load 1
Load
KLYINA
KLYINB
PKIB
PKIA
PKRA
PKRB
PLIA
PLRA
PSIA
PSRA
PEIA
PERA
PLIB
PLRB
PSIB
PSRB
PEIB
PERB
DCUP
DCDOWN
Bunker
HPDC
HPDC
DC
DC
HPDC
HPDC
HPDC
HPDC
HPDC
HPDC
Faraday Cups

## Slide 9: XBOX3 Configuration – HEK klystron

## Slide 10: XBOXES TO DO list
Finishing paper on High power tests
Knowledge transfer to next Grad, starting January
Writing documentation on Xboxes (guidelines)
Calibration of RF lines before leaving (at least PSI signals)
Changing a PXI card in XBOX3 (acquisition of pulse compressor chillers information) and making sure everything works
Not breaking a klystron during my last month

## Slide 12: XBOX3 Configuration
SSA A
SSA B
KLY A
KLY B
PC 1
HYB
PC 2
DUT 2
DUT 1
Load 1
Load 2
KLYINA
KLYINB
PKIB
PKIA
PKRA
PKRB
PLIA
PLRA
PSIA
PSRA
PEIA
PERA
PLIB
PLRB
PSIB
PSRB
PEIB
PERB
DCUP
DCDOWN
Bunker
HPDC
HPDC
DC
DC
HPDC
HPDC
HPDC
HPDC
HPDC
HPDC
Faraday Cups

## Slide 13: XBOX3 Configuration
SSA A
SSA B
E37117
E37123
PC 1
HYB
PC 2
DUT 2
DUT 1
Load 1
Load 2
KLYINA
KLYINB
PKIB
PKIA
PKRA
PKRB
PLIA
PLRA
PSIA
PSRA
PEIA
PERA
PLIB
PLRB
PSIB
PSRB
PEIB
PERB
DCUP
DCDOWN
Bunker
HPDC
HPDC
DC
DC
HPDC
HPDC
HPDC
HPDC
HPDC
HPDC
Faraday Cups

## Slide 14: X-Box 2: Reminder on BD classification
BD_Flag_PER_log
BD_Flag_DC_DOWN
BD_Flag_DC_UP
BD_Flag_PSR_log
BD_Flag_PKR_log
BD Load
BD DUT w. BDC
BD DUT wo. BDC
BD HyPC
No BD
DC Up >=140
DC Down>=100
TRUE
FALSE