## Slide 1: Internal monitoring for Crab-cavities
Vivien RUDE 2024-10-14
HL-LHC Collaboration Triumf-STFC-CERN

## Slide 2
From Teddy Capelli

## Slide 3
 | Step | 
Part 0 : CMM data | 0 | Validation of trolley
 |  | Position of the capacitive plates w.r.t. external references  CERN
 |  | Valve plate measurement
Part 1 : in clean room | 1-2 | Alignment in ISO5 (before connection in ISO4)
 |  | Alignment in ISO5 (after connection in ISO4)
Part 2 : before cryostating | 4-5 | Alignment of different equipment
 |  | Installation of FSI supports
 |  | Measurement of FSI supports
 |  | Installation FSI targets
Part 3 : cryostating (top plate) | 6 | Alignment of the cavities before cryostating
 |  | Measurement of the top plate cryomodule
 |  | Alignment trolley and top plate
Part 4 : cryostating (OVC cryomodule) | 10 | Measurement of the OVC cryomodule
 |  | Alignment top plate and OVC cryomodule
Part 5 : after cryostating | 10 | Installation of FSI heads on OVC and top plate
 |  | Measurement of FSI heads on OVC and top plate
 |  | FSI validation (comparison with a laser tracker)
 |  | Alignment with adjustment system
Part 6 : cold test at STFC or Triumf | 12 | Measurement ambient pressure
 |  | Measurement under vacuum
 |  | Measurement at cold
Part 7 : cold test at CERN |  | Measurement ambient pressure
 |  | Measurement under vacuum
 |  | Measurement at cold

## Slide 4
General principle : 3 objects  3 coordinate systems similar
Top plate
OVC
Trolley
X
Y
Z

## Slide 5
 | Step | 
Part 0 : CMM data | 0 | Validation of trolley
 |  | Position of the capacitive plates w.r.t. external references  CERN
 |  | Valve plate measurement

## Slide 6
 | Step | 
Part 0 : CMM data | 0 | Validation of trolley
 |  | Position of the capacitive plates w.r.t. external references  CERN
 |  | Valve plate measurement
IN
V3 mm
V1 mm
X
Y
Z
V2 mm
Coordinate frame :
Primary axis : Nominal axis (cavity line)  trajectory of the rails
Secondary axis : Normal vector to vertical plane of rails
Origin : Beginning of the rail
Advice :
Install fiducial points on both sides of the trolley.
Test the trajectory of the component on the rails
Test the impact of load on the trolley

## Slide 7
 | Step | 
Part 0 : CMM data | 0 | Validation of trolley
 |  | Position of the capacitive plates w.r.t. external references  CERN
 |  | Valve plate measurement

## Slide 8
 | Step | 
Part 0 : CMM data | 0 | Validation of trolley
 |  | Position of the capacitive plates w.r.t. external references  CERN
 |  | Valve plate measurement
Z
Y
X
Flange B
FPC
HOM-A
Final coordinate system (R-capacitive-plates):
Y-axis (primary axis): Defined by the capacitive plate axis
Z-axis (secondary axis): Defined from the projection of the center of circle FPC on line of the capacitive axis
Origin: projection of the center of circle FPC on line of the capacitive axis.
C1
C3
C2
C4
C5
C6
C7
C8
C9
C10
 | X (mm) | Y (mm) | Z (mm)
Flange_A_Inox | 0 | -500.857 | -0.039
Flange_B_Inox | 0 | 179.144 | -0.254
Flange_FPC_Inox | 0 | 0 | 234.834
Flange_HOM_A_Inox | 0.381 | -320.512 | 249.645
Flange_HOM_B_Inox | 89.659 | -274.265 | -250.260
Flange_HOM_C_Inox | -90.249 | -273.956 | -249.957
Flange_pick_up_A_Inox | 239.994 | -385.315 | -0.423
Flange_pick_up_B_Inox | 240.000 | 65.110 | -0.137
Line2 A | -194.522 | -508.330 | 0.298
Line2 B | -194.489 | 187.668 | 0.335
C1 | 121.217 | -0.148 | 214.856
C2 | 156.375 | -304.377 | 214.999
C3 | -184.520 | -410.281 | 215.239
C4 | 223.532 | -342.117 | 134.529
C5 | 223.371 | 47.401 | 134.695
C6 | 223.138 | -296.469 | -135.616
C7 | 223.011 | -43.997 | -135.468
C8 | 142.837 | 166.124 | 82.412
C9 | -41.267 | 166.172 | -144.759
C10 | -142.858 | 166.060 | 132.874
C11 | -305.115 | 30.952 | 135.617
C12 | -304.967 | -374.040 | 135.433
C13 | -305.525 | 69.491 | -109.469
C14 | -305.465 | -316.935 | -134.642
C15 | -41.117 | -486.970 | -145.082
C16 | -142.617 | -487.408 | 132.577
C17 | 143.050 | -487.184 | 82.045
Tunning_BOTTOM | -0.947 | -160.451 | -369.724
Tunning_TOP | 0.599 | -160.218 | 369.329
Capacitive_plate_CENTER | 0 | -160.238 | 0
Capacitive_plate_IN | 0 | -250.538 | 0
Capacitive_plate_OUT | 0 | -70.099 | 0
Origine | 0 | 0 | 0
Axe_X | 1000 | 0 | 0
Axe_Y | 0 | 1000 | 0
Axe_Z | 0 | 0 | 1000
IN | 0 | -500.857 | 0
OUT | 0 | 179.144 | 0
Flange A
Line2 A
HOM-A
FPC
Z
Y
C11
C12
C13
C14
C15
C16
C17
C3
C2
C1

## Slide 9
 | Step | 
Part 0 : CMM data | 0 | Validation of trolley
 |  | Position of the capacitive plates w.r.t. external references  CERN
 |  | Valve plate measurement
Valve Plate A
DQW_VP_A1
DQW_VP_A2
DQW_VP_A4
DQW_VP_A3
DQW_VP_B1
DQW_VP_B2
DQW_VP_B3
DQW_VP_B4
Valve Plate A
Valve Plate B
X
Y
Z
X
Y
Z
R_ValvePlate_A
R_ValvePlate_B
Coordinate system
For both Valve Plates,
X axis – along the line passing through the centres of the two cylinders (Secondary Axis)
Y axis – Normal to the measured 7 planes
Z axis – Perpendicular to XY Plane (Primary Axis)

## Slide 10
Valve plate A
Valve plate B
Cavity 2
Cavity 1
IN
OUT
IN
OUT
V3 mm
V1 mm
Trolley axis
B1
B2
 |  | X |  | Y |  | Z |  | Rx |  | Ry |  | Rz | 
Valve plate 1 | A1 | 0 | +/- 0.5 | -- | -- | 0 | +/- 0.5 | -- | -- | -- | -- | -- | --
 | A2 | -- | -- | -- | -- | 0 | +/- 0.5 | -- | -- | -- | -- | -- | --
 | Vect Na | -- | -- | N | N | -- | -- | 0 | +/-3.8 mrad | -- | -- | 0 | +/-2.5 mrad
Cav 1 | IN (Cap) | 0 | +/- 0.1 | -- | -- | 0 | +/- 0.1 | -- | -- | 0 | +/-1.2 mrad | -- | --
 | FPC | -- | -- | 1065.25 | +/- 2mm | -- | -- | -- | -- |  |  | -- | --
 | OUT (cap) | 0 | +/- 0.1 | -- | -- | 0 | +/- 0.1 | -- | -- |  |  | -- | --
Cav2 | IN (Cap) | 0 | +/- 0.1 | -- | -- | 0 | +/- 0.1 | -- | -- | 0 | +/-1.2 mrad | -- | --
 | FPC | -- | -- | 2027.74 | +/- | -- | -- | -- | -- |  |  | -- | --
 | OUT (cap) | 0 | +/- 0.1 | -- | -- | 0 | +/- 0.1 | -- | -- |  |  | -- | --
Valve plate 2 | B1 | 0 | +/- 0.5 | -- | -- | 0 | +/- 0.5 | -- | -- | -- | -- | -- | --
 | B2 | -- | -- | -- | -- | 0 | +/- 0.5 | -- | -- | -- | -- | -- | --
 | Vect Nb | -- | -- | 2504.5 | +/-5mm | -- | -- | 0 | +/-3.8 mrad | -- | -- | 0 | +/-2.5 mrad
Na
Nb
X
Y
Z
A2
V2 mm
Please note:
No thermal contraction or movement due to vacuum is considered at this stage
(V. Rude, T. Capelli).
A1
FPC
FPC
 | Step | 
Part 1 : in clean room | 1-2 | Alignment in ISO5 (before connection in ISO4)
 |  | Alignment in ISO5 (after connection in ISO4)

## Slide 11
 | Step | 
Part 2 : before cryostating | 4-5 | Alignment of different equipment
 |  | Installation of FSI supports
 |  | Measurement of FSI supports
 |  | Installation FSI targets
Objective / Results :
Cryogenic line : at 1 mm
Oblong bellows plate : at 1 mm
…
X
Y
Z (top)
X = 0.07 [0]
X = -137.46 [-137.5]
X = -137.55 [-137.5]
FPC
Y=149.88 [150]
Y=-1317.86 [-1318.5]
X = -0.08 [0]
X = -137.42 [-137.5]
X = -137.6 [-137.5]
FPC
Y=150.01 [150]
Y= 899.85 [900]
Y=- 2217.84 [-2218.5]

## Slide 12
 | Step | 
Part 2 : before cryostating | 4-5 | Alignment of different equipment
 |  | Installation of FSI supports
 |  | Measurement of FSI supports
 |  | Installation FSI targets

## Slide 13
 | Step | 
Part 2 : before cryostating | 4-5 | Alignment of different equipment
 |  | Installation of FSI supports
 |  | Measurement of FSI supports
 |  | Installation FSI targets
TROLLEY
ST4
TANK1
TANK2
ST1
ST2
ST3
ST5
ST6
2-3
1-4
5-8
6-7
10-11
9-12
13-16
14-15
Each FSI supports should be measured 3 times
Max dispersion : 15 μm per axis

## Slide 14
 | Step | 
Part 2 : before cryostating | 4-5 | Alignment of different equipment
 |  | Installation of FSI supports
 |  | Measurement of FSI supports
 |  | Installation FSI targets

## Slide 15
 | Step | 
Part 3 : cryostating (top plate) | 6 | Alignment of the cavities before cryostating
 |  | Measurement of the top plate cryomodule
 |  | Alignment trolley and top plate
Into taking in account vacuum, cold and hanging motions
Chart: Vertical position of the cavities w.r.t. Nominal
| 380.68200000000002 | 1260.4829999999999
TANK 1 | -0.931 | -0.793
TANK2 | -0.959 | -0.748
TANK1 Nominal position | -0.88 | -0.71
TANK2 Nominal position | -0.9 | -0.73
Chart: Radial position of the cavities w.r.t. Nominal
| 380.68200000000002 | 1260.4829999999999
TANK 1 | -0.951 | -0.768
TANK2 | -0.916 | -0.743
TANK1 Nominal position | -0.99 | -0.79
TANK2 Nominal position | -0.88 | -0.68
TANK 1
TANK 2
TANK 1
TANK 2
Nominal position
Position after alignment
Ry = -180 μrad [0]
Ry = -124 μrad [0]
Longitudinal axis (mm)
Longitudinal axis (mm)

## Slide 16
 | Step | 
Part 3 : cryostating (top plate) | 6 | Alignment of the cavities before cryostating
 |  | Measurement of the top plate cryomodule
 |  | Alignment trolley and top plate
R-top plate
Primary axis : Z : normal vector to top plate (on the edges)
Secondary axis : Y : normal vecor to lateral face
Origin : center of the 3 planes shifted of :
  X = 390 mm
  Y = 0 mm
  Z = 565 mm
Top face of the upper plate
Vertical reference
Lateral face of the upper plate Transversal reference
Extremity face of the upper plate
Longitudinal reference

## Slide 17
X
Y
Z
 | Step | 
Part 3 : cryostating (top plate) | 6 | Alignment of the cavities before cryostating
 |  | Measurement of the top plate cryomodule
 |  | Alignment trolley and top plate
The 2 coordinate systems should be at the same position

## Slide 18
 | Step | 
Part 4 : cryostating (OVC cryomodule) | 10 | Measurement of the OVC cryomodule
 |  | Alignment top plate and OVC cryomodule
R-top plate = R OVC
Primary axis : Z : normal vector to top plate (on the edges)
Secondary axis : Y : normal vecor to lateral face
Origin : center of the 3 planes shifted of :
  X = 390 mm
  Y = 0 mm
  Z = 565 mm
Advice :
Install top plate on the OVC.
The coordinate system of the top plate will be the coordinate system of the OVC

## Slide 19
 | Step | 
Part 4 : cryostating (OVC cryomodule) | 10 | Measurement of the OVC cryomodule
 |  | Alignment top plate and OVC cryomodule
OVC
The 2 coordinate systems should be at the same position

## Slide 20
 | Step | 
Part 5 : after cryostating | 10 | Installation of FSI heads on OVC and top plate
 |  | Measurement of FSI heads on OVC and top plate
 |  | FSI validation (comparison with a laser tracker)
 |  | Alignment with adjustment system

## Slide 21
 | Step | 
Part 5 : after cryostating | 10 | Installation of FSI heads on OVC and top plate
 |  | Measurement of FSI heads on OVC and top plate + TANK position
 |  | FSI validation (comparison with a laser tracker)
 |  | Alignment with adjustment system
EDMS : 2996397
Station
Cryomodule
Legend

## Slide 22
 | Step | 
Part 5 : after cryostating | 10 | Installation of FSI heads on OVC and top plate
 |  | Measurement of FSI heads on OVC and top plate
 |  | FSI validation (comparison with a laser tracker)
 |  | Alignment with adjustment system
 | Head | Target | FSI 3 D distance | Laser Tracker 3 D distance | Comparison
TANK 1 | 1 | 1 |  |  | 
TANK 1 | 2 | 2 |  |  | 
TANK 1 | 3 | 3 |  |  | 
TANK 1 | 4 | 4 |  |  | 
TANK 1 | 5 | 5 |  |  | 
TANK 1 | 6 | 6 |  |  | 
TANK 2 | 7 | 11 |  |  | 
TANK 2 | 8 | 12 |  |  | 
TANK 2 | 9 | 9 |  |  | 
TANK 2 | 10 | 10 |  |  | 
TANK 1 | 11 | 7 |  |  | 
TANK 1 | 12 | 8 |  |  | 
TANK 2 | 13 | 13 |  |  | 
TANK 2 | 14 | 14 |  |  | 
TANK 2 | 15 | 15 |  |  | 
TANK 2 | 16 | 16 |  |  | 
Should be less inside :
[-0.040 mm : 0.040 mm]

## Slide 23
 | Step | 
Part 5 : after cryostating | 10 | Installation of FSI heads on OVC and top plate
 |  | Measurement of FSI heads on OVC and top plate
 |  | FSI validation (comparison with a laser tracker)
 |  | Alignment with adjustment system
EDMS : 3065249
Should be less inside :
[-0.040 mm : 0.040 mm]

## Slide 24
 | Step | 
Part 6 : cold test at STFC or Triumf | 12 | Measurement ambient pressure
 |  | Measurement under vacuum
 |  | Measurement at cold
Station
Cryomodule
Legend

## Slide 25
 | Step | 
Part 6 : cold test at STFC or Triumf | 12 | Measurement ambient pressure
 |  | Measurement under vacuum
 |  | Measurement at cold
Station
Cryomodule
Legend

## Slide 26
 | Step | 
Part 6 : cold test at STFC or Triumf | 12 | Measurement ambient pressure
 |  | Measurement under vacuum
 |  | Measurement at cold
Station
Cryomodule
Legend

## Slide 27
Thank you
for your attention

## Slide 28
SPARE

## Slide 29
Alignment Objective n°1 : Fiducialisation
Bare Cavity
Cavity radio frequency axis  Approximated to mechanical axis of the capacitive plate
Dressed Cavity
8 Alignment targets :
FSI targets
Alignment requirement
Position : 50 μm (1)
Roll : 100 μrad (1)
Axis of the
capacitive plate
Axis of the
capacitive plate

## Slide 30
Alignment Objective n°2 : Internal monitoring
Cavities RF
Cryomodule
Alignment requirement
Position : < 100 μm (1)
Roll : < 400 μrad (1)

## Slide 31: Internal monitoring : Configuration
  FSI : Frequency Scanning interferometry
   Absolute distance measuring interferometric technique
Largest standard uncertainties | Uncertainty (1)
Position of the FSI Sensor in the framework of the cryomodule | 40 µm
Position of the FSI target in the framework of the crab-cavity | <15 µm
FSI target
on the flange of Helium TANK of the cavities
reflection
Reflecting
glass sphere
(FSI target)
Flange of Helium Tank of RF Crab-cavities
Support
FSI target
FSI Head (sensor)
on the Cryomodule
Optical fibre
FSI Acquisition system
Feedthrough
(FSI Head)
Fiber ferrule tip
8 FSI meas.
8 FSI meas.
Total : 16 observations
(3D distances)

## Slide 32: Internal monitoring : Challenges and prerequisites
Beam
All adjustments should be carried out at ambient pressure.
Ambient pressure
Measurement + Adjustment
The position of two cavities should be determined at cold at less than 0.1 mm (1sigma)
The position of two cavities at cold should be aligned at less than 0.1 mm (1sigma)
Two inner components
Non compensable misalignment
on the external envelope
No possibility to adjust
ANTICIPATION : The position of the cavities should be at their nominal position at warm with an accuracy of less than 40 microns.
The position of the FSI sensors w.r.t. cryomodule must be measured with an accuracy of less than 40 microns.
Cooling down
Anticipation
Anticipation of the Deformation of the outer envelope with an accuracy of less than 40 microns
(deformation up to 0.1 mm)
Anticipation of the Movement of the cavities :
(Complex movement up to 1.2 mm
7 DOF)
Anticipation of the Deformation of the outer envelope with an accuracy of less than 40 microns
(deformation up to 1.2 mm)
Anticipation of the Movement of the cavities :
(Complex movement up to 0.4 mm
6 DOF)
Under vacuum
Anticipation
Cryomodule
FSI Sensor
RF cavity
RF cavity

## Slide 33
Alignment analysis of the RFD prototype during cold test at CERN

## Slide 34
Ambient pressure
Under vacuum
Shape Repeatability
For 3 cycles :
Maximum dispersion : 67 μm
Average dispersion : 18 μm
Shape Repeatability
For 3 cycles :
Maximum dispersion : 57 μm
Average dispersion : 16 μm
Max deformation :
1.20 mm
Deformations of the cryomodule (due to pumping and cooling down)
No deformation was observed on the edge of the top plate during pumping.
Shape of the cryomodule under vacuum
Repeatable at +/- 0.040 mm (RMS)
Shape of the cryomodule at ambient pressure
Repeatable at +/- 0.040 mm (RMS)
Cold
Shape Repeatability
For 3 cycles :
Maximum dispersion : 68 μm
Average dispersion : 17 μm
Max deformation :
0.12 mm
During cooldown, the top plate observed a thermal contraction due to thermal convection between the thermal screen and the top plate --> not expected but repeatable--> Can be anticipated
Shape of the cryomodule at cold
Repeatable at +/- 0.040 mm (RMS)

## Slide 35
Chart: Relative vertical motion (mm)
| 0 | 3000
Objective at cold | 0 | 0
Cav 2 : ambient pressure | -0.855 | -0.781
Cav 2 : under vacuum | -1.262 | -1.122
Cav 2 : Cold | 0 | 0
IN
OUT
875 mm
Step 3 : Cavity with
cryomodule at cold
Step 1 :
Cavity with cryomodule
at ambient pressure
Step 2 : Cavity with
cryomodule under vacuum
Chart: Relative Radial motion (mm)
| 0 | 3000
Objective at cold | 0 | 0
Cav 2 : ambient pressure | -0.71 | -0.436
Cav 2 : under vacuum | -0.336 | -0.062
Cav 2 : Cold | 0 | 0
Step 3 : Cavity with
cryomodule at cold
Step 1 :
Cavity with cryomodule
at ambient pressure
Step 2 : Cavity with
cryomodule under vacuum
Relative Motions of the cavities
due to vacuum and cooling down (cycle 3)