## Slide 1: Intensity measurement
Marek Gąsior, Michał Krupa
CERN Beam Instrumentation Group
BI CAS 2025, Split, Croatia

## Slide 2
Introduction
Intensity is one of the most important beam parameters
Intensity measurement is essential for:
  operating an accelerator
  quantifying an optimizing accelerator performance
  scaling physics data (accuracy)
  (performance most important)
Intensity measurement can be critical for:
  accelerator protection
  equipment protection
  human health (e.g. medical applications)
  (reliability most important)

## Slide 3
A selection of topics related to operation and designing beam intensity measurement systems based on electromagnetic sensors working on the time scales ~  .. ~ 0.1 ns (frequency range ~ 0 .. ~ 10 GHz)
Faster techniques in two lectures ”Short Bunch Measurements I + II” by Thibaut
Outline
Electronics
Low frequency compensation
Calibration
Beam signal filtering
Cohabitation with power systems
Sensors
Wall Current Monitor
Fast Beam Current Transformer
Integrating Current Transformer
Wall Current Transformer
Faraday Cup
DC Beam Current Transformer
Cryogenic Current Comparator
Fundamentals
Wall charge and wall current
Beam current, charge and intensity
Low pass filtering and signal integrals
Intensity measurement time resolution
Numerical integration
Baseline restoration and correction
In spare slides and the proceedings
as well as Jeroen’s lecture

## Slide 4
wall charge density
beam charge density
wall current
beam current
RMS of the wall charge distribution
rBP – radius of the beam pipe
v, c – speed of the particles and of the light
Wall charge and wall current
(a recap of Gero’s lecture)

## Slide 5
Beam current, charge and intensity
N
T
IM
σ
= beam intensity = number of particles
= integration interval
= maximal (peak) current
= RMS bunch length
number of charges q1 during time interval T
charge of a Gaussian bunch
current distribution in a Gaussian bunch
current
charge
intensity
charge of one particle
Q1 = Q2= Q3
IM
σ1 = 1
σ2 = 2
σ3 = 4
IM1 = 1
IM2 = 0.5
IM3 = 0.25
i1 (t)
i2 (t)
i3 (t)
beam beam intensity charge
in BI world often
:

## Slide 6
A low pass filter should be used before any active electronics:
To pass only desired beam spectrum, from a “good part” of the sensor bandwidth
To limit the signal slew rate and exclude potential amplitude nonlinear response of the following active electronics
To adapt pulse lengths to the acquisition sampling rate and optimise the accuracy of the integration
Low pass filtering and integrals
Low pass filtering conserves the signal integral
low pass filtering does not affect beam charge calculated as the beam current integral

## Slide 7
Intensity measurement time resolution
T(t)
= revolution period
integration
rectangular surface calculation
one revolution period
trains
batches
bunches
DC beam current (!)

## Slide 8
first order hold
trapezoidal integration
second order hold
parabolic integration
parabola
section
Beam current numerical integration
integration
error
zero order hold
rectangular integration
symmetry
≈ conserved
zero order hold
rectangular integration

## Slide 9
sampling phase 2
sampling phase 1
Beam current numerical integration
Integration noise due to changing sampling phase
9
signal difference P1 – P2
Convenient symmetrical changes ONLY for symmetrical pulses

## Slide 10
Typical signals – pulses
Typical signals from intensity sensors with no DC response induced by rectangular beam pulses
Many intensity sensors are blind for DC signals
no DC net charge = 0
| positive charge | = | negative charge |
positive charge + negative charge= 0
“0 line”
10

## Slide 11
Typical signals – trains
11

## Slide 12
Typical signals – circulating beam
12

## Slide 13
Baseline restoration
integration windows
restored baseline
corrected baseline

## Slide 14
(restored) Baseline correction
integration windows
restored baseline
corrected baseline

## Slide 15
Baseline restoration and correction – examples
J
LLL

## Slide 16
Wall Current Monitor (WCM)

## Slide 17
Wall Current Monitor (WCM)
= one port output voltage
= beam current
= load resistor
= number of ports
= transfer impedance
= one port output current
vO1
iB
RL
n
ZTR
iO1
The beam pipe gap is loaded with typically 8 transmission lines, most often 50 Ω (RL = 50 Ω)
The gap connection (in red) is very critical for the WCM performance (must be kept very short)
For the vacuum gap the ferrite is in the beam vacuum (a serious issue for good vacuum, often a dedicated pump)
wall current

## Slide 18
Wall Current Monitor (WCM)
@,
fCL fCH
f'CH
iE
LE
iG
CG
LP
= cut-off low frequency
= “gap” cut-off high frequency
= “series L” cut-off high frequency
= enclosure current
= enclosure inductance
= gap current
= gap capacitance
= parasitic series inductance
wall current

## Slide 19
WCM – position signal dephasing
fastest risetime
(shortly an example)

## Slide 20
Wall Current Monitor (WCM)
To combine …
Good signal position independency (only position dephasing)
Signal deterioration by the combiner
… or not to combine
Largest position dependency
Best signal quality
Some decoupling between respective outputs
Saved money and thinking on the combiner

## Slide 21
WCM example 1 (CTF)
High frequency direct output bandwidth: 250 kHz – 10 GHz
Low frequency output with an active compensation: 10 kHz – 300 MHz
Courtesy of Patrick Odier
Curtesy of P. Odier

## Slide 22
WCM example 2 (LHC)
Gap 1.2 mm
High frequency cut-off ≈ 3 GHz
Courtesy of Tom Levens
ferrite distributed along the length

## Slide 23
Fast Beam Current Transformer (FBCT)
Workhorse of intensity measurements in many accelerators
In a circular machine one FBCT is enough, while linacs and transfer lines often require many of them
It is a current transformer, with the beam forming the primary single turn winding
Particular FBCT construction often determined (and sometimes limited) by the choice of commercially available ferromagnetic toroids
FBCTs are available commercially, along with their dedicated electronics
Courtesy of Romain Ruffieux
ferromagnetic toroid
ceramic insert
23

## Slide 24
Fast Beam Current Transformer
iB ~ 100 mA (CERN L4 proton linac)
iB ~ 100 μA (CERN L3 ion linac)
24

## Slide 25
Fast Beam Current Transformer (FBCT)

## Slide 26
FBCT signal droop
exp
exp
n = 5, 10, 20, 50
n = 5, 10, 20, 50
exp
26
v1 – output voltage for n=1
τ1 – time constant for n=1

## Slide 27
n = 5, 10, 20, 50
= 3, 10, 30, 100
exp
exp
FBCT signal droop
27
exp
v1 – output voltage for n=1
τ1 – time constant for n=1

## Slide 28
FBCT high frequency response
200 ns
“in flange” FBCT

## Slide 29
FBCT high frequency response

## Slide 30
Integrating Current Transformer (ICT)
ICT ≈ FBCT + good BUILT-IN low pass filter

## Slide 31
Integrating Current Transformer (ICT)
For bunches much shorter than the ICT pulse response, the ICT output pulse shape is independent of the bunch shape (so ICT integrates short beam current pulses)
For clean response ICT output may need to be further filtered by an external low-pass filter (e.g. on the input of its amplifier)
CERN AWAKE
ps electron bunches

## Slide 32
FBCT and ICT position dependency
≈ 1 % of the intensity
≈ 2 % of the pipe diameter

## Slide 33
Wall Current Transformer (WCT)

## Slide 34
Wall Current Transformer (WCT)
WCT ≈
WCM + FBCT
Invented for the LHC to assure bunch-by-bunch intensity measurement independent of the beam position and bunch length
The WCM gap load is replaced by small transformers allowing very small gap loads (RG = RL/ n2) (C of the ceramic insert not important)
Small ferromagnetic toroids assure very good high frequency operation
Position dependent signals on respective outputs are combined by high frequency combiners, internal on PCB and external
The RF bypass provides the control of the high frequency response and the impedance presented to the beam
Particular mechanical construction for the LHC to be mounted without opening the vacuum

## Slide 35
Wall Current Transformer (WCT)
35

## Slide 36
Frequency response comparison: WCT, ICT, FBCT
WCT
full BW
WCT
400 MHz

## Slide 37
Faraday cup
Measurements from DC to GHz with special designs
The only destructive sensor in the lecture
“Natural” operation if combined with a beam dump (dissipated power may be a challenge)
In any other location it requires some mechanics operating in the beam vacuum to put the sensor to/from the beam
Can be designed to measure very small and very large currents
The energy of the measured beam is very important for the design (low energy more difficult)
Available commercially, along with dedicated acquisition
Well described in the literature

## Slide 38
Faraday cup
https://doi.org/10.1063/5.0131890
www.dreebit.com

## Slide 39
IB
IDC
Vm
i1, i2
= beam current
= compensating DC current
= (triangular) modulation voltage
= T1, T2 currents
T1, T2
CCS
ASP
= ferromagnetic toroids
= controllable current source
= analog signal processing
DC Beam Current Transformer
39
Very small dynamic range of the sensitive detection circuitry addressed by using a compensation method
“Zero flux” operation
Excellent precision and dynamic range typical for DC measurements
In many accelerators a golden standard used to calibrate other intensity measurement systems

## Slide 40
DC Beam Current Transformer
IB
IDC
Vm
i1, i2
= beam current
= compensating DC current
= (triangular) modulation voltage
= T1, T2 currents
T1, T2
CCS
ASP
= ferromagnetic toroids
= controllable current source
= analog signal processing
Very small dynamic range of the sensitive detection circuitry addressed by using a compensation method
“Zero flux” operation
Excellent precision and dynamic range typical for DC measurements
In many accelerators a golden standard used to calibrate other intensity measurement systems
40

## Slide 41
DC Beam Current Transformer (DC BCT)

## Slide 42
Cryogenic Current Comparator (CCC)
CCC of the CERN Antiproton Decelerator
Photos from the left:
commercial SQUID system
Magnetic circuit assembly
CCC tank, Ø ≈80 cm, with ≈50 litres of liquid helium
Photos: curtesy of Miguel Fernandes
CCC ≈ DC BCT with a quantum supersensitive “zero” flux detector

## Slide 43
FBCT, ICT, WCT calibration
JJJ
L
The great one-turn calibration winding allows:
quickly testing the sensor anytime, anywhere, and if connected there, its whole acquisition chain
calibrating the amplitude scale of the whole acquisition chain in real-time, by applying precise reference signals
43

## Slide 44
No precise calibration with only an oscilloscope
A signal generator connected successively to each of the four inputs of an oscilloscope
Acquired waveforms stored in the oscilloscope memory
Identical channel settings, Identical generator settings, the same coax cable
12-bit, 1 GHz, 2.5 GS/s
Special techniques necessary:
DC precision measurement AND “zero detection” with an oscilloscope
See proceedings for details

## Slide 45
FBCT, ICT, WCT calibration errors
from the beam “point of view”

## Slide 46
Low freqnecy compensation: integrating amplifier
The first order high-pass response of the sensor is compensated by the “integrating” first order slope of the amplifier
Extending the sensor cut-off by a factor requires a proportional increase of the DC gain
The amplifier frequency characteristic must be well matched to the characteristic of the sensor. Spread in inductances of FBCTs in practice means that each FBCT must have a dedicated amplifier.
An input low-pass is easy to implement and, if required (e.g. radiation), the amplifier can be connected to the sensor through a long coaxial cable
In practice only the first amplifier stage has the DC response and the following stages are coupled capacitively with the cut-off frequencies much lower than the low-cut of frequency of the FBCT
FBCT, ICT WCT, WCM, …

## Slide 47
LF compensation: transimpedance amplifier
Very low input impedance of a transimpedance (TI) amplifier loading a FBCT output assures its good low frequency response. This is achieved at the expense of large amplifier DC gains. The very low input impedance is only within the amplifier bandwidth.
Spread in inductances of FBCTs connected to TI amplifiers results in a corresponding spread in the low frequency cut-offs
In general, a TI amplifier input must be directly connected to its corresponding FBCT (radiation issues)
Because of its very low input impedance, low pass filters are difficult to implement before the TI amplifiers. An adequate filtering can be achieved in the following classical stages, while limiting the gain of the input TI amplifier to an adequate value
In practice only the first amplifier stage has a DC response and the following stages are coupled capacitively with the cut-off much lower than the FBCT (for millisecond beam pulses even below 1 Hz)
Very low input impedance of TI amplifiers keep the calibration turn error very low
“virtual ground”
theory
practice

## Slide 48
one revolution period
trains
batches
bunches
(Sensor bandwidth) Summary
invasive
WCM FBCT ICT WCT FC DCBCT CCC
DC

## Slide 49
Intensity is one of the most important beam parameters
Beam intensity measurements are crucial for machine operation, optimisation, scaling physics data, sometimes also for protecting accelerator, equipment or even human health
A good intensity sensor is a requirement, but not a guarantee of a good beam intensity measurement system, as the sensor must be followed by adequate cabling, filtering, amplification, acquisition and software
To cover all operational needs often beam intensity is measured by a few complementary systems (unbunched beams, turn by turn, bunch by bunch, bunch temporal profile, …)
In general, any active electronics should be preceded by an adequate low pass filter
Low pass filtering does not affect beam charge calculated as the beam current integral
Intensity measurement accuracies better than a couple of percent require good calibration strategies
A beam current transformer “calibration winding” introduce an intensity measurement error
In general, oscilloscopes do not provide amplitude measurements with accuracy better than a couple of percents. For better accuracies special techniques should be used. Such techniques will be explained in the proceedings.
An adequate amplifier can dramatically improve the sensor low frequency response
BPMs can be used for intensity measurements and in special cases they can even overtake “true” intensity sensors. Examples will be provided in the proceedings.
For more on intensity measurements please read the proceedings (+ literature, electronics, calibration, cohabitation with power systems, intensity with BPMs, …)
A few more important points to take away
49

## Slide 50: Spare slides

## Slide 51: A mathematical proof that beam intensity is more important than beam position
position = 0  J
intensity = 0  L L L

## Slide 52
(FBCT, ICT, WCT, …) Signal filtering
to use
to throw
filter
classic reactive lossless LP filter
passband
stopband

## Slide 53
to use
to throw
passband LP
passband HP
filter
nonreflective absorptive constant impedance
filter
(FBCT, ICT, WCT, …) Signal filtering

## Slide 54
(FBCT, ICT, WCT, …) signal filtering
Z = 50 Ω
Z = 300 Ω
Z = 100 Ω
Z = 50 Ω
(100 || 100)
Z = 50 Ω
(60 || 300)
Z = 50 Ω