## Slide 1: LHC linear wire scanner CONS Project meeting mechanics update
10th Feb 2025
Harry Sullivan
Notes: Harry

## Slide 2: Overview
Instrument design:
Chamber changes (MME)
  Solid block machining
  Welded bottom cap
  Locations
Electrical contact
  Electrical feedthrough location
  Coating feedback
Ruler
  Changes to optical feedthroughs with MME
  Review complete
  Awaiting patterns
  Drawing
Changes to RF contacts
Test rig:
In house parts
  Material arrived
  Machining to commence end of next week
External parts
  Card water jet complete
  Bearings to arrive now
Next testing
  Ruler spec to be sent to Selba
  Test rig to be machined for Nabil’s rig
  Electrical contact prototype to be machined

## Slide 3: Test rig update
Updates:
Stainless steel 1mm panels installed
Initial results show negligible effect on motion
Next steps:
Integrate new components to test realistic loading case
Integrate new bearings and tensioning method
Internal parts:
  Material delivered
  Machining to start end of this week
External parts:
  Card water jet cutting complete
  Bearings delivered to CERN
Future testing
Design prepared for integrating glass ruler into Nabil’s test rig
Design required for integrating glass ruler into this test rig
Design required for integrating sliding electrical contact into this test rig

## Slide 4: Test rig observations
40k cycles report:
Residue/dust seen all four sides of the rail
  Most deposited on window side rail
  Could not be felt and was not removed with finger
Wipes off clean with alcohol and tissue and rail appears – after which rail appears undamaged
  With visible track marks in the right light, as seen after a few thousand cycles
Bearings were all marked with the same dust/residue
Carriage still feels rigidly held in place
Conclusions
Dust is in very small quantities, and appears to stick to rail – so should not provide a UFO risk
Particulates do not appear to affect the performance of the carriage
Carriage connection is unbalanced, due to the difference in deposition – to be rectified with next iteration
Deposition estimated to be less when new bearings integrated

## Slide 5: Instrument design updates
Detailed design started with MME
Meetings with workshop held to identify optimum machining processes
Material for chambers on order
Design progress advanced, estimated chamber to be finished next week

## Slide 6: Instrument design updates
Ruler
Ruler assembly concept complete
Ready to order test samples from Selba, once required patterns identified
Modifications to optical laser port assembly implemented

## Slide 8: Instrument design updates
Instrument chamber
To be machined from solid billet, due quantity required
CF feedthrough/ porthole features added
Rail alignment feature added

## Slide 9: Instrument design updates
Electrical contact
Fixed wires to connect sliding contact to feedthrough
Same connector used as seen on rotary scanner
Sliding contact
  Copper deposition on ceramic, with rhodium plating
  Silver plating on copper beryllium spring contacts
  Should have lifetime far exceeding requirement for scanner, and provide very low friction

## Slide 10: RF contacts

## Slide 11: Instrument design updates
Vacuum chamber
Welded assembly
Main body to be machined from solid, (EDM wire cut)

## Slide 12: Other updates
Meeting to be organised with vacuum to review outstanding design questions
12 x Magnets cleaned and ready for vacuum testing
Final model to be simulated with Impedance when complete (planned for next week)
Impedance working group next week