## Slide 2: Cryostat Joining Techniques and Leak Testing
T. Koettig on behalf of the Cryolab team

## Slide 3: Content
Mechanical joints
Solder joints and welded connections
Outgassing vs. leaks
Leak detection
Electrical breakdown

## Slide 4: Mechanical connections

## Slide 5: Transfer line examples
Thermal barrier
Insulation by foam, powder or MLI
MLI
Line 1
Line 2
Assembly weld
Rigid transfer line connection
Pumping port
MLI
Internal tube
Ext. tube
Adsorbent
Bellow
Support, spacer
Thermal barrier
Bayonet Johnston connection
Connection flanges
Warm zone
Joint, O-ring or gasket (Al, Cu, In)
Annular gap of 50 to 100 micron possible cold Teflon seal to reduce convection
Cold zone
MLI
Adapted from: Cryogenie, Institute International du Froid

## Slide 6: Low-temperature sealing techniques- demountable joints
Joint techniques Helicoflex and CF
From: Ekin, Experimental Techniques for Low Temperature Measurements
Leak tight pipe connections
www.swagelok.com: VCR, with wide variety of connectors, valves etc.
Kenol connector, metal flat seal, 1 side with convex shaped surface

## Slide 7: Low-temperature sealing techniques- demountable joints
Joint techniques Indium
From: Lim, Indium seals for low temperature and moderate pressure applications, Rev. Sc. Instr. 1986
From: Katheder, VDI course Cryostats
Compression curve for a 1 mm Indium wire
Force/length (N/cm)
Remaining thickness in mm

## Slide 8: Tube feedthrough and pressed mechanical contact
Tube feedthrough
Adapted from:
Ekin, Experimental Techniques for Low Temperature Measurements
Stainless steel tube
Stainless steel flange matching the main flange
Welded joint
Preferred solution double O-ring ?

## Slide 9: O-ring groove, and positioning
Inner overpressure: chose the O-ring approx. 1 to 2% larger than the outer groove diameter d7.
External overpressure: chose the O-ring approx. 1 to 3% smaller than the inner groove diameter d8.
Dimensioning of the O-ring groove for 15-20 % compression and allow deformation in groove width of +30 % (e.g. 5 mm O-ring depth 4.2 mm, width 6.7 mm)
https://www.rubbershop.ca/pages/o-ring-groove-design

## Slide 10: O-ring materials vs. temperature
From: https://www.barnwell.co.uk/what-is-an-o-ring/
HNBR
NBR
Rubber
PEEK
PTFE
Silicone
EPDM
Damaged KF 16 O-ring, reason? oil, chemicals…

## Slide 11: Joining techniques

## Slide 12: Material choice – solder joints
From: Ekin, Experimental Techniques for Low Temperature Measurements
Soft solder is better in thermal contact below 100 K
Attention with SC transition =>

## Slide 13: Material choice – joints and thermal contact
From: Ekin, Experimental Techniques for Low Temperature Measurements
Thermal conductance vs. T
data of solder, varnish and grease is per 1 cm2
Pressed joints are area independent and only vary with force!
Attention to thermal dilatation for pressed contacts!

## Slide 14: Welded connections – things to consider
From: Ekin, Experimental Techniques for Low Temperature Measurements
From: Parma, Cryostat design, CAS 2013
Fin for similar thermal mass

## Slide 15: Moving vs. fixing things
From: Tobler, Materials for Cryogenic Wind Tunnel Testing
Temperature dependence of self-friction coefficients for selected materials
Effect of lubricating films on the friction of copper.
Stainless steel and copper => silver coating for UHV applications, screws, gaskets

## Slide 16: Outgassing and leak rate

## Slide 17: Vacuum degradation – leaks, outgassing, diffusion
From: Ekin, Experimental Techniques for Low Temperature Measurements

## Slide 18: Vacuum degradation – leaks, outgassing, diffusion
From: Zapfe, Leak Detection, DESY
V = 1L

## Slide 19: Vacuum degradation – leaks, outgassing, diffusion
Example from BCCCA vacuum vessel

## Slide 20: Vacuum degradation – outgassing
A.G. Mathewson et al.
J.Vac.Sci. 7(1), Jan/Fev 1989, 77-82
Unbaked Al
Metallic surfaces q ~ q0/t
Plastic surfaces q ~ q0/√t
x 5 000
A.G. Mathewson et al.
J.Vac.Sci. 7(1), Jan/Fev 1989, 77-82
Baked Al
V. Baglin, CAS basic course, vacuum systems, 2021
 | H2 | CH4 | H2O | CO | CO2
Unbaked | 7x10-12 | 5x10-13 | 3x10-10 | 5x10-12 | 5x10-13
Baked | 5x10-13 | 5x10-15 | 1x10-14 | 1x10-14 | 1x10-14
Stainless steel after 10 h of pumping (Torr.ℓ/s/cm2)

## Slide 21: Vacuum degradation – leak rate
Poiseulle’s law describing the mass flow :
d hydraulic diameter
s wall thickness
density
dyn. viscosity
pressure diff. in Pa
Lebrun, CAS course proceedings, 1989
He II => Superleak?

## Slide 22: Vacuum degradation – back diffusion
From: Dauvergne, Backstreaming of Impurity Gas through a Leak in a pressurized Vessel, LHC Project Report 229
Above 5×10-9 mbar l/s:
helium flux is dominated by drift,
out-streaming gas effectively hinders the back-streaming of impurities,
maximal air flux appearing for helium leaks of the order of 10-4 mbar l/s.

## Slide 23: Leak detection

## Slide 24: Leak rate detection option
Lebrun, CAS course proceedings, 1989

## Slide 25: Mean free path of molecules
V. Baglin, CAS basic course, vacuum systems, 2021
The turbulent flow is established around the atmospheric pressure
In the low vacuum (103 - 1 mbar), the flow is viscous and laminar.
In the high vacuum (10-3 - 10-7 mbar) and ultra-high vacuum (10-7-10-12 mbar), the flow is molecular. The mean free path is much larger than the vacuum chamber size. Molecules interact only with the vacuum chamber walls.
Molecular
p~10-3 mbar
λ~10 cm

## Slide 26: Vacuum degradation – leaks, outgassing, diffusion
From: Zapfe, Leak Detection, DESY
Localized test
Integral test
Sensitivity

## Slide 27: Leak detector steps/models
Soap spray for large leaks, system in overpressure to atmosphere
Le pistolet:
measures He, CO2
He gas bottle with PRV to 50 mbarg and He sprayer, < 2 bubbles per sec!
Leak detector with its roughing pump connected to the cryostat pumping system, integral test.

## Slide 28: Mass spectrometer principle
Adapted from: https://chem.libretexts.org/Bookshelves/Analytical_Chemistry/Supplemental_Modules_(Analytical_Chemistry)/ instrumental analysis/mass spectrometry/
gas stream
e/m selection
gas stream
Step 1: Ionization to +1 by e- stream
Step 2: Acceleration to same speed
Step 3: Deflection in perpendicular magnetic field, Lorentz force
Step 4: Detection, neutralization by e- current
4He => 4He+
m/q=4

## Slide 29: Vacuum degradation – leaks, outgassing, diffusion
Sniffing mode => He inside the process vessel:
Keep the direction of He in air in mind
=> Start sniffing from the bottom!
Detection time is ~ 3-5 sec.
Separate parts from each other (plastic bag etc.)
Alcohol can temporarily close small leaks
Mustique can help as well
Careful with O-rings etc. 10-7 mbar l/s
From: G. Gisteau, INTRODUCTION TO HELIUM REFRIGERATION FOR SUPERCONDUCTING DEVICES

## Slide 30: Vacuum degradation – leaks, outgassing, diffusion
From: Zapfe, Leak Detection, DESY
Vacuum mode test procedure:
Pump vacuum system for one night
Connect and start leak detector, pump connection line!
Close V1 valve towards the roughing pump
Open slowly V2 valve to leak detector
Spray small amount of He from gas bottle p~50 mbar
Start from the top! (He gas rises)
Wait and check signal for 30 sec.
Repeat spray of He gas and check
Repair big leaks and start fresh! Small leaks can be temporarily closed by alcohol (Mustique, chewing gum)
Small leaks <10-8 mbar l/s need to be confirmed 3 times in a row to be valued => DR rule! esp. for 3He circuit
Progress to next location in downward direction
V1
V2

## Slide 31: Vacuum degradation – leaks, outgassing, diffusion
From: Zapfe, Leak Detection, DESY
Test procedure remarks:
Rubber O-ring seals and epoxy feedthroughs giving a signal of 10-7 to 10-8 mbar l/s after some time of He exposure
Close first V2 valve before stopping leak detection mode (possible venting of air)!
Open V1 valve to roughing pump
V1
V2
Good practice, values of leak tightness for:
insul. vacuum => 10-7 mbar l/s
4.2 K He => 10-8 mbar l/s
He II cooling => 10-9 mbar l/s
3He systems => 10-10 mbar l/s

## Slide 32
LN2 circulation
LN2
Cryostat flange
G10 supports
Sample
Thermal shield
DAQ
Pumping system and leak detector
Cylinder
GHe
LN2
Cryostat
Experimental setup with samples and list of tests:
Strip cable
(32 signals)
3D-printed epoxy or fiberglass structure
High density feedthrough flange:
Tests:
-Pressure and leak test at room temperature
-Pressure and leak test at low temperature
High density feedthroughs for future detector calorimeters of LAr

## Slide 33: Electrical breakdown fct. (p, distance and T)
Source Wikipedia
d
pgas
- +
Room temperature

## Slide 34: Electrical breakdown fct.(p, distance and T)
Source Wikipedia
d
pgas
- +
Transitions to liquid density

## Slide 35: Magnet helium conditions during Quench and HV testing
Further reading: E. Sili, Temperature Dependence of Electrical Breakdown Mechanism on the Left of the Paschen Minimum

## Slide 36: Summary
How to join different parts
Metal and polymer dismountable joints
Outgassing of materials
Leak detection methods
Example for FCC calorimeter cryostat
Electrical breakdown