## Slide 1
Iker LOBATO SOLA (TE-VSC-ICM)
TE-VSC Seminar
25/11/2025
Influence of humidity and temperature on Penning gauge measurement

## Slide 2: Outlook
Pfeiffer IKR070 Penning gauge
Found issues or inconsistencies
First findings and actions
Investigation
Solutions
Conclusion

## Slide 3: Pfeiffer IKR070 Penning gauge
LHC P4 vacuum application synoptic example
What it is?
The Penning gauge or cold cathode ionization vacuum gauge (VGP) is a vacuum pressure measurement instrument which provides measurements in the range of 10E-4 to 10E-11 mbar
Who is the manufacturer?
Distributed by Pfeiffer Vacuum
Designed and produced by INFICON AG (former Balzers)
What do we use them for?
To monitor the vacuum in our setups and accelerators
To give interlock to users and sector valves
How many are used at CERN?
More than 1000 units installed at:
cPS, SPS, LHC, experiments, etc.
Pfeiffer IKR 070
INFICON MAG 070

## Slide 4: Pfeiffer IKR070 Penning gauge
How it works?
Consists of two electrodes, an anode and a cathode, between which a high voltage (3.3 kV DC) is applied. When the negatively charged electrons leave the cathode at high velocity toward the anode, they ionize the neutral gas molecules in the chamber, igniting a gas discharge. This gas discharge current is proportional to the pressure
Pressure vs current transfer function
N
S
Anode
Cathode
Magnet
B
3.3 kV
N
S
+
-
A
IKR070 transfer function P-I (courtesy of N. Chatzigeorgiou)
IKR070 Specs
3pA = 1.0E-11mbar
3pA ≤ uR (Under range)
0,1mA = 2.3E-05mbar

## Slide 5: The issue
First observation
AD and ELENA OP observed an abnormal pressure fluctuation in the ELENA ring
AD Logbook screenshot
ELENA Logbook screenshot

## Slide 6: The issue
Our first observation
We noticed that an AD ring gauge showed the same behaviour
Timber plot of the AD and ELENA VGP readouts, July 2022
AD VGP
ELENA VGP

## Slide 7: The issue
First hypothesis
Controller, equipment or cabling issue or malfunction
Timber plot of the AD and ELENA VGP readouts and LNR %RH sensor, July 2022
First lead
Certain correlation between a %RH signal at the LNR and the pressure

## Slide 8: Logging hygrometry data
First action
We placed our own humidity and temperature Data Loggers “close” to the affected gauges
AD ring access PAD area
ELENA ring
AD/ELENA b.193
EL-USB-2 specs

## Slide 9: Logging hygrometry data
Logged Data
The warm weather, in addition to storms, affected the humidity and temperature levels in the hall, which affected the VGP readouts
AD hygrometry data from July and August 2022
AD DR4.VGP3806 and dew point
comparison (plot by J.A. Ferreira)

## Slide 10: Climatic Chamber
Climats EXCAL 1423-HA (2013)
  Temperature control from -70°C to +180°C
  Humidity control from 10 to 98%
  138L (550 x 500 x 500 mm)
  Admissible dissipation at 20°C = 2.5kW
  Location : b.13-S-011
  EP-ESE
Climats EXCAL 1423-HA climatic chamber
EXCAL 1423-HA touch panel and program
  Room controlled ambient conditions:
  ~ 20°C
  ~ 40%RH

## Slide 11: Setup
Experimental setup prepared for the tests
Objective to be fitted into the EXCAL machine
  Penning gauge (Pfeiffer IKR070)
  NEG cartridge
  Manual valve
  Pinch-off port
VGP testing setup
Features
  Baked-out
  Very low vacuum pressure
    E-12mbar range
    Penning reading under range (uR)
  Powered NEG cartridge
    Increase the temperature to ignite the gauge

## Slide 12: First test
Test performed by increasing 10% humidity every 3h. Fixed 25°C starting at 40%RH up to 90%RH and back
Standard IKR070 VGP measurements at the EXCAL machine
ΔI ~40pA Δ𝑃 ~6.0E-11 mbar

## Slide 13: Second test
Test performed by increasing 10% humidity every 24h. Fixed 25°C starting at 40%RH up to 80%RH and back
Standard IKR070 VGP measurements at the EXCAL machine
ΔI ~200pA Δ𝑃 ~1.0E-10 mbar

## Slide 14: LEMO connector and HV cable
  LEMO connectors
  Plug ref.: FFB.3S.460.CLAC10
  Socket ref.: EHP.3S.460.CLL
  Operation thresholds:
  Temp.: -55 °C to 250 °C
  Hum.: 95 %RH at 60 °C
LEMO socket on the left and plug on the right
  TFA3 cable – Triaxial FRNC
  Operation thresholds:
  Temp.: -20 °C to 60 °C
  Hum.: no information
Insulated LEMO socket and TFA3 cable with a LEMO plug at the EXCAL chamber at %80RH and 25 °C

## Slide 15: LEMO connector and HV cable test
Test performed by increasing 10% humidity every 24h. Fixed 25°C starting at 40%RH up to 80%RH and back
Insulated LEMO connector measurements at the EXCAL machine
 uR

## Slide 16: Heating collar
  Heating collar: DIMECA (F) 034 H60 230V 5W
    Adapter to ensure close contact
    Tested at room temperature
Heating collar temperature measurement versus time plot and setup photo
Heating collar with cable
Aluminum adapter

## Slide 17: Setup with the heating collar
VGP testing setup with the heating collar at the EXCAL climatic chamber

## Slide 18: Tests with and without the heating collar
Test performed by increasing 10% humidity every 24h. Fixed 25°C starting at 40%RH up to 80%RH
Standard IKR070 VGP with heating collar on/off comparison measurements at the EXCAL machine
ΔI ~10pA Δ𝑃 ~2.3E-11 mbar
ΔI ~200pA Δ𝑃 ~1.0E-10 mbar

## Slide 19: Araldite® 2012 adhesive
Characteristics and features
Two-component epoxy adhesive (mix ratio 1 : 1)
Bonds a wide variety of metals and plastics
Rapid cure at room temperature (pot life approx. 5-8 min, fixture in ~20 min)
High strength and toughness
Self-levelling
Cons:
Maximum continuous operating temperature ~70 °C
Impossible to remove and to reopen the connector
Not easy to apply with precision
Applying the Araldite® 2012 adhesive at the VGP PEEK and LEMO socket connector
VGP with freshly applied adhesive on the left, and after the bake-out on the right

## Slide 20: Araldite® 2012 tests
Test performed at high (>80%RH) humidity and fixed 25°C during 14 days with pre and after drying periods
Standard IKR070 VGP with the Araldite® 2012 adhesive measurements at the EXCAL machine

## Slide 21: Conclusion/actions after the first set of tests
ELENA LNR.VGPB.0606 with heating collar
  Conclusion
  Leakage current correlation with humidity level confirmed
    IKR070 not according to the specs
  Actions
  Heating collar installed at ELENA ring VGP
    Significantly reduces the leakage current
  Findings summarized on a Technical Note (EDMS3126377)
    Manufacturer (Inficon) contacted

## Slide 22: Investigation by Inficon
  Inficon confirmed our findings
  They concluded that the humidity condensates at the atmospheric side of the ceramic feedthrough, generating a conductive layer at the surface level
Inficon slide (courtesy of R. Salemme)

## Slide 23: Solution proposed by Inficon
  Main requirements
  To seal mechanically the inner volume of the HV feedthrough
  Solution to be installed without the need of breaking vacuum
  Mounting should be possible at any gauge mounting orientation
  Proposed solution
  O-rings to seal the critical feedthrough volume
  Modified PEEK part to host the O-rings and a ferrule to guide
Standard IKR070 on the left, modified PEEK solution on the right
O-rings
Ferrule (cone)
Modified PEEK solution parts
  Pros
  Easy replacement / installation
  No need to break vacuum
  LEMO wall socket connector untouched
  Cons (due to the FKM O-ring)
  Continuous max. bakeout temperature reduced to 200°C resp. 250°C
  Radiation tolerance reduced to 105 Gy resp. 107 Gy

## Slide 24: Modified setup with the PEEK solution
VGP testing setup with two IKR070 gauges; Inficon modified PEEK on the left, standard PEEK with 1mm drilled hole on the right
VGP testing setup with two IKR070 gauges at the EXCAL

## Slide 25: Modified PEEK solution tests
Test performed at high (>80%RH) humidity and fixed 25°C during 14 days with pre and after drying periods
VGP setup measurements at the EXCAL machine with two IKR070 gauges; Inficon modified PEEK and standard PEEK with 1mm drilled hole
ΔI ~210pA Δ𝑃 ~1.1E-10 mbar

## Slide 26: Conclusion and future work
  Leakage current correlation with humidity level confirmed at the VGP
  Pfeiffer IKR070 / Inficon MAG070 not according to the specs
  Conditions to see the Leakage Current effect at the standard IKR070 measurements
  Very low vacuum pressure (E-11mbar range)
  High humidity (>60%RH)
    Machines on the surface with openings (e.g. AD/ELENA, LEIR)
    Gauges close to cooling circuits and no ventilation (e.g. LHC P5 – CMS yoke)
  CERN in-house solution – heating collar
  Requires 230VAC near the VGP
  Significantly decreases the leakage current
  INFICON solution – modified PEEK with O-rings
  Seals the HV feedthrough internal volume
  Easy installation
    Including retrofitting on installed gauges
  What now?
  10 modified PEEK units to be received soon
    To be installed at AD, ELENA and LEIR during YETS25/26
  Upon CERN acceptance*
    A part number will be assigned by INFICON for ordering the modified PEEK part solution
    A part number will be assigned by INFICON to order the MAG070 Penning gauge with the assembled modified PEEK
    *We need to acknowledge the max. bakeout temperature of 200°C and the radiation tolerance of 105 Gy

## Slide 27: Acknowledgments
Thanks to the main contributors; A. Sinturel, J.A. Ferreira, G. Pigny and ICM colleagues
Especial thanks to other colleagues for the productive conversations, shared knowledge, and collaboration;
TE-VSC
    IVO
    BVO
    VSM (S. Meunier)
    SCC (B. Teissandier)
    DLM (H. Rambeau)
BE-OP-AD
EP-ESE (C. Richard)
BE-CEM-EPR (S. Fiore)
INFICON AG (Roberto, Oskar and Tim)
...and many others

## Slide 28
Thank you for your attention !
Any questions ?

## Slide 29: NEG cartridge
  NEG was heated during the EXCAL climatic chamber tests
    Powered with ~10.12W (5.06V / 2.00A)

## Slide 30: IKR 070 disassembled

## Slide 31: MAG 070 gauge model with new PEEK adapter
  A groove milled around the external surface of new PEEK adapter allows visually recognising and distinguishing the new part
CAD model courtesy of R. Salemme

## Slide 32: Modified PEEK solution tests
Test performed at the b.30 ICM electronics lab under natural room conditions for both the controller and the gauges
VGP setup measurements at the b.30 lab with two IKR070 gauges; Inficon modified PEEK and standard PEEK with 1mm drilled hole

## Slide 33: VFDS shrink tube test
Test performed at high (>80%RH) humidity and fixed 25°C during 14 days with pre and after drying periods
Standard IKR070 VGP covered with the Raytronics VFDS shrink tube measurements at the EXCAL machine