## Slide 1: Status of North Area ISTs – Beam Instrumentation
Inaki Ortega on behalf of the SY-BI and BE-EA teams
01/04/2025 – 182nd EATM

## Slide 2
Overall, a smoother IST than last year (do you remember the issues with the gas for the wire chambers??)
Procedure:
Test full acquisition chain with a Sr90 source when possible, measure noise levels…
Test control systems: motorisation, high voltages…
XSCI:
  Some motor problems addressed with normal maintenance. One purple frame replaced.
XWCA:
  Replaced electronics of one XWCM in M2.
  Replaced a broken chamber in H6
  Some motor problems (shared with XSCI)
DWC:
  OK
FISC:
  Some hiccups with the motor movement
  3 more motors broken -> replaced with 3 last spares. From now on, new motors

## Slide 3
XION:
  We are testing the same readout as the East Area (from SY-BI-PM section).
EMC:
  Should be OK now.
CEDAR in H6:
  We are currently working on it
    New voltage dividers
    We have noticed degraded quality in cables (40 years old!)
  Second CEDAR installation: I propose we request new cables.
XCET:
  Infrastructure ready for new pressure sensor (Druck). They will be installed this week.
  Status by PPE:
    PPE 138 : installed, no gas
    PPE 148 : installed, CO2
    PPE 158 : installed, CO2
    PPE 152 : out of beam, He
    PPE 142 : out of beam, no gas
    PPE 136 : out of beam, He