## Slide 1
Characterization of a simple particle detector using cosmic particles
Sune Jakobsen (some material from Christian Joram)
Additional material: AIDAinnova 2nd Annual meeting, 2 lectures on introduction to photo detectors
Summer student workshop 2024
Program
Discussion of the detector parts
Scintillators
Wavelength shifters
Photodetectors
Experimental Part
Determine the 1 photo-electron charge and calculate the gain of a photomultiplier.
2a. Measure the light yield of a MIP in a scintillator coupled to a photomultiplier using a light guide.
2b. Measure the light yield of a MIP in a scintillator coupled to a photomultiplier using a wavelength shifter.
3. Measure the Quantum Efficiency of a PMT in the wavelength interval 200 to 800 nm (optional)

## Slide 2
Scintillators
scintillator
photodetector
Inorganic
(crystalline structure)
Up to 70000 photons per MeV
High Z (good for photoeffect Z5)
Large variety of Z and r
Undoped and doped
ns to ms decay times
Expensive
Fairly Rad. Hard (100 kGy/year)
E.m. calorimetry (e, g)
Medical imaging
Two categories
Organic
(crystals, plastics or liquid solutions)
Up to 10000 photons per MeV
Low Z (not good for photoeffect)
Low density r~1g/cm3
Doped, large choice of emission wavelength
ns decay times
Relatively inexpensive
Medium Rad. Hard (10 kGy/year)
Energy deposition by an ionizing particle or photon (g)
 generation
transmission of scintillation light
detection
2

## Slide 3
Plastic scintillators
Often they consist of a solvent + scintillator and a secondary fluor as wavelength shifter.
A fluor has its absorption and emission spectra shifted. The difference of the two peaks is called Stokes shift
Solvent
DE = dE/dx·Dx
Scintillator
Fast and local energy transfer via non-radiative dipole-dipole interactions (Förster transfer).
fluorescence light
wavelength
shifter (‘fluor’)
radiative transfer
UV
(~300 nm)
Visible
( 400 nm)
>
~
ionizing particle
3

## Slide 4
Wavelengh shifters - Two / One Dopant scheme
solvent + scintillator + wave shifter
Förster Radiative
solvent + large stokes shift scintillator Förster
Dopants in toluene: large Stokes shift dopants feature a much smaller self-absorption
0.025 molar sol.
0.0025 molar sol.
Abs. and emission spectra
A
B
A
B
B
4

## Slide 5
Scintillator readout
Readout has to be adapted to geometry, granularity and emission spectrum of scintillator.
Geometrical adaptation:
Light guides: transfer by total internal reflection (+outer reflector)
Wavelength shifter (WLS) bars
“fish tail”
adiabatic
5

## Slide 6
Photomultipliers
Photodiodes
Hybrid Photon Detectors
Gaseous Devices
Photodetectors
6

## Slide 7
Photodiode (PD)
  No internal gain, direct conversion of optical photons to electron-hole pairs that are simply collected. External amplification is required.
Avalanche PD (APD)
  They incorporate internal gain through higher electric fields that increases the number of charge carriers that are collected.
Geiger Mode APD (GM-APD = SiPM)
  They are operated with a voltage above the voltage breakdown. Breakdown is triggered by a photoelectron or a thermal electron. It has to be quenched.
Multipixel GM-APD
  They are made of many small area GM-APD on the same substrate and electrically connected in parallel.
Photodetectors
7

## Slide 8
Cosmic rays are energetic charged subatomic particles, originating in outer space. Most primary cosmic rays are protons, atomic nuclei, or electrons.
When cosmic rays enter the Earth's atmosphere they collide with mainly oxygen and nitrogen molecules  cascades of lighter particles, called air showers. Photons and electrons form electromagnetic showers. Muons don’t! In addition they have a long enough lifetime to reach the ground level. So, what we detect are mainly muons (and parts of e.m. showers).
The number of particles that hit the ground is dependent on several factors including location with respect to the earth's magnetic field, solar cycle, elevation, and the energy of the particles.
As a rule of thumb, the flux of muons above 1 GeV/c at sea level is :
F = 70 m-2 s-1 sr-1  1 cm-2 min-1
 “Cosmic muons” are very useful and cheap for testing particle detectors.
Cosmic Rays / Cosmic Radiation
8
In 2012, it’s just 100 years, that Victor Hess discovered the cosmic rays by means of electrometer (=ionization) measurements during balloon flights up to 5300m. He concluded "The results of my observation are best explained by the assumption that a radiation of very great penetrating power enters our atmosphere from above."  Nobel prize 1936.

## Slide 9
1 photo-electron
Task 1: Determine the 1 photo-electron charge and calculate the gain of a photomultiplier.
Materials:
Classic photomultiplier
Oscilloscope
Pulse generator
LED
High voltage power supply unit
Cables
9

## Slide 10
Light yield using light guide
Task 2a: Measure the light yield of a MIP in a scintillator coupled to a photomultiplier using a light guide.
Materials:
Classic photomultiplier
Oscilloscope
Scintillator
Light guide
High voltage power supply unit
Cables
10

## Slide 11
Light yield using wavelength shifter
Task 2b: Measure the light yield of a MIP in a scintillator coupled to a photomultiplier using a wavelength shifter.
Materials:
Classic photomultiplier
Oscilloscope
Scintillator
Wavelength shifter
High voltage power supply unit
Cables
11

## Slide 12
Absolute Measurement of the Quantum Efficiency of a Classical PMT
Task 3: Measure the Quantum Efficiency of a PMT in the wavelength interval 200 to 800 nm. Discuss the result, its precision and possible error sources.
Set-up:
- PMT
- Xe-lamp
- monochromator
- reference photodiode
- Keithley picoampere meter
- PC (Labview)
12

## Slide 13
Principle of the QE determination
(DUT = Detector Under Test
REF = Reference Detector)
Photon flux, unknown !
Use a reference detector with known (calibrated) eQ to determine the photon flux
13

## Slide 14
Xe lamp
monochromator
Calib. PD
PMT under test
mirror
14

## Slide 15
D1
K
Schematic representation of dynode structure
15
Iphoto
D1
K
+V

## Slide 16
Model | 818-UV
Detector Type | Semiconductor
Spectral Range | 200 to 1100 nm
Active Diameter | 1.13 cm
Detector Active Area | 1 cm2
Material | Silicon-UV Enhanced
Power Density, Average Max w/ Attenuator | 0,2 W/cm2
Power Density, Average Maximum w/o Attenuator | 0,2 W/cm2
Pulse Energy, Maximum - w/ Attenuator | 0,1 µJ/cm2
Pulse Energy, Maximum - w/o Attenuator | 0.1 nJ/cm2
Uniformity | ±2 %
Shunt Resistance | ≥10 MΩ
Calibration Uncertainty | 4% @ 200-219nm 2% @ 220-349nm 1% @ 350-949nm 4% @ 950-1100 nm
Calibration Uncertainty, w/ Attenuator | 8% @ 200-219nm 2% @ 220-349nm 1% @ 350-949nm 4% @ 950-1100nm
NEP | 0.45 pW/√Hz
Reverse Bias, Maximum | 5 V
Linearity | ±0.5 %
Connector Type | BNC
The 818-UV Low-Power UV Enhanced Silicon (Si) Photodetector is supplied with a NIST traced calibration report that details individual detector responsivity measured with and without attenuator over the 200 to 1100 nm wavelength range.
16

## Slide 17
17

## Slide 18
18