## Slide 1: Performance Evaluation of TOF-PET Readout Electronics: Progress & Contributions #4
Cesar Cardona*
Boston University | Group: EP-CMX-DD
Mentors: Carsten Lowis✲ , Giulia Terragni※
* sciences@bu.edu | ✲ carsten.lowis@cern.ch | ※ giulia.terragni@cern.ch
Cesar Cardona
April 8, 2025

## Slide 2: FastIC+ ASIC Team
Cesar Cardona
April 8, 2025

## Slide 3
FastIC+ ASIC Team
• Barcelona team (David Mazzanti) visited last week
• Gave hands-on tutorial on how the FastIC+ board works
• Explained signal chain and board features
• Boards had already been tested by them
• Now it’s our turn to use and evaluate them
Cesar Cardona
FastIC+ ASIC Team
April 8, 2025

## Slide 4
FastIC+ ASIC Team
• Start of hands-on collaboration
• Our mechanical setup worked perfectly
Cesar Cardona
FastIC+ ASIC Team
April 8, 2025

## Slide 5: Measurements
Cesar Cardona
April 8, 2025

## Slide 6
• Started by calibrating boards using SPTR
• SPTR = Single Photon Time Resolution — board’s response to single-photon events
• Used new laser test bench — built in one week
• Bench lets us simulate single photons with precise timing
• Calibration helps us understand ASIC timing performance before full CTR tests
Cesar Cardona
Measurements: SPTR
April 8, 2025

## Slide 7
• Setup for Coincidence Time Resolution (CTR) measurements
• Using two 2×2×3 mm LYSO:Ce codoped crystals (top-grade)
• Crystals face each other — single bar configuration
• Goal: measure time difference between photon detections
• Setup inside dark box, mechanically stable
Cesar Cardona
Measurements: CTR Bench Setup
April 8, 2025

## Slide 8
• To set threshold, perform ITH* scan → get staircase-like plot
• Y-axis: photon count rate | X-axis: internal threshold (ITH)
• Each step = additional photons triggering detection
• First step (8–10 mV here) = single-photon detection threshold
• Set threshold here for best time resolution
• Especially critical for BGO due to lower light yield
*(ITH stands for Internal Threshold. An ITH scan is a procedure where you vary the internal discriminator threshold of the ASIC to determine at what signal level (in mV or LSB) the system starts registering photon events.)
Cesar Cardona
Measurements: Threshold Scan
April 8, 2025

## Slide 9
LYSO Ce:Ca Preliminary Results
• Used David’s code for initial CTR extraction — may adapt later
• FWHM Coincidence Time Resolution (CTR): 87.6 ps
• Both crystals show good energy resolution
• Double peaks in spectra likely LYSO Lu-176 natural radioactivity
• Promising start for FastIC+ with codoped LYSO
Cesar Cardona
Measurements: Preliminary Results
April 8, 2025

## Slide 10
BGO Preliminary Results
• One LYSO codoped crystal (reference) + one BGO crystal
• BGO signal looks distorted — poor energy resolution
• DT distribution not symmetric → fit is unreliable
• Code still David’s — may improve with our own analysis
• Reported CTR: –252.3 ps, clearly incorrect
• Highlights challenge of working with low-light-yield crystals like BGO
Cesar Cardona
Measurements: Preliminary Results
April 8, 2025

## Slide 11
Matrix Setup
• Transitioned to LYSO 4x4 crystal matrix measurements
• Carefully coupling matrix to SiPM array
• Alignment is challenging — must ensure full optical contact
• Precision here is critical for uniform light collection
Cesar Cardona
Measurements: Matrix Setup
April 8, 2025

## Slide 12
• Setup: 4×4 LYSO crystal matrix on one side, single LYSO crystal (reference) on the other
• Matrix coupled to SiPM board, carefully aligned
• Matrix board connected to FastIC+ and mounted in dark box
• Used handwritten channel map to identify matrix channels
Cesar Cardona
Measurements: Matrix Setup
April 8, 2025

## Slide 13
Cesar Cardona
Measurements: Matrix ITH
April 8, 2025

## Slide 14
Preliminary Results
• Currently no proper way to analyze matrix data
• FastIC+ has 16 channels — split across 2 ASICs (8 channels each)
• Problem: channels are mirrored — e.g., channel 5 = ASIC1 CH5 + ASIC2 CH5
• This creates overlap/confusion in the data
• In the plot, one side appears clear while the other is jumbled
• Working on custom analysis to isolate individual ASIC channels
Cesar Cardona
Measurements: Matrix
April 8, 2025

## Slide 15: Lower Voltage Reference
Cesar Cardona
April 8, 2025

## Slide 16
More Reference Measurements
• FastIC-compatible SiPM saturates above 42V bias
• Initial reference scans done at 45–48V — not usable
• Need to repeat reference measurements at 40–42V
• Preparing new runs to match FastIC’s voltage limitations
• I'm a pro at coupling single crystals now.
Cesar Cardona
Lower Voltage Reference
April 8, 2025

## Slide 17: Analysis
Cesar Cardona
April 8, 2025

## Slide 18
ROOT Comeback
• Boards must be controlled from two separate terminals
• Our computer struggles with very large data volumes
• Reference board records every event → ~10⁸ events in 2 hours
• Storage and processing becoming a bottleneck
• Moving forward with ROOT-based analysis for efficiency and scalability
Cesar Cardona
Analysis
April 8, 2025

## Slide 19: Miscellaneous
Cesar Cardona
April 8, 2025

## Slide 20
Poster Reeboot
• Tasked with removing old posters from lab hallway walls
• Carefully scanned and uploaded them to the CERN Box
• Now digitally immortalized… until a high-energy EM pulse erases history
• RIP wall posters — your contributions will not be forgotten
• (At least not until the next intern clears the Box folder)
Cesar Cardona
Miscellaneous
April 8, 2025

## Slide 21
Cesar Cardona
Miscellaneous
April 8, 2025

## Slide 22
Edible Radioactivity
• Word got out I was Mexican…
• Naturally, this turned into a lab challenge
• Tried one drop of a 1,000,000 Scoville hot sauce — child’s play
• Confidence was high → added three drops
• Regret was… also high
Cesar Cardona
Miscellaneous
April 8, 2025

## Slide 23: Questions?
Cesar Cardona
April 8, 2025