## Slide 1: Status of East Area XSEC Calibration
Dani Metin, Aurélie Goldblatt, Michel Duraffourg, Federico Roncarolo
07 April 2025

## Slide 2: Secondary Electron Emission Chamber (XSEC)

## Slide 3: East Experimental Area Beam Lines and Powering

## Slide 4: XSEC Calibration factor and details
Name​ | 2024 Calibration factor​ | ​Notes (last updated Dec 2024)
T08.XSEC070-I​ | 1.0923e+08​ | 900pF --> 10 x signal foil + 11 x bias foil​
F61.XSEC023-I1​ | 1.15e+08​ | 2 x signal foil + 3 x bias foil​
F61.XSEC023-I2​ | 1.15e+08​ | 2 x signal foil + 3 x bias foil​
F62.XSEC022-I​ | 5.08e+08​ | 10 x signal foil + 11 x bias foil​
F63.XSEC008-I​ | 5.08e+08​ | 2 x signal foil + 3 x bias foil​
T8.XSEC095-I​ | 4.3697e+08​ | 900pF --> 10 x signal foil + 11 x bias foil​ ​
Signal depends on number of foils and construction1
1 Ravotti, F., Ollacarizqueta, A.G. and Ruffieux, R.B., 2022. Calibration of the T08. XSEC devices by means of Al-foil activation and BCTF measurements (No. EP-Tech-Note-2022-001).

## Slide 5: Calibration factors for F61, F62 & F63

## Slide 6: XSEC relative signal and Calibration Factor for F61 (2025)

## Slide 7: XSEC relative signal and Calibration Factor for F61 (2024)

## Slide 8: Today’s Conclusion for Calibration Factors
Name​ | 2024 Calibration factor​ | 2025 calibration factor (TBC) (95% BCT transmission target for F61) | ​ Notes (last updated Dec 2024)
T08.XSEC070-I​ | 1.0923e+08​ |  | 900pF --> 10 x signal foil + 11 x bias foil​
F61.XSEC023-I1​ | 1.15e+08​ | 6.39e+08 | 2 x signal foil + 3 x bias foil​
F61.XSEC023-I2​ | 1.15e+08​ |  | 2 x signal foil + 3 x bias foil​
F62.XSEC022-I​ | 5.08e+08​ | 1.06+08 | 10 x signal foil + 11 x bias foil​
F63.XSEC008-I​ | 5.08e+08​ | 6.39e+08 | 2 x signal foil + 3 x bias foil​
T8.XSEC095-I​ | 4.3697e+08​ |  | 900pF --> 10 x signal foil + 11 x bias foil​ ​
Must observe vacuum environment

## Slide 9: New Foil - T08.XSEC095

## Slide 10: Scan of new foil
Measurements taken morning of 13th March 2025
Position on foil moved from 0 to over 30mm away from center
Examples from logbook are shown below

## Slide 11: Scan of new foil - response

## Slide 12: Scan of new foil - response
Norm. XSEC signal only shifted by more than 2% when beam was on very edge of foil

## Slide 13: Comparison to more stable beam position
Changes are more dramatic than when the beam position is stable, but remain small

## Slide 14: Comparison to more stable beam position
Changes during beam position scans are similar in magnitude to those observed over weeks without beam position variation Will compare with another beam position test once the foil has aged

## Slide 15: Supplementary Slides

## Slide 16: 2024 XSEC

## Slide 17: Transmission for F61

## Slide 18: 2025 Vacuum for F61

## Slide 19: 2025 Vacuum for F62 and F63