## Slide 1: Development of BPMs for the Linac4
J. TAN, M. SORDET, L. SØBY, F. GUILLOT-VIGNOT, D. GERARD, M. LUDWIG, D. STEYAERT
BE/BI
Linac4 Instrumentation Review
18th October 2011

## Slide 2: Outline
BPM
  Layout
  Functional specs
  The monitor and low beta beams
STATUS
  Drawings and prototyping
  BPM test bench and Acquisition chain
  First BPM characteristics
  Todo list
Spare policy
Planning
Summary

## Slide 3: BPMs layout and specs
Line | # of Monitors | Beam position |  | Rel. Beam Int. | Beam phase | TOF | Comments
 |  | Resolution | Accuracy |  |  |  | 
L4D | 2 | 0.1mm | 0.3mm | 1% wrt to peak current | 0.5 | 1 per mille | 
L4C | 7 |  |  |  |  |  | 
L4P | 6 |  |  |  |  |  | 
L4T to PSB | 27 |  |  |  |  |  | 
Observable 1: Beam position
Observable 2: Relative beam intensity between two BPMs
3 MeV
50 MeV
102 MeV
160 MeV
Transfer Line to PSB
H- source
LEBT
RFQ
MEBT (chopper)
DTL
CCDTL
PIMS
Dump Line
L4L ~ 12 m
L4D ~18 m
L4C ~22 m
L4P ~14 m
L4T- LT- LTB-BI ~178 m
L4Z ~10 m
Observable 3: Beam phase wrt RF reference
Observable 4: Energy via TOF between two BPMs

## Slide 4: Linac4 BPM (1)
Shorted Stripline
Linear
Fair sensitivity
Compact
Reasonable price
Resonating
b-dependent
Shorted
45
Z0= 50

## Slide 5: Dealing with low-b beams
Movable test bench BPM
Transfer impedance changes with b and bunch length
PU sensitivity changes with b
These simulations confirm
Shafer’s theorem (1994)
The values obtained from the simulations will be implemented in the software
Sensitivity at 352MHz
b

## Slide 6: Outline
BPM
  Layout
  Functional specs
  The monitor and low beta beams
STATUS
  Drawings and prototyping
  BPM test bench and acquisition chain
  First BPM characteristics
  Todo list
Spare policy
Planning
Summary

## Slide 7: Multiple BPM designs
Transfer Line to PSB
H- source
LEBT
RFQ
MEBT (chopper)
DTL
CCDTL
PIMS
Dump Line
L4L ~ 12 m
L4D ~18 m
L4C ~22 m
L4P ~14 m
L4T- LT- LTB-BI ~178 m
L4Z ~10 m
Beam aperture: 34 mm
Electrode length: 72 mm
Quad
DTL
Beam aperture: 34 mm
Electrode length: 117 mm
TRANSFER LINE
Beam aperture: 100 mm
Electrode length:140 mm
CCDTL and PIMS
Beam aperture: 39 mm
Electrode length:117 mm
3 BPMs with welded feedthroughs onto the body
39 BPMs with feedthroughs mounted on CF flanges

## Slide 8: Movable Test Bench BPMs
Modular commissioning of RFQ, MEBT and DTL tank1
Transverse : profiles, emittances, halo, position
Longitudinal : transmission, av. energy, halo, DE/E, bunch phase spread
Three BPMs
Beam aperture: 67 mm
Electrode length: 60 mm

## Slide 9: Prototype of a DTL BPM (1)
Body
Electrode
Feedthrough
detail
Mounted unit with
Welded feedthroughs
Feedthrough

## Slide 10: Prototype of a DTL BPM (2)
Issues
During e- bombardment welding
Three feedthroughs affected by :
Manufacturer’s procedure:
glass-ceramic seal
No sand blasting
Ultrasonic cleaning +acetone rinsing before
Scanning electron microscopy
Standard 304 grade stainless steel
Surface not inlayed with dust particles
Traces of Si, Na, Mg, Ca
Welded
feedthroughs
Sparks
Material spray
Multi-pass welding needed

## Slide 11: BPM test bench
Wire technique w/ standing waves
Optical sensor for mechanical centering : 0.01mm
352MHz sine wave with a N.A.
Data acq by steps of 0.1mm
Commissioning issues
Black anodized coating =poor grounding.
Solution : Improve grounding. Done
Wire not terminated = Antenna
Solution : terminate the wire by a short circuit. Done
Optical rail : C-shape & 25mm cross section yield a lack of stiffness and 1mm electrical offset
Solution : Swap to a full frame of 34mm cross section. Pending action.

## Slide 12: Analog Front End Board
FEATURES
Input Low pass filtering @1GHz
Down-mixing with LO
Variable gain with switchable attenuators
Band-pass filtering @22MHz
BPM Calibration
LO
U
D
L
R
A
D
C

## Slide 13: Signal Processing
IR QR –IR -QR
VR, ADC
VL, ADC
time
22.0125MHz Carrier Signal
Tcarrier =45.42ns
IL QL –IL -QL
Clock jitter ~10ps
MR-ML
MR+ML+MU+MD
Dx =
 = beam - LO
I = MR+ML+MU+MD
TOF = PU2 - PU1

## Slide 14: Stripline Characteristics
Beam aperture= 67mm
Electrode length = 60mm
 | THEORY | H plane | V plane
Slope [mm] | 40.32 | 41.79 s=0.42 | 41.02 s=0.56
Elec. Offset [mm] | - | -0.77 s=0.18 | -0.19 s=0.13
Coupling w/ adjacent electrodes | 7.3% | 6.56 | 6.6

## Slide 15: First Measured characteristics
352MHz signal injected via the wire
Measured voltages amplitudes from electrodes : ~65mV
From the simulations, this is equivalent to a signal induced by a
  nominal beam of 1.14x109 H-
  98 longitudinal phase (nearly debunched beam)
Position (averaged over pulse length 740ms  ~32k positions)
  Precision= 0.1mm
  Resolution = 0.1mm
  Resolution with few samples : to be done
RF phase wrt LO
  Precision = 1 (room for improvement)

## Slide 16: Todo List
Understand welding issues
Improve stiffness of lab test bench
Transfer impedance and characteristic impedance
Resolution with short pulses
Improve ADC clock jitter : the goal is ~200fs
Commissioning of BPM calibration procedure
Software : algorithm for I, Q, -I-Q determination
Measurement with a pair of BPMs
  TOF
  Relative beam intensity
Signal processing : Choose between
  SIS3302 (16bits – 13.4ENOB) : 6400CHF per board x25
  FMC (14bits – 11.7ENOB) : Status ? 25 boards. Unknown final price

## Slide 17: Outline
BPM
  Layout
  Functional specs
  The monitor and low beta beams
STATUS
  Drawings and prototyping
  BPM test bench and Acquisition chain
  First BPM characteristics
  Todo list
Spare policy
Planning
Summary

## Slide 18: Spare policy
Monitors
  Movable Test Bench : 3 +2 spares for lab commissioning
  Linac : 15 + 11 spares (one per version)
  Transfer line : 27 + 2 spares
  TOTAL : 55 BPMs
Acquisition chain
  Front end board : 45 + 15 spares
  ADC or FMC : 23 + 4 spares

## Slide 19: Planning
Start of Machine Installation
Start of TL construction & Installation
LS1
Start of Machine Commissioning

## Slide 20: Summary
Very encouraging results yet still a lot of work ahead of us.
Ready for the Movable Test Bench commissioning
Planning in phase with Linac4 installation and LS1
Budget sligthly underestimated : overrun by ~200kCHF

## Slide 21: Thank you for your attention
…Questions ?

## Slide 22: Linac 4 – Basic architecture
 | Movable test bench 3-12MeV | DTL 50MeV | CCDTL 102MeV | PIMS 160MeV | New TL -> LT.BHZ20 160MeV | Old TL -> PSB 160MeV
Number of BPMs | 3 | 2 | 7 | 6 | 10 | 17
Beam Ap. [mm] | 67 | 34 / 39 | 39 | 39 | 100 | 100
Long. Phase [ rms] | 6.4 /98 | 3 | 3 | 2.5 | 25 | 55
RMS length [ps] | 50 / 777 | 24 | 24 | 20 | 200 | 433
Bunch spacing : 2.84 ns

## Slide 23: Simulation of non-relativistic beams
Nominal beam phase width evolution (1RMS)
Simulations with CST Particle Studio
  Multi-bunch
  Space charge effect
  Low 
 | MEBT | Linac4 | PSB strip. foil
Sigma [ps] | 200 | 20 | 433
Rel. beta | 0.0798 | 0.52 | 0.52

## Slide 24: High Frequency Losses
Power flow module
Electric field module
External diameter of the
RF feedthrough is important

## Slide 25: High Frequency Filter
gap 0.5mm
Body
Electrode
Capacitive effect : 9pF
C