## Slide 1: Performance Evaluation of TOF-PET Readout Electronics: Progress & Contributions #5
Cesar Cardona*
Boston University | Group: EP-CMX-DD
Mentors: Carsten Lowis✲ , Giulia Terragni※
* sciences@bu.edu | ✲ carsten.lowis@cern.ch | ※ giulia.terragni@cern.ch
Cesar Cardona
April 29, 2025

## Slide 2: April 6 – April 25
Cesar Cardona
April 29, 2025

## Slide 3: Fixing CTR Bench
Cesar Cardona
April 29, 2025

## Slide 4
The CTR Bench broke.
This is not the bench we are using (we use a double sided CTR bench which can also do normal ctr runs, and the FastIC+ bench)
The bench seems to fry the boards and SiPMs. It has broken 3 that I know of.
I volunteered myself to fix with Thomas Meyer. Mostly to learn how the electronics work.
Cesar Cardona
Fixing CTR Bench
April 29, 2025

## Slide 5
Cesar Cardona
Electronics Circuit Diagram
April 29, 2025

## Slide 6
Cesar Cardona
6V Power Supply Problems?
April 29, 2025
Checked to see if Power supply had any spikes when:
Turned On/Off with the button.
Turned On/Off with physical switch.
Dis/Connected directly off the wall.
On/Off Button
Plug (offscreen)
On/Off Switch
Power Supply
Scope

## Slide 7
Cesar Cardona
6V On/Off Button
April 29, 2025
Good board. On Button.
Good board. Off Button.
50 ms
1 V
50 ms
1 V

## Slide 8
Cesar Cardona
6V On/Off Switch
April 29, 2025
Good board. Off then On Switch.
2 s
1 V
Power supply resets when switched off

## Slide 9
Cesar Cardona
6V Disconnected from Wall
April 29, 2025
Good board. Disconnected.
200 ms
1 V
Good board. Dis then Connected.
5 s
1 V

## Slide 10
Cesar Cardona
6V Power Supply on Bad Board
April 29, 2025
There are 3 Broken Boards that I can check.
Same procedure as with the good board, same cables even.
I have been told that all three have a dead time branch.

## Slide 11
Cesar Cardona
6V On Button Bad Boards
April 29, 2025
Bad Board #1. On Button
50 ms
1 V
6V
5.3V
Current is limiting voltage.

## Slide 12
Cesar Cardona
6V On Button Bad Boards
April 29, 2025
Bad Board #2. On Button
Bad Board #3. On Button
50 ms
1 V
50 ms
1 V
6V
6V
4.75V

## Slide 13
Cesar Cardona
Time Branch
April 29, 2025
Good Board Time Branch (35V Biased)
10 ns
20 mV
Bad Board #2 Time Branch (35V Biased)
10 ns
20 mV

## Slide 14
Cesar Cardona
Capacitors? Amplifiers?
April 29, 2025
When measuring resistance across capacitors C11, C12, and C13, each board reads a different resistance.
Board | Resistance
Good Board | 487 Ω
Bad Board 1 (V2-C001-A3) | 76 Ω
Bad Board 2 (V2-C003-A2) | 377 Ω
Bad Board 3 (Caseless) | 45 Ω

## Slide 15
So what's wrong with the bench?
We have discounted the power supplies and energy blackouts as a possible culprit.
What takes the boards out seems to be cheap electronic components mixed with user error.
More test will take place.
Cesar Cardona
Conclusion?
April 29, 2025

## Slide 16
We got new boards!
Just this morning we received 5 new boards. I checked their resistance over the capacitors and found this:
We will change components U1 and U2 from Bad Board 1.
Cesar Cardona
Another Update
April 29, 2025
Board | Resistance
Good Board | 487 Ω
00-43455-e1-0001 | 481 Ω
00-43455-e1-0002 | 482 Ω
00-43455-e1-0003 | 480 Ω
00-43455-e1-0004 | 480 Ω
00-43455-e1-0005 | 483 Ω

## Slide 17: FastIC+ Updates
Cesar Cardona
April 29, 2025

## Slide 18
Deadline Thursday May 8, 2025
We are submitting an abstract for conference.
Topic being:
Exploring Scalable Electronics for Heterostructured Scintillators Using FastIC+
Cesar Cardona
Writing Abstract for IEEE Conference
April 29, 2025

## Slide 19
We have gotten mostly used to it…
FastIC+ records a lot of Data (~10 GB per single pixel 20 minute run).
Most of it is useless, but we still get considerable statistics on coincidences.
Since the FastIC+ boards output .txt files, we need to use a BEAST computer to first filter the data before we can do analysis in our laptops.
Cesar Cardona
FastIC+ Operation
April 29, 2025

## Slide 20
Heterostructure Timing Studies with BGO + EJ232
We are investigating a heterostructure crystal composed of BGO and EJ232 (3×3×20 mm), paired with a reference LYSO:CeCa crystal (2×2×3 mm) on a Broadcom NUV-MT SiPM at 40 V.
Goal: Improve Coincidence Time Resolution (CTR) while maintaining high detection efficiency.
Motivation for heterostructures:
BGO: High stopping power, but slow scintillation.
EJ232: Very fast light emission, but low stopping power.
Combining them could yield the best of both worlds: better timing + good sensitivity.
Cesar Cardona
Detector Setup and Motivation
April 29, 2025

## Slide 21
What is Event Classification?
In heterostructures, gamma rays may interact in:
Just the BGO, Just the EJ232, Or both (shared events)
We use signal features (e.g., rise time, amplitude, time-over-threshold) to classify the event type.
Classification enables us to isolate fast events and study their timing separately.
Cesar Cardona
Event Classification
April 29, 2025

## Slide 22
DOI (Depth of Interaction)
We are also using the FastIC+ to study DOI
I helped by designing a 3D mount to set the crystals at different orientations.
Two techniques used:
Double Sided Readout (Carsten)
Light Sharing Matrix (Giulia)
Cesar Cardona
Other Measurements
April 29, 2025

## Slide 23
In standard PET, detectors assume gamma photons always interact near the front of the crystal, but they can actually interact deeper inside, leading to parallax error.
Parallax error worsens as the angle between the incoming photon and the detector face increases — especially in systems with large ring diameters.
Depth of Interaction (DOI) techniques aim to measure the exact depth at which the photon interacts inside the crystal.
Cesar Cardona
Why DOI Matters in PET?
April 29, 2025
vs
SiPMs
Scintillators

## Slide 24
The FastIC+ got a little hot…
When using matrices, the current is high enough that it causes the SiPMs to get a little toasty, and melt the Meltmount supporting the crystals… Not Good.
We decided to build a fan to keep things cold.
We came up with many ideas as to how to mount it, but then came my genius!!!
Cesar Cardona
My Best Work Yet…
April 29, 2025

## Slide 25
Only 3D printed a 10g. Brace, and used parts from the lab to create this beauty!
But, it was too short to provide the SiPM with good airflow…
Until…
Cesar Cardona
My Best Work Yet…
April 29, 2025

## Slide 26
CARSTEN’S GENIUS ARRIVED!!!
Cesar Cardona
Carsten to the Rescue
April 29, 2025

## Slide 27: Questions?
Cesar Cardona
April 29, 2025