## Slide 1: Enabling Novel Sensor Designs with Timepix
Medipix and Timepix symposium
22 September 2025
Martin van Beuzekom

## Slide 2: Long history of sensor development with Timepix1-4
InGrid/Gridpix on Timepix1
Silicon
Sensor
Timepix4
Wirebonds
We use Timepix primarily for charged particle tracking
But also some photon (visible & X-ray) detectors
Focus changed from Gaseous to Semiconductor sensors (mainly silicon)
Silicon sensor on Timepix4
LHCb VELO sensor
qualification with Timepix3

## Slide 3: Why timing: 4D tracking
Timing
2 ns → 30 ps
© LHCb collaboration
Single bunch crossing
Future accelerators have an increasingly complicated tracking environment
Time is important for the track reconstruction; separation of tracks in time
Fewer tracks per time ‘frame’ → simplifies reconstruction
Courtesy Tim Evans
0.5 mm
2 meter

## Slide 4: Requirements for 4D trackers
Time resolution of order 50 ps per hit/pixel
Spatial resolution of 10 mm
High readout bandwidth of tens of Gbit/s/cm2
Radiation >20 MGy (5x1016 1 MeV neq /cm2)
Large systems with 107 to 109 pixels
Up to several square meters
of sensitive area
LHCb Vertex Locator
(Upgrade-II)
1 m

## Slide 5: Silicon pixel sensors
Fast silicon sensors consist e.g. of p-type silicon bulk with n-type pixel implants
Effects relevant for good time resolution mostly happen below the surface
Pixels have a non-zero thickness -> ‘voxel’
Timepix

## Slide 6: Time resolution
Need:
Large signal
Short rise time
Low threshold
Need:
High Signal-to-Noise
Short rise time
Limited by pixel
area and power
Timewalk [ns]
Signal [a.u.]
σt2(sensor) +
σt2(jitter) + σt2(timewalk)
+ σt2(TDC)
Time resolution =
[K. Akiba, DOI 10.1088/1748-0221/18/02/P02011

## Slide 7: Time measurement with Timepix4 ASIC
Base clock of 40 MHz
Hit starts local 640 MHz oscillator (VCO)
  1.56 ns bins nominally
  Count #clock cycles = fToA (count)
  VCO per group of 2x4 pixels
Oscillator is stopped by first rising edge of 40 MHz clock
Time-over-Threshold is representative for signal amplitude
In addition the internal state of ring oscillator is captured –> 195 ps bins (56 psrms)

## Slide 8: Does Timepix4 have a good resolution out of the box?
The nominal VCO frequency is 640 MHz
However, there is a significant frequency spread across the matrix (VCOs are not identical)
Without correction, time resolution can increase by by hundreds of picoseconds
-> Have to (be able to) determine the VCO frequency of each pixel
Col
Row
Frequency [MHz]
400 ps
10 MHz difference
= 400 ps over 25 ns
25 ns
arXiv:2503.15207
K. Akiba, arXiv:2503.15207

## Slide 9: How we test our sensors

## Slide 10: Timepix4 beam telescope at SPS H8 beam line

## Slide 11: Telescope performance
180 GeV/c mixed hadron beam
2-3 mm pointing resolution
  Allows intrapixel studies
12 ps time resolution with MCP
  ~ 90 ps with timepix4 only
4D-tracking demonstrator
Timepix4

## Slide 12: How sensors affect timing

## Slide 13: A closer look at the sensor: signal induction
Current is induced in the electrodes when charges move
Both the electrons and holes induce current
described by Ramo-Shockley theorem
Weighting field Ew describes the coupling between location of the charge and the collecting electrode(s)
Ew defined by the geometrical arrangement of electrodes
= 1

## Slide 14: How to get good time resolution (electronics viewpoint)
Time resolution
Minimise capacitance C
  Small pixel -> can fit less functionality per pixel
  Small electrode -> less uniform fields -> more sensitive to drift time differences
Maximise signal Q in shortest possible time (dQ/dt = current)
  Short drift distance, high drift speed (large bias voltage), avalanche gain
Large power P
  Last resort, ‘costly’ in terms of material needed to provide power and cooling

## Slide 15: Overview of different hybrid sensors
Small signal if thin
Medium capacitance
Reasonably rad. hard
Planar
Large signal (gain)
Medium capacitance
Modest rad. hard
Low Gain Avalanche Detector (LGAD)
Large signal (thick)
Rad. Hard (shown 3E16)
High capacitance
More material
3D

## Slide 16: What can be achieved with planar sensors?
Mean time resolution of the TPX4 analog front-end versus injected charge
Planar (electron collecting) sensors on Timepix4 as benchmark
Though thicker sensors give more signal, time resolution does not improve
  Drift time is too long; ‘extra’ charge arrives too late
Very thin sensor do not provide enough signal; front-end jitter dominates
[K. Heijhoff, DOI: 10.1088/1748-0221/17/07/P07006]
50 mm
100 mm
200 mm
Testpulse
No sensor
Preliminary

## Slide 17: Low Gain Avalanche Detector (LGAD)
Gain layer
Avalanche multiplication in a high electric field
High E-field from high (acceptor) doping level
  Sensitive to radiation damage
Electrons ( ) accelerate, collide and liberate extra electrons, which in turn liberate more electrons
Most of the induced current is from the
backward travelling holes ( )
Aim for gain of order 10
Atlas / CMS achieve ~30 ps resolution with
1 mm ‘pixels’ in a 50 mm thick sensor
Accelerate
Accelerate
Accelerate
Accelerate
Collide
Collide

## Slide 18: What about LGADs with small pixels?
Regular LGADs have a ‘dead’ (no-gain) region between pixels of > 50 mm
Junction termination extension (JTE) because of high E-fields
  Not a problem for ATLAS/CMS size pixels
  But very low efficiency for small pixels
Alternatives: trench isolation (TI-LGAD), inverted LGAD
'regular’
(JTE + p-stop)
Trench isolated
Inverted LGAD
JTE
p-stop

## Slide 19: TI-LGAD on Timepix4
Tested multiple sensor variants
  1 or 2 trenches
  Different distance gain layer to trench
  Different doping levels, Carbon co-implantation (radiation hardness)
here they are
Gain of ~11
Best time resolution achieved is 99 ps
  For the innermost 40% of the pixel
Expect improvement by cooling the sensor, and higher voltage operation
Top view, single pixel
1 and 2 trenches
Preliminary

## Slide 20: Inverted LGAD (provided by Glasgow University)
Continuous gain layer on back side -> uniform response across pixel
Hole collecting devices -> Timepix4 front-end not optimal
Currently we only have 250 mm thick sensors, with a gain of only ~4
  Hence time resolution is not great, ~370 ps
  Unfortunately, two thin sensor productions failed
These are still very interesting devices for studies, not only timing
No gain region
Gain region
Am-241 spectrum
14 mm
14 mm
Preliminary
Preliminary

## Slide 21: 3D silicon sensors
The electrodes of 3D sensors penetrate the silicon bulk
Electron-hole pairs drift perpendicular to track direction
  Track length (= signal magnitude) decoupled from drift distance
O(5 mm) holes etched via Deep Reactive Ion Etching (DRIE), then doped
Shown to be sufficiently radiation hard

## Slide 22: 3D sensor on Timepix3 and Timepix4
Recent sensor batches suffer from excessive leakage current
Hence, transplanted old 3D sensor (2012) from TPX3 to TPX4 (by Advafab)
Unfortunately hole collecting, and large diameter pillars
Timepix3
Timepix4

## Slide 23: 3D time resolution
New sensor productions are being launched
Narrower and shorter pillars (lower capacitance)
Electron collecting
Different layouts
Preliminary
Preliminary

## Slide 24: Summary
The Timepix family provides an excellent platform for testing novel sensors in realistic conditions
  Large number of pixels, which brings corresponding challenges
  In addition to ‘single pixel’ performance tests
Why Timepix?
  ASICs with superior performance, thanks to a world class design team
  Many operation modes, tuning options -> flexibility / versatile
  Excellent support
  Benefits of a large and very active collaboration
What’s next?
  More sensors with gain: thin iLGAD, TI-LGAD, AC-LGAD, SPAD, SiEM
  3D sensors with lower capacitance, and possibly with gain
  Other sensor material such as SiC

## Slide 25: Thank you!

## Slide 26: Backup slides

## Slide 27
Timepix4 telescope configuration
12 ps time res.
Writing paper with performance of Timepix4 telescope

## Slide 28: Timepix4 and SPIDR4 readout
Sensor
Timepix4
Wirebonds
512x448 pixels, 55 um pitch

## Slide 29: Simulation and example from a real sensor
Disconnected pixel
3D sensor, influence of capacitance on jitter
With testpulses
R. Balabriga, https://doi.org/10.1016/j.nima.2022.167489

## Slide 30: iLGAD at grazing angle
Current iLGADs are 250 µm thick
Perpendicular tracks not ideal for good timing:
  Initial signal is small , and amplified signal is late / slow
  And TPX4 threshold can not be set high enough to trigger on amplified signal
  And, very non-uniform weighting field due to small pixels
Hence using grazing angle tracks to (partially) mitigate thickness effects
E. Dall'Occo et al 2021 JINST 16 P07035
[DOI: 10.1088/1748-0221/16/07/P07035]

## Slide 31: Timewalk as function of depth in iLGAD
Timewalk curve is now a strong function of depth, while charge distribution is the same
  For tracks near pixel implant, curve is dominated by small (non amplified) electron signal
  For tracks near gain layer, curve is dominated by amplified hole signal
preliminary
preliminary
preliminary