## Slide 1: Post-Mortem report – Experimental Areas
Inaki Ortega on behalf of the Experimental Areas team
26/06/2025 – SY-BI Technical Board on Post-Mortem reports

## Slide 2: East Area
Smooth commissioning overall, thanks to the ISTs, which were enormously helpful. We can use radioactive sources during the IST, allowing us to test the entire system, both hardware and acquisition chains.
XBPFs: OK
Scintillators: OK
BCCGA (Longitudinal spill monitor): relocated to a new position upstream, closer to the PS -> worked well since the very beginning.
MWPC in IRRAD and CHARM: OK
XCETs: new cables installed to simplify synchronization with beamline scintillators -> worked OK since beginning.

## Slide 3: North Area
Smooth commissioning, like for the East Area. We use the same procedure for the ISTs: radioactive source to test entire system
MWPC:
  Replaced one set of local electronics in TT83, damaged by radiation. The number of spares is good.
  Replaced a broken chamber in the H6 beamline.
  We had a few issues with the in/out motorisation, which is shared with the scintillators.
Delay Wire Chambers: OK
FISC:
  Some issues with motor movement, which were caused by incorrect addressing of VME boards. It occurred after the boards were removed without updating the changes in CCDE.…
  3 motors failed and were replaced with the 3 last spares. From now on, we will use the new motors: everything is already prepared for this (electronics and FESA are ready)

## Slide 4: North Area
Scintillators: OK, except for in/out motorisation issues shared with the MWPC
XION: OK. We are doing tests with the readout of the East Area XSEC, from the SY-BI-PM section
Electro-magnetic calorimeter: Hiccups with control system. OK.
CEDARs:
  CEDARs in the H2 beamline: Operational and critical for the experiments. Users are satisfied.
  CEDAR in H6: Encountered issues with 40-year-old voltage dividers. Only 4 out of 8 PMTs could be made to work, so this CEDAR could not be commissioned.
    Mitigation: New voltage dividers and PMTs have been purchased and have recently arrived.

## Slide 5: North Area
XCET (Cherenkov Threshold Counters): OK
  Installed new digital pressure sensor, replacing the 40-year-old Hamilton pressure sensor.
    Required cabling modifications, new electronics…
    Successfully completed for the 8 XCETs in the EHN1 hall.