## Slide 1: Equipments for electrical tests in SM18
By S. Juberg and G. Pichon
15/05/2025

## Slide 2: Summary
Test plan
High voltage tests
Capacitive discharge and AWT CD
Transfer function and AWT TF

## Slide 3: Electrical tests in SM18
Test plan
    HV tests :
    At reception
    After installation on the test bench
    At cold before and after the powering tests ( repeated if thermal cycle )
    At 100 K during the final warmup
    At room temperature before departure
    Transfer functions and capacitive discharges: Once at warm and twice at cold.

## Slide 4: High voltage tests
Stand alone Megger S1-568 (or MIT-525)
    Positive terminal directly connected to the equipment to be tested (Busbars or vtaps).
    The other circuits connected to the ground:
    along the test bench
    On the cold mass
    IR(t) or SV

## Slide 5: High voltage tests
Stand alone Megger S1-568 (or MIT-525)
    Check if the circuits are properly connected:
    At the minimum voltage of the Megger (40 V)
    A multimeter is connected on the circuit and the ground
    => Voltage measured = Voltage Megger
    => Insulation resistance displayed = multimeter impedance
Connection check

## Slide 6: High voltage tests
Stand alone Megger S1-568 (or MIT-525)
    Discharge after test :
    Internal load of the Megger
    all circuits grounded after each configuration
Connection check

## Slide 7: High voltage tests
Stand alone Megger S1-568 (or MIT-525)
    The data are recorded with a laptop through the application of the manufacturer (USB connection) : PowerDB
    One point per second
    Plot the graph in live
PowerDB interface

## Slide 8: High voltage tests
Stand alone Megger S1-568 (or MIT-525)
    The data are exported in a .csv file :
    One point per second
    Date
    Relative time
    Voltage measured
    Current
    Resistance
Raw data

## Slide 9: High voltage tests
Stand alone Megger S1-568 (MIT-525)
    Data conversion with EMU application in DIADEM
    ( National Instrument software ) :
    Generate :
    Plots in *.PNG files
    A text file to export to the data base Carpenter (QA)
Post-processing

## Slide 10: High voltage tests
High voltage rack
Composition :
  PXI National instrument :
  Remote control of the Megger
  Commutation of circuits
  Data acquisitions
  2. Electronic crate
  Interface between the PXI and the relay box
  3. Megger S1-568
  4. Relay box : 16 channels / 3 positions (HV+;HV-;Float)
1
2
3
4

## Slide 11: High voltage tests
High voltage rack
Automatization of the tests
Repeatability
Interlock with the PLC of the test bench
Cycle (ramp + plateau)
LabVIEW software programmation and support by BE-CEM
Labview interface

## Slide 12: High voltage tests
High voltage rack
Configuration
Selection of the polarity for each output
Test voltage at cold or at warm
Maximum current allowed
Expected capacitance
Cycle : ramp + plateau
Labview interface

## Slide 13: Capacitive discharge test
CDG-7000
CDG-7000 device
CDG-7000 with 7 kV max output voltage, 2 kA max output current and 49 J max pulse energy
Tests performed up to 600 V
Tester connected to the IFS boxes to discharge in full coil, coil pairs and individual coils
Manual setup, configuration and saving on USB stick
Prone to human errors during tests

## Slide 14: Capacitive discharge test
AWT – Automatic Winding Tester CD
Project initiated in 2024 to automate, save time and free operator to perform other tests
Thanks to AWT team: Unai Martinez Hernandez & Roberto Gabriel Dias Goncalves
Remote control of CDG 7000
Automatic configurations and saving process
Relays and driver board for individual relay control

## Slide 15: Capacitive discharge test
AWT – Automatic Winding Tester CD
Independent controlled relays by driver board
Modular relay chassis 1:128
Cable adaptors for each magnet type
NI I/O module, DAQmx, Modbus and LabVIEW environment
Multimeter

## Slide 16: Capacitive discharge test
AWT – Automatic Winding Tester CD
Simple GUI for user
Easy setup, “Plug and play”
Outputs dataset in .txt and graphs in .png
NI I/O module, DAQmx, Modbus and LabVIEW environment

## Slide 17: Capacitive discharge test
AWT – Automatic Winding Tester CD

## Slide 18: Transfer function
LCR meter : HIOKI IM3533-01
4-wires measurements
Manual setup, configuration and saving on laptop
Prone to human errors during tests
Connected through a test box
Prone to human errors during tests
Controlled with HIOKI software
HIOKI connection

## Slide 19: Transfer function
AWT TF
Same HW setup as AWT CD with double amount of relays
Premade drivers from manufacturer was not robust so new home-made drivers in progress
Same SW with almost similar setup as AWT CD
Connected through a test box
Controlled with HIOKI software

## Slide 20: Transfer function
HIOKI IM3533-01 & AWT TF
Data file HIOKI and AWT TF : .csv or .bpm
Frequency, impedance and phase
HIOKI data

## Slide 21: Transfer function
HIOKI IM3533-01 & AWT TF
EMU
Merges all the data files
Creates plots
Data compilation