## Slide 1: Performance Evaluation of TOF-PET Readout Electronics: Progress & Contributions #6
Cesar Cardona*
Boston University | Group: EP-CMX-DD
Mentors: Carsten Lowis✲ , Giulia Terragni※
* sciences@bu.edu | ✲ carsten.lowis@cern.ch | ※ giulia.terragni@cern.ch
Cesar Cardona
May 13, 2025

## Slide 2: April 27 – May 10
Cesar Cardona
May 13, 2025

## Slide 3: CTR Board Fix
Cesar Cardona
May 13, 2025

## Slide 4
• Out of 5 TOF-PET readout boards, only one was fully functional by May 9.
• Boards began failing progressively — symptoms first appeared in the timing branch.
• I developed a systematic, multi-step diagnostic method to identify, classify, and repair faulty boards.
Cesar Cardona
Initial Board Failures and Diagnostic Approach
May 13, 2025
Example of a TOF-PET readout board.

## Slide 5
• The first sign of board failure was a missing timing signal.
• I compared a known-good board (Board-01) with the four failing boards.
• Oscilloscope traces clearly distinguished functional timing behavior from flat/noisy output.
Cesar Cardona
Timing Branch: Good vs. Broken
May 13, 2025
Board-01: Normal Timing Response
Board-05: Broken Board No signal
10 ns
20 mV
10 ns
20 mV

## Slide 6
• The first diagnostic step was suggested by Thomas Meyer: measure the resistance across capacitors C11, C12, and C13.
• This quick test allowed me to isolate suspect boards immediately.
Cesar Cardona
First Test: Resistance Across Timing Capacitors
May 13, 2025
Board | Resistance
Board-01 (Good Board) | 481 Ω
Board-02 | 377 Ω
Board-03 | 76 Ω
Board-04 | 84 Ω
Board-05 | 383 Ω

## Slide 7
• I figured out I could use pin-to-pin resistance to diagnose amplifier damage.
• This gave a reliable fingerprint of a functioning BGA616 amplifier.
• Board-01 served as the reference for all others.
Cesar Cardona
Amplifier Pin Checks: Board-01 Reference
May 13, 2025
U1 Board-01 | 
Pin-to-pin Check | Resistance
1-2 (IN to GND) | 445 Ω
1-3 (IN to OUT) | 478 Ω
1-4 (IN to GND) | 445 Ω
2-3 (GND to OUT) | 502 Ω
2-4 (GND to GND) | 0 Ω
3-4 (OUT to GND) | 502 Ω
BGA616 Silicon Germanium Broadband MMIC Amplifier

## Slide 8
• Damaged amplifiers typically showed 0 Ω across multiple pin pairs.
• Board-05 had a shorted U1, matching this pattern. However, all other broken boards also had unusual readings in U1.
Cesar Cardona
Amplifier Pin Checks: Board-05
May 13, 2025
BGA616 Silicon Germanium Broadband MMIC Amplifier
U1 Board-05 | 
Pin-to-pin Check | Resistance
1-2 (IN to GND) | 0 Ω
1-3 (IN to OUT) | 383 Ω
1-4 (IN to GND) | 0 Ω
2-3 (GND to OUT) | 383 Ω
2-4 (GND to GND) | 0 Ω
3-4 (OUT to GND) | 383 Ω

## Slide 9
• All four broken boards showed abnormally low resistance across the U1 amplifier pins.
• I replaced each U1 amplifier using a precision soldering pen and a microscope.
• This process required careful alignment, solder wick cleanup, and steady reflow soldering of all 4 pins.
Cesar Cardona
Replacing Amplifiers
May 13, 2025
1 mm

## Slide 10
Here are some Photos. It took me back to assisting in Ophthalmic Surgeries.
Cesar Cardona
Hands On Repair at the Soldering Station
May 13, 2025

## Slide 11
• With U1 replaced on Boards 02, 04, and 05, timing signals were fully restored.
• Board-03 required U2 replacement instead — it also recovered its timing branch (But more on this board later).
• All four repaired boards now match the signal profile of Board-01.
Cesar Cardona
All Timing Branches Fixed
May 13, 2025
10 ns
20 mV

## Slide 12
• After restoring all timing branches, I proceeded to test the energy output on each board.
• The energy branch is typically more stable, but timing was prioritized first because it showed clearer signs of failure and was needed for CTR tests.
• Here’s the test setup I used, with the SiPM coupled to a crystal and a Sodium-22 source positioned nearby. Very scientific.
Cesar Cardona
Energy Branch Diagnostics
May 13, 2025
Sodium-22 Source

## Slide 13
• All four repaired boards showed energy signals that matched the reference profile from Board-01.
• This confirmed that both amplifier channels (U1 and U2) were operational and that the energy pathway was unaffected post-repair.
Except…
Cesar Cardona
Energy Signal: Restored Boards
May 13, 2025
Typical Energy Signal (Repaired Board)
100 ns
200 mV

## Slide 14
• Board-03’s energy branch was the only one that didn’t recover — the waveform showed no clear pulses.
• Timing branch was previously fixed by replacing U2, but energy remains unresponsive.
• My working theory is that either the SiPM is not properly soldered, or the inductor (L3) is damaged.
• I ran out of time to explore further, but I plan to test both hypotheses in the coming weeks.
Cesar Cardona
Energy Branch Failure: Board-03
May 13, 2025
Board-03 Energy Output (Unfixed)
100 ns
200 mV

## Slide 15
• After restoring most boards, I began searching for the underlying cause of the failures.
• I suspected the bias voltage supply might be unstable.
• I tested both supplies with an oscilloscope and captured strange behaviors in one unit — shown on the next slides.
Cesar Cardona
Investigating the Root Cause
May 13, 2025

## Slide 16
• When ramping from 0 V to 10 V, the supply did not transition smoothly.
• The scope showed oscillations and sharp intermediate transients before settling.
• This kind of instability can create momentary overcurrent or signal injection, damaging sensitive components like U1.
Cesar Cardona
Problem #1: Voltage Ramp Instability
May 13, 2025
0 to 10 V ramp – unstable behavior.
10 ms
2 V

## Slide 17
• Turning the power supply ON from the physical switch with 0 V setpoint did not yield 0 V output.
• The voltage unexpectedly jumped upward before settling back — an unsafe startup transient.
• This could easily inject a voltage into downstream components even when “OFF” conditions are expected.
Cesar Cardona
Problem #2: Voltage Spikes at Power-On
May 13, 2025
Power-On Jump with 0 V Setpoint.
1 ms
2 V

## Slide 18
• When toggling quickly between +1 V and -1 V, the scope shows a transient “bubble.”
• While small in amplitude, this change is extremely fast — and all damaged boards had U1 near the HV entry path.
• My working theory: fast transients near the transformer (balun) couple into the U1 amp, stressing it despite filtering.
• Further tests would be needed to confirm this.
Cesar Cardona
Problem #3: Fast Voltage Toggling Bubble
May 13, 2025
+1 V to -1 V Toggle – Transient Bubble.
500 µs
500 mV

## Slide 19
Cesar Cardona
System Level Schematic
May 13, 2025

## Slide 20: Questions?
Cesar Cardona
May 13, 2025