## Slide 1: Design review for the LHC-BGI magnets
D Bodart
27/05/2025

## Slide 2: Introduction
The Beam Gas Ionization (BGI) monitor is a non-invasive instrument for measuring transverse beam profiles. Building on the success of BGI devices currently operational in the PS and tested in the SPS, the high potential of this technology has been well demonstrated. In the 2022 review, BGI was selected as the baseline profile monitor for the High Luminosity LHC (HL-LHC).
To meet the stringent operational requirements—achieving a transverse profile measurement accuracy within 2.5% and an emittance error below 5%—the new Timepix-based BGI requires a magnetic field of 0.6T. This marks a significant improvement over the older MCP-camera-based BGI, which operated at only 0.2T, making the reuse of existing magnets unfeasible.
With the increased magnetic field, operational considerations necessitate an integrated B-field of zero to maintain a closed orbit and avoid transverse offsets. To achieve this, a triplet configuration was chosen, consisting of one main magnet and two trim magnets.
Working within a constrained budget, the project has prioritized cost optimization by recycling existing power converters, magnets, and infrastructure wherever possible. Alternatives, including permanent magnets 3 in 1 magnet and MDX doublet configuration, were thoroughly evaluated before finalizing the triplet configuration, which requires a new main magnet for each instrument.
In total, four instruments will be deployed—one for each transverse plane of the two LHC beams—ensuring precise and reliable beam profile monitoring for the HL-LHC.

## Slide 3: BGI-LHC Technical requirements
4 Beam Gas Ionization detectors in total are located in Point 5: 2 in 5R4, 2 in 5L4.
2 configuration are used: One with a horizontal field and one with a vertical field per beam.
A setup consist of 1 BGI magnet and 2 “field cancelling” trim magnets located upstream and downstream of the BGI magnet with opposite field so that they cancel the B field between the three magnets. each of trim magnets provide half of the integrated B field of the BGI magnet.
The requested B Field in the detector gap is 0,6 T and the Good Field Region is defined as a volume within the magnet center of X,Y,Z: 30 x 50 x 30 mm with Y offset of 1,5 mm.
The field quality is defined as DB/B <1 e-3 in the defined volume above.
The magnet gap is 140 mm as a result of the instrument size and related vacuum chamber and assembly clearance (see Fig. 1, 2 and 3).
The magnet design shall accommodates with the environment constraints defined by the machine layout, where basically the adjacent beam pipe and the QRL limit the transverse space (see fig. 4 to 7).
Easy access to the detector is needed requesting the BGI magnet to be movable in the longitudinal axis of minimum 300 mm.
Operating power converter parameters:
  The BGI magnet converters provide 600A, 40V, 1 unit per BGI magnet, already available.
  The Trim magnets converters provide 125A and 125V, 1 unit per 2 trim magnets (presently in operation and available).
Available cooling water: P=11 bars, DP 8 bars, inlet temp= 27 °C. (1/2” piping)

## Slide 4: Designs overview OPTION 4: LHC-BGI RECT INSTR & RECT CHAMBER ASSBLY - TPX3
ST1904954
Detector: TPX3
Envelope ~ 180 mm x 352 mm x 355 mm RF requirements:
  SPS-BGI type
Beam vertical offset +1,5 mm
BGI-LHC Technical requirements
Fig.1 instrument chamber size side view
Fig.2 & 3 instrument chamber size cross section

## Slide 5: Magnet integration and layout
5L4 configuration top view:
5R4 configuration top view
Fig.4 Operating configuration
Fig.6 Operating configuration
Fig.5 Instrument access configuration
Fig.7 Instrument access configuration
The configuration of the magnets is the following: The BGI magnet (in green) is set between 2 IMHH magnets (in grey) at identical distance for symetry purposes. To access the instrument the BGI magnet is slided longitudinaly thanks to available movable supports.

## Slide 6: Magnet integration and layout
5L4 configuration top view:
5R4 configuration top view
Fig.8 Horizontal field configuration
Fig. 10 Horizontal field configuration
Fig.9 Vertical field configuration
Fig.11 Vertical field configuration
The BGI magnet yokes are shaped so that they do not interfer with the adjacent equipment. It allows installation / removal with vertical lifting devices such as crane.

## Slide 7: BGI magnet design proposal
Fig.12 2D drawing with instrument integration
Fig.13 2D FEMM magnetic simulation
Fig.14 B field in transverse axis
The BGI magnet used to generate the B field to the beam instrumentation is a H window solid yoke with rectangular pole shape made of ARMCO ©. The magnet design fits to the 4 configurations. 2 coils are used to generate the field and match the power converters specifications. The optimization of the pole profile is tapered on the last 30 mm. 70562 amperes turns provides 0,6T in the magnet gap at 485 amperes. The magnetic efficiency is 94,7%. The return yokes are asymmetrical due to space constraints mainly due to the detector instrumentation supply and measurement cables.

## Slide 8: BGI magnet design proposal
The pole rectangular. The yoke length is larger than the pole tip by 60 mm at each side to reduce iron saturation. Negligible areas on the yoke has a magnetic field of 1,6T (see Fig.17). The field homogeneity is dependent of the vertical position By (see Fig.16). The optimization of the pole tip allows the By field to be 1.1 e-4 and 6,6 e-5 within the 1 e-3 in the defined volume of x=30, y=56 (including the offset), z=30 mm.
Chart: 
Fig.15 3D model of the BGI magnet
Fig.16 3D model of the BGI magnet
Fig.17 3D Field intensity in the yoke (transverse view)

## Slide 9: BGI magnet performances
The ʃ B.dl = 0,211 Tm at 0,6T for the full magnet length.
Absolute and normalized harmonics values are given for half magnet length R=17 mm L= 1 m
Coils data | 
# of Turns/coil | 72
Conductor size | 10x10x4 mm
I density | 5,66 A/mm2
Cooling circuits | 2/coil
Water velocity | 2,5 m/s
Water flow/coil | 7,5 l/min
Total dissipated power | 9,19 kW
R magnet | 0,039 Ω
Temperature increase | 18 degC
Table 1: Harmonics
Table 2: Coils data
 | Absolute values | Normalised values
B1 [mTm] | 105,21 | 
B2 [units] | 6.34E-05 | 3.94E-07
B3 [units] | 7.71E-04 | 4.79E-06
B4 [units] | 2.26E-05 | 1.41E-07
B5 [units] | 1.73E-05 | 1.08E-07
B6 [units] | 8.64E-06 | 5.37E-08
B7 [units] | 8.72E-06 | 5.42E-08
B8 [units] | 3.80E-06 | 2.36E-08
Fringe field disturbing the adjacent beam max intensity is of 2,3 mT at 0,3 m from the magnet center. Magnetic shield will be wrapped around the adjacent beam to reduce disturbance to negligible levels being lower than 1.e-4 T.

## Slide 10: IMHH magnet design
The IMHH magnet is a H type bending laminated yoke magnet. The magnets presently installed and operating in the LHC will be used for the Hi Lumi project. Their integration to the new configuration requires supporting structures modifications. The magnet aperture of 200 mm is compatible with the beam pipe of 85 mm in diameter. The magnet yoke dimensions are 750 x 656 x 646 mm. B field cancelling by upstream and downstream B field= 0,00434 Tm/A. 0,211 x 0,00434/2= 45,7 A. Detailed information of the magnet are available in the Norma magnet database here: https://norma-db.web.cern.ch/magdesign/idcard/520/
Fig.18: IMHH magnet 3D model isoview
Fig.19: IMHH magnet pictures

## Slide 11: Single IMHH magnet performances
Coils data | 
# of Turns/coil | 180
Conductor size | 6,5x6,5x3,5 mm
I density | 1,4 A/mm2
Cooling circuits | 4/coil
Water velocity | 2,35 m/s
Water flow/coil | 710 l/min
Total dissipated power | 0,38 kW
R magnet | 0,1936 Ω
Temperature increase | 0,5 degC
Table 4: Harmonics
Table 5: Coils data
The magnet configuration made of 2 IMHH and one BGI generates a B field = 0 Tm when the operating current is set to self cancelling being independent from the particle energy. The BGI magnets produces a beam deflection angle of 0,010 and 0,148 mrad at a Ek of 450 and 7000 GeV respectively.
IMHH magnet fringe field disturbing the adjacent beam max intensity is of 0,8 mT at 0,35 m from the magnet center. Magnetic shield made of electrical steel sheets or Mu metal sheets of 0,5mm double layer will be wrapped around the adjacent beam to cancel the disturbance to negligible levels being lower than 1.e-4 T.
 | Absolute values | Normalized values
B1 [mTm] | 53,033 | 
B2 [units] | -2.26E-03 | -4.26E-05
B3 [units] | -0.03394 | -6.40E-04
B4 [units] | -3.51E-04 | -6.62E-06
B5 [units] | -1.07E-03 | -2.01E-05
B6 [units] | 9.13E-05 | 1.72E-06
B7 [units] | 1.31E-04 | 2.48E-06
B8 [units] | 2.75E-05 | 5.19E-07
The ʃ B.dl = 0,1055 Tm at 0,188T. The use of two IMHH magnets provide a By field equivalent to the BGI By field. Absolute and normalized harmonic values are given for half magnet length R=17 mm L= 1 m

## Slide 12: Operating voltages
The following tables shows the total requested voltage per circuit. It includes the magnets and the supply cables.
Table 2: Coils data
Total voltage drop of the BGI magnet + supply cables |  |  |  |  |  |  |  | 
Magnet ID | BGIV.A5L4 | BGICV.A5L4 | BGIH.A5L4 | BGICH.A5L4 | BGIV.A5R4 | BGICV.A5R4 | BGIH.A5R4 | BGICH.A5R4
Length PC to M [m] | 64 | 73 | 62 | 70 | 64 | 73 | 59 | 67
Total cable length [m] | 128 | 146 | 124 | 140 | 128 | 146 | 118 | 134
Cable cross section [mm2] | 480 | 480 | 480 | 480 | 480 | 480 | 480 | 480
Rho copper | 1,72E-08 |  |  |  |  |  |  | 
Cable R [ohm] | 0,009 | 0,010 | 0,009 | 0,010 | 0,009 | 0,010 | 0,008 | 0,010
Operating voltage magnet + cables [V] | 23,2 | 23,9 | 23,1 | 23,7 | 23,2 | 23,9 | 22,9 | 23,5
The maximum operating voltage for the BGI 23.9V below the 40V limit.
The 2 IMHH magnets connected in parallel 84,8V below the 125V.
Table 6: BGI magnet operating voltage
Table 7: MDX magnets operating voltage
Total voltage drop of the power cables + 2 IMHH magnet |  |  |  |  |  |  |  | 
Magnet ID | BGIV.A5L4 | BGICV.A5L4 | BGIH.A5L4 | BGICH.A5L4 | BGIV.A5R4 | BGICV.A5R4 | BGIH.A5R4 | BGICH.A5R4
Length [m] | 64 | 73 | 62 | 70 | 64 | 73 | 59 | 67
Total calbe length [m] | 128 | 146 | 124 | 140 | 128 | 146 | 118 | 134
Cable cross section [mm2] | 95 | 95 | 95 | 95 | 95 | 95 | 95 | 95
Rho copper | 1,72E-08 |  |  |  |  |  |  | 
Cable R [ohm] | 0,046 | 0,053 | 0,045 | 0,051 | 0,046 | 0,053 | 0,043 | 0,049
Operating voltage magnet + cables at I max [V] | 84,3 | 85,1 | 84,1 | 84,8 | 84,3 | 85,1 | 83,8 | 84,6

## Slide 13: Cost / Planning
The cost estimation for the procurement of the magnets is detailed as follow:
One pre-serie magnet for magnetic measurements / optimization: 75 kCHF, lead time 9 months.
Series production of three magnets: 150 kCHF Lead time 1 year if the magnets are similar to the pre-serie unit.
One spare coil set: 20 kCHF
IMHH magnets: Available and already located at the spot.
Magnet support structures: BGI magnets table modification: 30 kCHF
Total: +/- 300 kCHF
Procurement lead time is 2.5 years from the design validation to commissioning.

## Slide 14: Conclusions
The magnetic model of the BGI-LHC magnets provides the requested B field of 0.6 Tesla with the required field quality of 10e-3 in the defined Good Field Region.
The MDX magnet used as trim magnets operate as field canceling magnets and ensure the beam is set back to its orbit. As those magnets are available they reduce procurement costs.
The magnets are compatible with the available space, the power and cooling constraints, using available power converters for all magnets and power cables for trims only.