## Slide 1: My activities - the ET pilot sector
Purnalingam Revathi
Working with Carlo Scarcia, Jan Hansen, Ivo Wevers, Sophie Meunier, Manjunath Dakshinamurthy, Paolo Chiggiato
05/06/2025

## Slide 2: Introduction: ET pilot sector components
The ET pilot sector will be assembled and tested in two phases.
Phase 2 / ET Pilot Sector
Phase 1 / Background
Pumping module
End cap with Instrumentation
Bellows

## Slide 3: Introduction: Impact of valves and instrumentation
The ET pilot sector (also in ET) will be equipped with instrumentation (RGA, BA gauges) and isolating valves.
The isolating valves are essential to replace any defective instrumentation and pumps without the
need for additional bakeouts.
In the ET pilot sector, we have FKM (viton) sealed gate valves (x 2) and all-metal right-angle valves for isolation (x 2)
along with an RGA and a BA gauge.
It is necessary to know the impact of these components on the pilot sector outgassing rate budget.
Gate valves
Angle valve
Pumping module
RGA
Angle valve
BA gauge

## Slide 4: Outgassing rate and calibration Inventory [08/04/2024 – 23/04/2025]
Element | Type | Degassing treatment | Measurement | 
 |  |  | Total (Open/Closed) | Background
DN40 AV | All metal | 250°C for a week |  | 
DN63 AV |  |  |  | 
DN100 GV | FKM seal | Valve flange at 200°C and valve body at 140 °C for a week |  | 
DN150 GV |  |  |  | 
Objectives:
Conditioning and outgassing rate measurements of all the valves and non-ferritic StS components.
Outgassing rate measurements and calibration of the instrumentation (RGA, BA gauges).
AV: Angle Valve; GV: Gate Valve; RGA: Residual Gas Analyzer; BA gauge: Bayard-Alpert gauge

## Slide 5: Outgassing rate – Angle valves
DN40 angle valve
DN63 angle valve
The results shown are for a single valve.

## Slide 6: Outgassing rate – Gate valves
DN100 gate valve
DN150 gate valve*
*CO2 outgassing rate in open position measured using the throughput method.
The results shown are for a single valve.

## Slide 7: Instrumentation – Outgassing Rates
RGA: QMA125
B-A gauge: SVT

## Slide 8: Outgassing Rates - Equivalence
Outgassing rate of RGA, BA gauge (SVT), and DN40 angle valve equivalent to the total length of the pilot sector tube made of AISI 441.
The outgassing rate of AISI 441 is considered here: 1E-15 mbar.l.s-1.cm-2

## Slide 9: Partial Conclusions (1)
Credits: Paolo Chiggiato, Peer Review of the ET pilot sector, Jan 2024
Credits: Carlo Scarcia, Peer Review of the ET pilot sector, Jan 2024
The measurement campaign for the outgassing rate of the valves and instrumentation has been concluded.
The total outgassing rate from the RGA, SVT and the isolation valves (in open position) of the pilot sector is ~ 1E-8 mbar.l.s-1 which is equivalent to 360 m of the beampipe (10 x the pilot sector).
The ultimate pressure (with the components) will be determined not by the tube but by instrumentation and valves, so the planned accumulation measurements (without the components) are necessary.
Instrumentation and valves
36 m beampipe

## Slide 10: Instrumentation Calibration: B-A gauges
* Results shown for filament 1 of B-A gauges | Electron energy: 100 eV.
Calibration factor or Sensitivity [1/mbar]
Gas species
Calibration factor of B-A gauges
Performed with Thomas Tillier

## Slide 11
* Results shown for filament 1 of the B-A gauges | Electron energy: 100 eV | Cross section data: NIST Standard Reference Database 107
Relative Ionization Cross Section (w.r.t N2)
Relative Sensitivity (w.r.t N2)
H2
N2
Ar
CH4
CO2
CO
He
Instrumentation Calibration: B-A gauges

## Slide 12: Instrumentation Calibration: RGA
RGA | Calibration Factor [A/mbar] | 
Species | Faraday | SEM
H2 | 1.03E-04 | 19.65
He | 3.11E-05 | 3.46
N2 | 7.34E-05 | 6.48
CH4 | 6.97E-05 | 6.93
CO | 7.52E-05 | 6.48
CO2 | 5.51E-05 | 3.95
* Results shown for filament A of the RGAs | Type: QMA 125 | Electron energy: 100 eV.
Pressure [mbar] (He eq.)
Ion Current [A]
He Sensitivity – Faraday mode

## Slide 13: Partial Conclusions (2)
The measurement campaign for the outgassing rate of the valves and instrumentation has been concluded.
The total outgassing rate from the RGA, SVT and the isolation valves (in open position) of the pilot sector is ~ 1E-8 mbar.l.s-1 which is equivalent to 360 m of the beampipe (10 x the pilot sector).
The ultimate pressure will be determined not by the tube but by instrumentation and valves, so the planned accumulation measurements are necessary.
The calibration of the instrumentation has been completed and it is important to have a sensitive and at the same time, stable instrumentation for precise measurements in the pilot sector.

## Slide 14
Mock-up Installation procedure document [15/10/2024 – 06/01/2025]
The mock-up is the Φ 850 mm chamber to test the sleeve welding process and to assess the dust generated during the welding.
The document includes procedure for handling during the drying phase after cleaning, packing, unpacking, assembling, and welding the components.
mock-up
https://edms.cern.ch/document/3205259/1
Performed with Manjunath Dakshinamurthy

## Slide 15
Leak detection tooling test [12/02/2025 – 24/02/2025]
The tooling is required to leak test components of the pilot sector.
It was tested on one of the ET pilot sector bellows.
Top: The counter flange and the O-ring on top of the bellow.
Bottom Left: Housing with the flat rubber seal.
Bottom Right: The housing and supporting ring are bolted together.
Performed with Jan Hansen, Carlo Scarcia

## Slide 16
Leak detection tooling test
Pump down and leak detection
The tooling can be used for leak detection of 1E-10 mbar l s-1 for the components of the pilot sector.
Performed with Jan Hansen, Carlo Scarcia

## Slide 17
ET Pilot sector Phase1 support installation [27/02/2025 - 29/04/2025]
The phase1 support installation test was performed with the help of handling team.
End frame preparation
Installation test
Performed with Jan Hansen, Carlo Scarcia

## Slide 18
ET Pilot sector Phase1 support installation
Installation procedure document
Current Status
It has been reassembled to facilitate the bottom line first and the topline next sequentially.

## Slide 19
ET Pilot sector Phase2 support installation [28/03/2025 - 01/04/2025]
The phase2 support installation test was performed with the help of handling team.
Ground plates preparation
Support installation
Performed with Jan Hansen, Carlo Scarcia

## Slide 20
ET Pilot sector Phase2 support installation
Cradle supports installation and testing

## Slide 21
ET Phase1 extremity leak detection [21/05/2025 - 02/06/2025]
The leak detection of the two extremities of phase1 has been completed and is tested leak-tight ( 1 x 10-10 mbar l s-1).
The report writing is underway.
Leak detection of the end cap
Leak detection of the bellows
Leak detection of the pumping module
Performed with Jan Hansen, Carlo Scarcia

## Slide 22
Question: If ferritic alloys are not chosen as the material of choice for the ET beampipe for reasons such as corrosion,
can we follow the LIGO solution of doing an air bake of the AISI 304L coils to reduce the H2 Outgassing?
Potential Issue: Does the tightly packed coil structure prevent the depletion to the required level (1E-15 mbar.l.s-1.cm-2) because of the increase in H2 partial pressure (1E-4 mbar in air) in the gap between stainless steel layers in the coil is very small (~ 1 -10 μm)?
In LIGO :
Material: AISI 304L 3 mm thick as coils
Degassing treatment: Air bake at 454 °C for 36h +
21 days of bakeout at 168 °C
Measured H2 specific outgassing rate: 5.2E-14 mbar.l.s-1.cm-2
Small gap between coil layers
Plan B for beampipe materials [28/03/2025 - present ] ​

## Slide 23: Analysis
QH2
Qpump
g
Coil
The most important parameter that affects hydrogen depletion is the concentration (equilibrium pressure PH2) of hydrogen around the layers.
Questions to be answered:
What is the contact force between the layers so there is no increase in H2 equilibrium pressure between the layers?
How does the outgassing rate of hydrogen at room temperature change as a function of the contact force between the layers?
How is the gap affected because of?
Clamping force,
Weight of other layers above,
Expansion during the air bake,
Cleanliness between the layers.
PH2= ?
PH2= 1E-4 mbar

## Slide 24: Basic considerations
Roughness Ra [3]:
Finish 1D/HRAP: 3.5 – 7.5 µm
Finish 2B: 0.1 - 0.5 µm
Initial hydrogen concentration in the bulk c0 = 1 wt. ppm = 0.097 mbar. l/cm3 at 20 °C
Initial hydrogen concentration in the surroundings is 1E-4 mbar equivalent to cw = 1E-4 wt. ppm = 9.9E-6 mbar. l/cm3 at 20 °C
Air bake temperature: 450 °C
Hydrogen equilibrium pressure at 450 °C: 40 mbar
The design of the stacked flat sheets experiment requires two parameters
Contact pressure between the layers: determined by the pulling force.
Dimensions of the sheet: determined by the conductance.
https://www.worldstainless.org/files/issf/non-image-files/pdf/euro_inox/roughnessmeasurement_en.pdf
[3]

## Slide 25: 2D Steady state analysis
Allowing flow in both directions:
1. Boundary conditions: P0 = 10-4 mbar at all four edges.
2. Initial conditions: qH2(T = 450 °C) = 7.5E-7 mbar.l/s.cm2 and constant
l [cm]
w [cm]
h [cm]

## Slide 26
Conclusions and Outlook
I will continue my work on pilot sector installation and measurements.
For the Plan B for beampipe materials, time dependent analysis and contact force modelling are the next steps.

## Slide 27
Thank you for your attention!

## Slide 28: Appendix

## Slide 29: DN40 angle valve assembly
Throughput system
DN40 Angle Valves (open/closed)
Blank Flange
VLV
Crosses
Throughput system
Background measurement
Blank Flange
VLV
Crosses
Assembly: 4 angle valves + 2 DN40 crosses + 5 DN40 blank flanges + VLV
Background: 2 DN40 crosses + 5 DN40 blank flanges + VLV

## Slide 30: DN63 angle valve assembly
Conditioning: Degassing valve flange at 200°C and valve body at 140 °C for a week.
DN63 Angle Valves (open/closed)
Throughput system
VLV
Cross
Throughput system
VLV
Cross
Assembly: 6 DN63 angle valves + 1 DN63 cross + 3 DN63 tees + 6 DN63 blank flanges + DN63/DN35 transition + VLV
Background: 1 DN63 cross + 3 DN63 tees + 6 DN63 blank flanges + DN63/DN35 transition + VLV

## Slide 31: DN100 gate valve assembly
Throughput system
Blank Flange
VLV
Transitions
DN100 Gate Valves (open/closed)
Conditioning: Degassing valve flange at 200°C and valve body at 140 °C for a week.
Assembly: 2 gate valves + 1 DN100/DN100 transition + DN100/DN35 transition + 1 blank flange + VLV
Throughput system
Blank Flange
VLV
Transitions
Background: 1 DN100/DN100 transition + DN100/DN35 transition + 1 blank flange + VLV