## Slide 1: Crab Cavity BPMs Status and outlook
M. Krupa for WP13
15th HL-LHC Collaboration Meeting, CERN, 02/10/2025

## Slide 2: Introduction
Multi-purpose pick-up required for CC diagnostics:
  A: cavity phasing (WP4)
  B: cavity pick-up antena signal filtering (WP4)
  C: cavity feedback (WP4)
  D: direct monitoring of residual crabbing and instabilities (WP13 with RHUL)
Technology review (for D) in January 2025 (Indico, Review report, DMR):
  Integrated button and stripline BPMs, full dual-plane implemenation
A+B
C
D
Beam direction

## Slide 3: Specifications
WP4 specification defined in LHC-BPM-ES-0015
Criterion | Key | Target | Units
Single bunch, single pass resolution at bunch centre for pilot bunch intensity | 100 | 50 | um
Single bunch, single pass resolution at bunch centre for nominal bunch intensity | 10 | 5 | um
Precision1 of the measurement for nominal bunch intensity | 10 | 5 | um
Long term stability2 of the offset for nominal bunch intensity | 50 | 20 | um
High frequency cut-off (-3dB) | 5 | 10 | GHz
Low frequency cut-off (-3dB) | ≤ 1 | ≤ 0.5 | MHz
In-band (between -1dB low and high cut-off roll-off) response variation | ≤ 1 | ≤ 1 | dB
Time resolution for single bunch, single pass measurement | 50 | 25 | ps
Acquisition length for a single bunch measurement on successive turns | > 1000 | 10000 | turn
WP13 specification defined in LHC-BPW-ES-0001 – extrapolation from the existing head-tail monitor performance

## Slide 4: Other requirements & constraints
Installation next to CCs: high beta functions and favourable phase advance
Very tight angular alignment against the CCs required by WP4
  Manual UAP – independent alignment of two beam lines (unique in the LHC)
  Tests needed to validate the spec
Au + Cu coating for impedance reduction
  Known process but tests required
NEG coating for e-cloud mitigation
  No experience with NEG-coated BPMs
  Discussions with TE-VSC started
Permanent bake-out jackets:
  Heat- and rad-tolerant SiO2 cables
  “Standard” RF termination – short access needed before and after each bake-out
CC
BPM
Q4 / arc

## Slide 5: Design
Preliminary design done by WP13
  Advanced prototyping ongoing
  Design freeze (and ECR) requested by December
  Short stripline length to be confirmed by WP4
  Hand-over to EN-MME for manufacturing drawings in 2026
Electrodes based on cryo BPMs
  Thermomechanical simulations and tests needed – electrodes and feedthroughs during bake-out
All required tooling identified, design activities started
  Biggest challenges: RF assembly, electrical qualification, handling
Loann on Wed AM (WP13)

## Slide 6: Impedance
Electromagnetic design verified and optimized with the Impedance Working Group
Impedance reduced by a factor of 9
Effective longitudinal broad-band impedance:
0.06% of the total LHC budget

## Slide 7: Expected performance (1)
WP4 functionalities: Noise floor will be dominated by the processing electronics
  Intrinsic pick-up noise: < 0.1 μm in (400 MHz ± 68 kHz) BW
WP13 functionality:
  Improved impedance matching of stripline electrodes
  High-fidelity copy of the bunch longitudinal profile, improved BPM bandwidth
Daniel on Thu AM (WP2/WP4/WP13)

## Slide 8: Expected performance (2)
Development of a new procedure for correction of imperfections of the front-end RF electronics
All existing head-tail monitors, including front-ends, measured and modelled
Technique validated offline using beam data, further computational optimization possible

## Slide 9: Expected performance (3)
Signal digitizer remains the bottleneck of the system
  Commercial high-end oscilloscopes used currently
New digitizer will be procured as late as possible to leverage (slow) technology progress
  Current technologies do not fully meet the WP13 specification
Electro-optical time-stretching read-out techniques under consideration to replace the traditional RF front-end
Andreas on Wed AM (WP13)
~ 50 dB DR
2 GHz

## Slide 10: Budget and schedule
Budget of 852 kCHF defined based on in-house production of cryo BPMs and with inputs from technical groups (TE-VSC, BE-GM)
  Partially financed through HL-UK2 contribution
  Plan to commit most budget by mid 2027
Aggressive production schedule to maintain the momentum from cryo BPMs production and to have reasonable contingency
  2025: design review and freeze, order of raw materials, prototyping
  2026: pre-series production of BPMs, coating studies and tests, production of alignment platforms, assembly tooling production
  2027: series production of BPMs, installation and handling tooling production
  2029: tunnel installation (after CCs)

## Slide 11: Summary
Operation of crab cavities requires new and challenging diagnostics to be deployed in the LHC
  BPM design being finalized by WP13
  Novel RF front-end solutions under study by WP13
  Electronics design ongoing by WP4 and WP13
BPM based on “standard” technology but with significant technical challenges
  Precise alignment
  Electrical assembly
  Coating
  Handling
Budget and schedule defined and successfully followed so far

## Slide 12: Thank you very much for your attention
Special thanks to colleagues from RHUL, ATS-DO, BE-ABP, BE-GM, EN-MME, SY-RF, TE-VSC for providing their support and expertise

## Slide 13: EOMTS-Prototype in LHC Point 4
DC bias voltage
EO modulator
ADC (RFSoC)
Stripline
Gallery (UA47/U0-001)
Surface (SX4 R-002)
Tunnel (2438/U0-2438)
+12.5 ns
RF signal
Attenuators (> 30 dB)
90:10 splitter
DCM 20 km
Femtosecond
laser
Chirped volume
bragg grating
DCM 5 km
PM fibre
Balanced Photodetector
SM fibre
stretched time
500 m
1 ns
4 ns