8x8 SiPM array + WLS + Atlas tiles May 2025 beam test at DESY
Jiri Kvasnicka1), Ivo Polak1), Gerald Eigen2), Andre Wilhahn2), Yannik Buch2), Jaroslav Moravec3), Matej Rehak3)
1) FZU, Prague
2) Uni Goettingen
3) Foton, Nova Paka
kvas@fzu.cz

---
Testbeam setup
- 32 Atlas TileCal tiles (same size, 224x97 mm)
- ATLAS WLS fibre samples (provided by Stanislav Nemecek)
  - Cut to 55 cm length
  - Polished ends via 3M lapping papers (down to 0.3 um)
  - loose fibre ends coated with silver via magnetic RF sputtering (in-house)
- Each tile has a plastic side “cap” holding the WLS fibre in the center of the tile edge.
- All tiles positioned by a plastic plates with cutouts for fibres and tiles (“Tile stand”), held together by four M3 screw rods
- SiPM Matrix HPK S13615-1025:
  - 8x8 elements, 1.2mm pitch (0.2mm gap between SiPM elements)
  - 25 um pixel pitch (1584 pixels), 47% fill factor, 25% P.D.E. (Low PDE !)
- CAEN DAQ DT5202, connected via “dupont” cable to SiPM, ethernet readout
- Cooling: DAQ via black box side wall and adhesive thermal pad, SiPM air cooled
- Temperature logging: Onboard LM35 sensor, later temperature logger
beam

---
Testbeam programme
- Main goals: measure the LY (light yield) and crosstalk of the whole chain, acquire an experience with CAEN DAQ
- 3 testbeam positions in the black box:
  - Vertical perpendicular (hitting all tiles),
  - Vertical longitudional (hitting 1 tile in length of 97 mm) - difficult alignment,
  - Horizontal longitudional (hitting 1 tile in length of 97 mm)
- Two trigger setups: Internal (autotrigger), later external trigger (coincidence of 3x1.5 cm2 finger beam scintillators)
- Energy scans: 2.6 (highest rate), 1, 2, 3, 4, 5, 6 GeV e-
- Uniformity scan 15+2 positions
- Bias voltage scan (59V down to 51.8V in 0.4V steps)
- Positional scans in longitudional orientation
  - 7 positions, beam is wide enough to cover several tiles
- Highly collimated beam (low rate)
- Application of optical grease between SiPM and Fibre bundle
- Second SiPM board
- High statistics overnight runs (No automated positioning, unfortunately)
- Short run with a tungsten absorber in front of the black box

---
Crosstalk / pedestal shift issues
- When looking at the data during the beam test, we noticed a severe signal degradation - SPS not clearly visible
- We made a test of 32 runs, where we enabled only “n” scintillators (and 2*n channels) via individual channel voltage trim
  - Voltage trim of up to 4.5V done at the SiPM input of the ASIC via “Input DAC”
  - Quickly confirmed, that we are experiencing a massive signal shift to the left in the histogram.
  - Signal of 2nd p.e. shifted left and merged with pedestal and 1st p.e.
  - Selected 8 channels operation as the maximum operation where we see SPS on all channels
- Every measurement has to be made as a batch of 8 measurements with different set of enabled channels (balanced between 2 ASICs: 4ch + 4ch)
  → event building for a single electron is not possible in all 32 tiles
- Possible source of signal degradation
  - NOT the Optical crosstalk (measured separately in the lab, <1% on geometric neighbors) – we would
    still see the SPS nicely
  - Internal in the ASIC- can be probably excluded, as this has not been seen with 20 individual SiPMs (see later)
  - Electrical crosstalk on long wires (might contribute partly)
  - Primary suspect: High voltage distribution + caen DT5202 connection
    - Our board has not ideal SiPM connection. It has a single bias input + filter and array of decoupling capacitors. Anode signal was supposed to be measured as a current to the ground (it has 64 dedicated ground pins)
    - CAEN DT5202 has 64 individual cathode pins and 64 anode pins with no ground
    - Even when the adaptation if the SiPM board to CAEN was not ideal, we did not expect that huge impact

---
Towards calibration and analysis
- Purpose: we want to measure number of photons in each channel
- Several parameters need to be put in the equation:
  - Gain
  - IC factor between HG (High Gain) signal and LG (Low Gain) signal (next slide)
  - Pedestal value (default 100, but may shift according signals in other channels)
  - TODO: saturation correction
  - TODO: crosstalk from other channels
- Gain extraction: a robust, but less accurate method was implemented
  - Signal treated as a logscale waveform in time domain
  - Butterworth bi-directional bandpass filter applied
    around the expected gain frequency (+-35%)
    - filtfilt from scipy.signal
  - Average of 7 peaks above a threshold is calculated
  - Very robust, works almost on every run
  - Systematic decrease of gain when more channels
    are connected

---
Intercalibration factor HG2LG
- Both HG and LG signals are recorded by the ASIC. Example shown: Lab test with LED.
- HG range is about ~90 p.e., LG range ~10x more
- Non-trivial conversion, linearity is not great (or not well understood)
- Fit does not go through [0,0] (actually [100,100]). Typical factor ~10
- When fitted the higher values, the error around pedestal level is ~2 p.e.
- Method used: fit a HG range of 2000-4000; use HG values for HG<2000; converted value for HG>4000, smooth blend in between

---
Results preview
- All channels can be calibrated
- Chip 2 has higher noise
  - Especially pronounced with the longer shaping time
  - Same noise increase observed with other DT5202 in the institute
- Longitudional Light Yield: 75~135 p.e.
  - With optical grease applied
  - Translates to 2.3~4.2 p.e. for perpendicular e- impact.
  - Consistent to the results measured perpendicularly (2~4 p.e. by eye)
Picked active channels
from 7*8 runs
Calibrated p.e.

---
Pedestal shift investigation
- Comparison: we do have a data from the lab at FZU with 20 individual SiPMs (HPK 6x6 mm2)
  - Not confirmed with this setup, shift <<0.05%
- SiPM array with the LED light: shift ~2 %
  - Similar order of magnitude observed during testbeam
- Might not be as linear as we think
SiPM Array,
LED
Individual SiPMs,
Sr-90 source
100 p.e.
1000 p.e.
Electrical crosstalk overtakes ?
(neighbor cable)
Individual SiPMs,
Cosmics
2000 p.e.

---
Summary
- No “disaster” issues with the CAEN DT5202 operation, only issues
  - Quite some heat being burned – we had to cool and stay away from SiPM
  - No statistics, no temperature during a “batch mode” – SW bug
  - Missing timing information from the data → no information which channel passed the threshold in the event. FW bug?
  - External trigger latency @testbeam: too slow. We had to put the ASIC into 50 ns shaping time (2nd asic has increased noise in this mode)
  - Only moderate datataking rate at the testbeam (2MBytes/s, ~5k events/s)
  - 900 GB of data for 6 days of beam
- Temperature very stable during the whole testbeam (within 1.5 DegC)
- Huge electrical crosstalk: 2~3% to some/all channels? Need to understand the mechanism (SiPM connection + cable?)
  - We were able to see simultaneus SPS (single photon spectra) only in 8 channels at once. Not for all channels at the time
    => significant slow down of data taking (in group of 8 channels),
    => no complete events in all channels (always only a subset)
  - We think this can be fixed for next time with different SiPM PCB design with electrical connection suited to CAEN DT5202 pins (direct Cathode+Anode)
- Light yield of 75-135 p.e. for longitudional e- hit, 2.3~4.2 p.e. for perpendicular e- hit
  - mind the PDE of our SiPM (25%) and spread possibly due to field application of the grease
- Data analysis: Andre Wilhahn (PhD student @uni Goettingen) might start looking at the data

---
Backup and details

---
Mechanics for Tile testing
- Tile “cap”: WLS holder (in scale)
- Tile holder (OpenSCAD parametric)
  - Screwed together via M3 thread rod
  - Rubber padding inside to put pressure on the WLS, otherwise they slide too freely

---
Further reports
- Lab tests (Goettingen, Prague)
  - 2023: single fibre and 4x4 flange, single channel readout (report)
  - 2024: single fibre and 8x8 flange, 64 channel readout (report, report)
  - 2025: construction of WLS fibre holder, fibre polishing, testing 32-tile stack with cosmics and Sr90 (report)
- DESY Testbeam 05/2025 with 32 tiles (early report)

---
SiPM Matrix parameters

---
WLS fibre holder