## Slide 1
Pushing the Limits: Advanced SiC Detectors for Operation in Harsh Environments

## Slide 3: The answer at the end of the talk!

## Slide 4: Main potential applications in harsh environments
https://protomag.com/medical-history/flash-radiotherapy-finally-way/
SiC Radiation detectors
High Energy Physics
Fusion
Space
Medical dosimetry
Nuclear Physics
Plasma diagnostic
in Fusion reactors
Nuclear Physics
Radiotherapy
HEP
Space

## Slide 5
https://www.precedenceresearch.com/silicon-carbide-semiconductor-devices-market
Silicon carbide is expected to improve energy efficiency in key sectors of the economy. Europe has many large companies leading this technology.
SiC devices are widely used in fields such as rail transit, ultra-high voltage power grids, charging stations, photovoltaic power generation, new energy vehicles, high-speed trains, consumer electronics, and more.

## Slide 6: Why use SiC radiation detectors?
F Nava et al 2008 Meas. Sci. Technol. 19 102001
J.M. Rafí et al 2018 JINST 13 C01045
Same area
150mm wafers
Insensitive to visible light

## Slide 7: Few comments
Very recently Silicon-Carbide 2um CMOS technology is available in the EUROPRACTICE portfolio, provided by Fraunhofer IISB.
Epitaxial layer thickness is limited by increasing defect density (< 250 μm) -> kV for full depletion…
Standard wafer size is 150mm but now it is possible to find high quality 200mm wafers.
R&D necessary to understand radiation hardness for specific applications. Defect engineering is not clear for high neutron or ion fluences.
Low charge generation (Ionization energy) per μm and MIP SiC: 57eh; Si: 80 – 90eh -> Gain may be necessary for mips.
Simulation very well established for power devices, no very accurate for detectors *.
Most advanced microelectronic processes are available for SiC, and large-scale production may be feasible within European industries—provided there is sufficient interest.
* Review: https://arxiv.org/pdf/2410.06798

## Slide 8: High Energy Physics experiments
Large interest from the DRD3 communities.
Last meeting on 2-6 June 2025 in Amsterdam, https://indico.cern.ch/event/1507215/

## Slide 9: HEP
Fabrication run RD50 project
Second fabrication completed. Three wafers fabricated:
one with 50 µm thick epitaxial layer,
one with 100 µm thick epitaxial layer,
and one dummy wafer (high doped substrate).
Details on the design were presented by S. Weid, 2nd DRD3 workshop

## Slide 10: Mapping and yield
Yield 68.8%
Yield 78.6%
Full depletion voltages:
50um ≈ 300V
100um ≈ 2200V

## Slide 11: Charge collection vs Fluence (1)
RD50 collaboration in 2003 already demonstrated the lost of charge collection in SiC and GaN detectors exposed to high neutron fluences.
They also attempted to use triode (transistor-based) detector structures to achieve gain, but without success (M. Bruzzi on behalf of the RD50 Collaboration – IEEE NSS MIC 2003, October 19-24, 2003).
J.M. Rafí et al., 10.1109/TNS.2020.3029730
New high-quality SiC material !

## Slide 12: Charge collection vs Fluence (1)
RD50 collaboration in 2003 already demonstrated the lost of charge collection in SiC and GaN detectors exposed to high neutron fluences.
They also attempted to use triode (transistor-based) detector structures to achieve gain, but without success (M. Bruzzi on behalf of the RD50 Collaboration – IEEE NSS MIC 2003, October 19-24, 2003).
New high-quality SiC material !
Same results!
J.M. Rafí et al., 10.1109/TNS.2020.3029730

## Slide 13: Charge collection vs Fluence (2)
T. Bergauer, RD50 workshop
J.M. Rafí et al., 10.1109/TNS.2020.3029730
Note: all SiC measurements at room temperature!
Much higher bias may be needed to recover the charge! Problems with the electronics.
G. Casse, et al., Update on Thin Silicon Detectors Results. 21st RD50.
Silicon
2x1015 neq cm-2

## Slide 14: How can we turn SiC into a better detectors?
SiC LGAD -> like silicon it is possible to implement a multiplication layer to mitigate the small signal before and after irradiation.
  Impact ionization in multiplication layer produces gain  larger signals.
Challenges:
  Creation of deep gain layer very difficult to achieve by “normal” ion implantation (as usually done for Si-LGAD).
  Gain layer could be implemented during epitaxial growth
    involvement of wafer supplier necessary in formation of gain layer
    We know that wafer supplier can grow sandwich of N-/P layers.
  Alternative: high-energy implantation (tens of MeV).
Yang Tao, 39th RD50 workshop
G. Pellegrini et al.,
10.1016/j.nima.2014.06.008
Silicon

## Slide 15: However: Gain is lost after irradiation!
Gain calculated as a ratio of TCT signal amplitudes/integrals between LGAD1 (LGAD2) and no-gain PN diode irradiated to the same total fluence.
Jiri Kroll et al., Radiation effects in 4H-SiC PN and LGAD, 3rd DRD3 workshop.

## Slide 17: Signal multiplication in irradiated SiC p-n diodes
TPA-TCT z-scan of irradiated sample
I.Vila, 2nd DRD3 workshop, CERN, December 5th, 2024.
Indication of hole multiplication due to a high electric field deep into the active thickness
Φ=1015 neq /cm2 at ATI in Vienna
P++
N++

## Slide 18: Timing resolution
-TCT measurements with UV laser, Signal equiv. to 5.5MeV alpha in SiC show time resolution of 10ps.
  Comparison made between 50um thick 1mm DUT detector and Reference detector.
  G. Pellegrini et al, 3rd DRD3 workshop
-MIP measurements shows time resolution of 94ps with 5x5mm2 diodes at 500V.
-T. Yang et al., 10.3389/fphy.2022.718071

## Slide 19: Conclusion HEP
Si retain much more charge after irradiation but requires cooling even at very low fluences.
Gain is lost after irradiation; defect engineering may help to mitigate the negative effect (DRD3 WG3).
IBIC and TPA TCT show high signal when operated in forward bias -> holes multiplication (DRD3 WG5).
  Can we engineer (smarter design) this effect to harness it for practical applications?
Time resolution is better for SiC than Si (without gain) but more measurements are necessary after irradiation and gain is very likely to be necessary for 4D tracking.

## Slide 20
Nuclear Physics

## Slide 21: Nuclear physics: ISRS (ISOLDE Superconducting Recoil Separator)
DE = 30 keV Dt < 1 ns
No cooling
The SiC detectors are a promising tool not just for a focal plane, but also as beam monitor and/or tagging system (till 1x10-7 pps).
ISRS consists of a superconducting storage ring that enables mass identification of nuclei via time-of-flight measurements. The decays of exotic nuclei can be observed by detecting particles or photons using high-resolution detection systems.
System requirements:

## Slide 22: Collaboration with the IEM-CSIC, Madrid
ISRS Focal plane detectors:
Linear spectrometer mode
SiC detection technology
SiC prototypes developed by
CNM-IMB Barcelona
Silicon prototype

## Slide 23: Preliminary results
Good energy and time resolutions:
30 keV at 5.8 MeV
Expected time resolution: tens of ps.
Irradiation campaign ongoing (lower fluences than HL-LHC).

## Slide 24: Conclusion for Nuclear experiments
SiC offers lower leakage current and better timing resolution compared to Si.
Signal amplitude is not an issue, as ions produce a large signal.
Cooling is not required.
A new batch of detectors is nearly ready for testing with ion beams.
One wafer will be thinned to 100um.

## Slide 25: Radiotherapy

## Slide 26: Flash radiotherapy
Vozenin et al., Clin Cancer Res 25 (2019)
Schüller et al., Physica Medica 80 (2020)
FLASH effect (Favaudon et al., Sci Transl Med 6 (2014)):
  Irradiations with Ultra-High Dose Rate (UHDR) pulsed radiation reduce adverse effects in healthy tissues
Therapy doses typically range from 45 to 60 Gy divided into several fractions
With FLASH it is possible to treat radioresistant tumours with higher dose.
  FLASH dose rate ≥ 40 Gy/s (conventional radiotherapy ≈ 0.05 Gy/s)

## Slide 27: FLASH conditions
1 Gy in water of 160 MeV protons ≈ 1E9 p/cm2
per pulse of ~1µs
(dose = fluence × mass stopping power)
Saturation of conventional active dosimeters (gas and silicon)
Need real-time, precise, radiation hard dosimeters
Motivation: FLASH radiotherapy
w | pulse width | [0.5 – 2.2] µs
f | pulse repetition frequency | [10 – 200] Hz
D*p | dose-rate in pulse | [103 – 5 · 106] Gy/s
Dp = D*p · w | dose per pulse | [10-3 – 5] Gy
D *m = D*p · w · f | mean dose-rate | [10-2 – 1000] Gy/s
Notes: 1 Gy in water

## Slide 28: Semiconductor diode dosimeters
p–n (or Schottky) junction in:
  Current mode operation
  Unbiased conditions
The radiation-induced current is measured by an electrometer
The current is proportional to the dose rate and its integration provides the absorbed dose in the semiconductor
Shi et al., Med Phys 30 (2003)
I. Lopez, Nature Scientific Reports |, First use of silicon carbide detectors with graphene‑enhanced contacts for medical dosimetry

## Slide 29: SiC diodes for FLASH dosimetry
P-N junction implanted diodes
4H-SiC, 100 mm wafers
3 µm high resistivity n-type doped epilayer
1 and 2 mm diameter with multi guard ring structures
Technology based on our SiC power diodes
Adapted for radiation detection with high temperature processes and optimized metal contacts
Patented technology

## Slide 30: Electrical tests
Large single diodes
Forward I-V
Reverse I-V

## Slide 31: Dosimetry of FLASH electrons (2)
PTW’s microSilicon
SiC
Detector linearity
Response independent both of Dose Per Pulse (DPP) and of instantaneous dose rate
SiC diode
  1 mm diameter
  Encapsulation by PTW
  Without external bias
C. Fleta et al., Phys Med Biol 69 (2024) 095013
20 MeV FLASH electrons at PTB reference beam (Germany)
SiC diode in water phantom at PTB
Dosimetry of FLASH: electrons
Notes: Response independent both of DPP and of instantaneous dose rate

## Slide 32: Dosimetry of FLASH electrons (3)
Long-term radiation hardness
(initial irradiation of SiC diode (10s kGy) not shown)
Linear decrease of sensitivity of 0.018% / kGy of 20 MeV electrons
 Not as radiation hard as diamond, much better than commercial silicon diodes
Temperature depencence of sensitivity
Sensitivity reduction of –0.08% /°C
Better than most commercial Si diodes + opposite trend.
Notes: Advantage -> does nto require calibration

## Slide 33: Radiation hardness with FLASH protons (1)
1 and 2 MeV protons at CNA (Sevilla)
FLASH beam structure: 1-10 ms pulses, 500 Hz – 10 kHz, up to 5.6 Gy per pulse (4.6 MGy/s)
30 µm diameter planar microdiodes
Vacuum chamber
SiC dosimeter
Sample holder
To Electrometer
To Brookhaven current
integrator
Pulsed beam current measured to calculate dose
M.C. Jiménez-Ramos et al., submitted to IEEE Trans Radiat Plasma Med Sci
UNPUBLISHED DATA – DO NOT COPY OR DISTRIBUTE
Dosimetry of FLASH: Protons

## Slide 34: Radiation hardness with FLASH protons
Diode #1
Diode #2
-1.34 %/kGy
-0.04 %/kGy
Commercial silicon dosimeter:
-5 %/kGy with 62MeV protons
NIEL/LET (62 MeV/2MeV) ≈ 1.3
Radiation hardness with FLASH protons (3)
UNPUBLISHED DATA – DO NOT COPY OR DISTRIBUTE
L. Raffaele et al., NIMA 891 (2018)
NIEL/LET (62 MeV/2MeV) = 1.3
62 MeV protons cause 1.3 times more damage than 2 MeV for the same dose
Notes: NIEL/LET (62 MeV/2MeV) = 1.3
62 MeV protons cause 1.3 times more damage than 2 MeV for the same dose

## Slide 35
Linear behaviour for FLASH dose rates is maintained after irradiation
Diode #1
Diode #2
Radiation hardness with FLASH protons (4)
UNPUBLISHED DATA – DO NOT COPY OR DISTRIBUTE
UNPUBLISHED DATA – DO NOT COPY OR DISTRIBUTE

## Slide 36: DOSImetry monitor for FLASH therapy
Arrays of diodes: 0.6 x 0.6mm2.
3um epitaxy Thickness.
SiC microdetector array development and electronics -> ≈400 channels with an area of 20 x20 cm2.
For treatment verification before patient irradiation
C. Guardiola & C. Fleta (IMB-CNM-CSIC)
Like IBA MatriXX ionization chamber

## Slide 37: Different technologies
Arrays of Microdiodes
10 to 30 µm diameter
For dosimetry and microdosimetry with 2D spatial resolution
Planar and 3D (“mesa”) configurations

## Slide 38: Trenched and planar microdiodes
Breckdown down to 30-40V due to curvature

## Slide 39: Wafer mapping (1mm diodes)
(In white: rejected diodes, with I@(-10V) > 10 pA. Total 65/1140 = 5.7%)
13.2 cm
4.2 cm
Breakdown at 300-350 V

## Slide 40: Conclusion for FLASH dosimetry
Sensors can be operated without external bias
Signal linearity of SiC is one order of magnitude higher than commercial Si diodes.
Linear performance is maintained after MGy doses
Sensitivity reduction due to temperature is better than commercial Si diodes
Sensitivity degradation with dose slower than for commercial Si diodes -> less need for re-calibration

## Slide 41
Plasma diagnostic
in Fusion reactors

## Slide 42: Fusion: plasma diagnostic
X-rays: dominant source of radiation (hn=hc/l~Te):
100 eV<Te<20 keV ⇒ 0.5 Å < l <130 Å ⇒ X-rays!
20 keV< Ee< 200 keV ⇒ SXR to HXR
D-T reaction produce suprathermal alpha particles (3.5 MeV). These alphas are essential for self-heating of the plasma.
Neutron emission is directly proportional to the fusion reaction rate of deuterium-tritium (D-T) reactions, which produce 14.1 MeV neutrons.
Long pulse operations (LPO) open a new era of expertise, training (e.g. physics and engineer) & collaborations for radiation-hardened detectors.
See talk of Luis Felipe Delgado Aparicio at the DRD3 workshop June 3rd , 2025.

## Slide 43: Need to develop radiation-hardened sensors and electronics for this harsh enviroment
Main issues: circumvent gamma- and neutron-induced noise, T- and B-limitations, as well as define LIFETIMES!
Si-detectors:
Used due to the availability of good quality homogeneous material and high charge carrier transport properties. Can only withstand 1012-1013 neq/cm2.
Lifetimes shortened:
  neutron damage
  Survivability at high temperatures (200-300ºC)
Forced to look for new solutions compatible with CERN’s very-high-luminosity experiments (up to 1016-1017 neq/cm2 ???).
ITER
DT
1 MW of DT neutrons: ~4.4×1017 n/s
For d=15 m ⇒ Fn ~ 1.5×1010 n/cm2/s
Dt=5-10 s ⇒ ~ 1011 n/cm2/shot

## Slide 44: Measurement at high temperature SiC diodes
M.C. Jiménez-Ramos et al., https://doi.org/10.1016/j.radphyschem.2023.111283

## Slide 45: Irradiation at high temperature
C.Torres-Muñoz et al., 3rd DRD3 workshop, 2-6 June 2025, Amsterdam

## Slide 46: CCE versus voltage and irradiation T
C.Torres-Muñoz et al., 3rd DRD3 workshop, 2-6 June 2025, Amsterdam

## Slide 47: Conclusion Fusion
SiC detector can be operated at high temperature.
SiC detector response improves when irradiated at higher temperatures
Operating the device in high-temperature environments results in less spectrometric degradation compared to room temperature, effectively extending the device’s lifetime.
Urgent need of rad-hardened sensors for long pulse operations.
Need better neutron monitors for DD and DT (2.5 vs 14.1 MeV): flux and spectra

## Slide 48: Is SiC more radiation hard than Si detectors?

## Slide 50: Acknowledgement
Funding:
This project has received funding from the European Union’s Horizon Europe research and innovation programme under grant agreement No 101075672, FOR2ENSICS Project.
UHDPulse: the EMPIR programme cofinanced by the Participating States and from the European Union's Horizon 2020 research and innovation program (18HLT04).
This work has been funded by the Spanish Ministry of Science and Innovation (MCIN/AEI/10.13039/501100011033/) and by the European Union’s ERDF program “A way of making Europe”. Grant references: PID2021-124660OB-C22.
NEWDOSI: Desarrollos tecnológicos para nuevos detectores de estado solido para dosimetría avanzada (PID2021-123484OB-I00).
CARACTERIZACION: Caracterización de dosímetros avanzados fabricados en el IMB-CNM para aplicaciones médicas (202250I185).
Microdetectors for medical applications. Proyectos Intramurales Especiales CSIC (20225AT008).
Dosimetry monitor for FLASH therapy. LaCaixa foundation (HR23-00718)

## Slide 51: Microprobe beam line at CNA

## Slide 52: Fast neutron detection with SiC
Measurements with cuasi-monoenergetic
fast neutrons (HiSPANoS CNA)
SiC detectors
Pérez, Martín, et al. "Characterization of new silicon carbide neutron detectors with thermal and fast neutrons." Nuclear Instruments and Methods in Physics Research Section A: Accelerators, Spectrometers, Detectors and Associated Equipment 1069 (2024): 169968.
Experimental setup CNA
The total deposited energy (red curves) is the sum of the energy produced by the recoil nuclei of 28Si and 12C and also of the protons generated after the interaction of fast neutrons and the plastic holder that surrounds the Detector.
Protons deposit energy over the detection threshold of 0.7 MeV.
Measurements with AmBe source SiC+Polyethylene (Bariloche Atomic Center -Argentina)

## Slide 53: SiC vs Si test at IEM-CSIC:
Mesytec MPR-1, Cividec (for diamond module) along to DAQ-Mesytec and Maestro Multichannel
MPR-1 + ORTEC 671 Amp on Maestro
Comparison between ORTEC chain + Maestro data taking Mesytec digitizer for SiC and Si detector. Both cases use the MPR-1 readout

## Slide 55
Therapy doses typically range from 45 to 60 Gy. Typically, these doses are divided into multiple smaller doses that are given over a period of one to two months. The total radiation dose is usually divided into several fractions. For most patients who require radiation therapy, the total dose is broken up into daily doses of five times a week for a total period of five to eight weeks

## Slide 56
Cross-section of the proposed N-LGAD structure in SiC. It is designed to optimize
the multiplication of hole charges