| ISOLDE |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Machine Supervisor last week |  |  | Erwin Siesling / Emiliano Piselli |  |  |  |  |  |  |
| Machine Supervisor this week |  |  | Emiliano Piselli / Simon Mataguez |  |  |  |  |  |  |
| Beam Scheduled |  |  |  |  |  |  |  |  |  |
| GPS | Yes | HRS |  | No |  | HIE-ISO |  |  | Yes |
| Beam Availability by Destination (AFT) |  |  |  |  |  |  |  |  |  |
| GPS | 96.8 % | HRS |  | - % |  | HIE-ISO |  | 99.7 % |  |
| Facility Status |  |  |  |  |  |  |  |  |  |
| Summary | Report WK26: Monday 30/06/2025 General: - From technical as well as physics point of view a relatively good week GPS / REX/HIE: - Continuing delivering Ytterbium beam, 162Yb41+ at 4.1MeV/u, to the HIE ISOLDE Miniball experiment IS761. (Target #902 Ta). - Interruption of the run due to the technical stop starting Monday 23/06 15h. - Beam supposed to be back at the experiment Wednesday 25/06 at 12h but due to issues at PSB this was not the case until 23h. - IS761 was running Wednesday-night with an increased Yb intensity and approved (F. Pozzi, RP) derogation for the hall PAGIS501 radiation monitor up to 50uSV/h to recover the lost beamtime. - The Miniball IS761 experiment finished on Thursday-morning at 10h30 with a successful run. Technical: - Friday-afternoon 27/06 Nicolas Guillotin (TE/CRG) changed the CM4 He level from 54% to 47% as requested by Daniel Valuch (SY/RF). The 47% has proven to be more stable for the SRF (setting last year). - Daniel Valuch will be working on the max gradients for the SRF (weekend & early next week) GPS: - Target change to #900 UCx-VD7 on Thursday-afternoon 26/06 - Stable beam and yield checks on Friday - Physics to IDS (IS687, 22O) and TAS (IS693, 17N) foreseen during the weekend - Note: The target did not produce any Oxygen and it was decided to run different isotopes to profit from the allocated physics time (He, Ne, Ar). HRS: - HRS in standby and restarting Monday-afternoon 30/06 in preparation of Polonium beams for CRIS - Restarted HRS for CRIS on Monday. |  |  |  |  |  |  |  |  |
| Issues | HIE ISOLDE: - Only a few SRF trips to report – estimated ½ hr total. Technical: - For the upcoming run the max gradients will be needed- instabilities at higher gradients: Daniel Valuch will spend time to work on these instabilities and define the final max gradients REX: - The amplifiers went down on Monday 30/06 due to a swap between the mixed water chillers (maintenance: leak- and functional tests). Should normally not happen (Xavier Decrand CV) GPS: - Over the weekend the target heating once dropped, an HT trip occurred and a few electrostatic elements in CAO went down and needed a reset - GPS.BFC4900 shows sometimes a strange behaviour: starts with correct current which then slowly reduces. Not seen on the following FC’s Technical: - The tape of the tapestation was found broken on Monday 30/06 when Rn yield checks were supposed to be carried out (Isa Frank STI). Repair to be scheduled. - HRS/ISCOOL: - At ISCOOL there are issues getting the beam through: The switch between EXTA voltage and GND would go to EXTA when no timing is applied. With then the power supply set to 0V it seems the plate charges up and forms a barrier for the beam. This also happens if timing is applied but voltage set to 0V. But not always.. We have now located the issue and will find a solution (terminator on the timing input if not in bunch mode so the GND from the switch is directly connected to the plate). |  |  |  |  |  |  |  |  |
| Plans | GPS run to stop on Monday 30/06 HRS setting up to start on Monday 30/06 |  |  |  |  |  |  |  |  |
| Intervention Request |  |  |  |  |  |  |  |  |  |
| Yes / No | Duration |  |  |  | Preferred date/time |  |  |  |  |
| Reason |  |  |  |  |  |  |  |  |  |
| Impact |  |  |  |  |  |  |  |  |  |