## Slide 1: Measurement and alignment needs for the BDF Target
A. Huncikova, C. Vendeuvre
23.07.2025

## Slide 2: Measurement methods study – BDF Target
Laser Tracker Leica ATS800 with full 1,5 inch spheres
Laser Tracker Leica ATS800/ATXXX with Newport retroreflectors

## Slide 3: Measurement methods study – BDF Target
Question:
Robotic solutions for the installing of the removable retroreflective targets/laser tracker – accessibility?

## Slide 4: Spatial Requirements - Survey network
Link between TCC8 and ECN3 preferably on both sides
Jura
Opening through the chicane (installation of brackets on the wall up to 15 m up- and downstream)
Salève
Link on the floor level/
In height? Is there space along the wall?
Position and number of the fire doors and confinement walls and the time of installation

## Slide 5: Spatial Requirements - Vacuum Vessel
Spots for the positioning of the instrument with a permanent pillar – possible collision with the conveyor and shielding manipulation
BDF Target – openings in the shielding on both sides and enough margin
Clear line of sight

## Slide 6: Environmental conditions inside the vacuum vessel
Radiation dose – gross estimate 10 MGy/yr for materials with density of 7,85 g/cm3
Temperature – up to 100°C for the Au and Al coated retroreflective prisms by manufacturer
Vacuum – conditions inside the vessel for the tests
Radiation tests for retroreflective targets

## Slide 7: Error Budget
Measurement |  | mm |  | 
Network | Network simulations (geometry+instrument) | 0.19 | R 0.250 L 0.050 Z 0.050 mm - simulation part. masked ST20 REF10 M40 ECN3+TCC8 - WORST CASE | 
 | Stability of the network overtime (mainly building sitting) | ? | Re-measuring the network, measuring the network before and after the installation | 
Instrument | ATS800 | 0.03 | 2cc 5 μm + 2 ppm for ATS800 - distance cca 5 m from the target ca. 0.015 d, 0.024 for angl., | 
Survey target | Single station - single point | 0.06 | OPTIMISTIC | 
 | Radiation/heat damage | ? | Full spherical targets, radiation tests on retroreflective prisms | 
Target |  |  |  | 
Fiducialisation of the target | Configuration of the fiducials within the target | ? | To be simulated according to the possible geometry within the vacuum vessel | 
 | Fiducialisation | 0.05 | standard empirical value | 
 | Deformation of the target during operation | ? | What is the expected range? Better observation from two stations | 
Precision of the alignment system | Capability of the alignment system | ? | Ability to align with respect to the nominal | 
 | Precision of the repeatability of the position | ? | Compensated by measurement of the target inside the vacuum vessel | 
External |  |  |  | 
 | Refraction caused by heat instability after opening the vacuum vessel | ? | What is the air temperature after opening and during the measurement? | 
 | Mechanical vibrations | ? | Cooling, ventilation, machinery...? | 
 | Installation and desinstalation of the shielding | ? | We do not know | 
 | Vacuum vessel deformations | ? | Movement of the target caused by the vacuum vessel movements |