## Slide 1
Dr. Mark D. Bird
Chief Technology Officer
National High Magnetic Field Laboratory
Florida State University, Tallahassee, Florida, USA
The NHMFL 40 T All-SC Magnet Project

## Slide 2
1 Lab. 3 Sites. 7 User Facilities | Users
DC @ FSU: continuous fields. | 631
Nuclear Magnetic Resonance @ FSU: solution and solid-state NMR and animal MRI. | 341
Advanced Magnetic Resonance and Spectroscopy @ UF: high-resolution and solid-state NMR, in-vivo imaging and spectroscopy, MRI microscopy, and diffusion. | 262
Ion Cyclotron Resonance @ FSU: ultra-high resolution and high mass-accuracy Fourier transform ICR mass spectroscopy. | 246
Electron Magnetic Resonance @ FSU: continuous & pulsed techniques associated with unpaired electron spins. | 173
Pulsed Field @ LANL: short, ultra-intense field pulses. | 155
High B/T @ UF: high fields combined with ultra-cold electron temperatures to 0.4 mK. | 18
Total = 1826 users in 2023

## Slide 3
Chart: Year
| 1961 | 1962 | 1963 | 1964 | 1965 | 1966 | 1967 | 1968 | 1969 | 1970 | 1971 | 1972 | 1973 | 1974 | 1975 | 1976 | 1977 | 1978 | 1979 | 1980 | 1981 | 1982 | 1983 | 1984 | 1985 | 1986 | 1987 | 1988 | 1989 | 1990 | 1991 | 1992 | 1993 | 1994 | 1995 | 1996 | 1997 | 1998 | 1999 | 2000 | 2001 | 2002 | 2003 | 2004 | 2005 | 2006 | 2007 | 2008 | 2009 | 2010 | 2011 | 2012 | 2013 | 2014 | 2015 | 2016 | 2017 | 2018 | 2019 | 2020
 | 5 | 5 | 5 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 14 | 14 | 16.5 | 16.5 | 16.5 | 17.5 | 17.5 | 17.5 | 17.5 | 17.5 | 17.5 | 17.5 | 17.5 | 17.5 | 17.5 | 18 | 18 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 21.1 | 21.6 | 22.3 | 22.3 | 22.3 | 22.3 | 23.4 | 23.4 | 23.4 | 23.4 | 23.4 | 23.4 | 23.4 | 23.4 | 23.4 | 23.4 | 23.4
 | 20.7 | 20.7 | 20.7 | 21.1 | 21.1 | 21.5 | 22.8 | 22.8 | 22.8 | 22.8 | 22.8 | 22.8 | 22.8 | 25 | 25 | 25 | 25 | 26.8 | 33.8 | 33.8 | 33.8 | 35.4 | 35.4 | 35.4 | 35.4 | 35.4 | 41.4 | 45.4 | 45.4
 | 24 | 32
The first Ultra-High Field High-Tc user magnet.
The Development of Superconducting Solenoids
MagLab 40 T
Magnets
Low-Tc User Magnets
MagSci 40 T Challenge
“A design . . . study should be conducted for the construction of a 20 T, . . . 65 cm diameter . . . magnet suitable for . . . human subject research.” – MagSci Report, 2013, pg. 18
Hi Tc Test Coils
Hi-Tc User
Hi-Strength Hi-Tc Tape
High-Tc Materials
MagLab 32 T
Field | Conductors | Maker | Location | Year
25 T | Bi-2223 + LTS | Toshiba | Sendai | 2017
25.8 T (1.1 GHz) | REBCO + LTS | Bruker | Nashville | 2019
28.2 T (1.2 GHz) | REBCO + LTS | Bruker | Florence + 12 | 2020
32.1 T | REBCO + LTS | MagLab | Tallahassee | 2017
26 T | Bi-2223 + LTS | IEE-Beijing | Huairou | 2023
30 T | REBCO + LTS | IEE-Beijing | Huairou | 2023
Existing User Magnets > 23.5 T Worldwide
There is only 1 commercial magnet supplier building magnets at 30 T, for NMR only.
Several Other Projects Underway Worldwide!

## Slide 4
Key Performance Parameters:
  Central field: 40 T
  Operating lifetime: 50,000 cycles or 20 years
  Field set-ability: ≤ 10 mT across the range of field
  Cycling repeatability: ≤ 10 mT
  Operating temperature: 4.2 K in LHe
  Cold bore diameter: ≥ 34 mm
  Ramp rate: ≥ 0.5 T/min
  Homogeneity: ≤ 500 ppm over a 1 cm DSV
  Stabilization time: < 3 minutes to stabilize at a set point
  10 Gauss fringe field: < 5.4 m from the axis of the magnet at the mid-plane
40 T All-SC Magnet Requirements
NI-REBCO not likely
REBCO SCIF generates ~0.2T
SCS & fatigue testing are important

## Slide 5: Stored Energy in HTS Coils
Chart: HTS Coil Stored Energy 40 T vs Coils prior to 2018
| NI-REBCO | I-REBCO | ICF | Bi-2212 | Bi-2223
To Date | 9 | 16 | 0 | 0.001 | 0.36
40 T | 91 | 84 | 100 | 99.999 | 99.64
26 T
32 T
25 T
34 T
As magnets get larger, stored energy increases along with potential for damage during quench!
Quench Damage
1 MJ ~ 1 stick of dynamite.
We need more development to handle increase in stored energy.
Org. | Field | Energy | Year
MIT | 30.5 T | 80 kJ NI-REBCO | 2018
RIKEN | 28 T | 30 kJ I-REBCO x 2 | 2017
Tohoku U | 25 T | 350 kJ I-REBCO | 2016
Numerous other HTS test coils around the world have destroyed themselves.
Some coils destroyed during quench

## Slide 6: Tape Magnets: Screening Currents

## Slide 7: Screening Currents: Tape Conductors
Br
Bz
Jθ
Jt
Jt = transport current in θ direction.
During charging of the magnet, Br creates screening currents, Js, in the tape.
Axis
Js
Fr
Fr
Js reacts with Bz to give radial forces Fr, in addition to those created by the transport current.
This gives rise to a Diamagnetic Torque and Twist and strain that is not uniform across the width of the tape.
Tape removed from test coils can display plastic deformation!
Photo credit: Paul Hu
Confidential FSU/NHMFL
D. Kolb-Bond, M.D. Bird, I.R. Dixon, T.A. Painter, J. Lu, K.L. Kim, K. Kim, R.P. Walsh, F. Grilli, SuST, 34, 095004 (2021).
Y. Yan, P. Song, C. Xin, M. Guan, Y. Li, H. Liu, and T. Qu, SuST. 34 (2021) 085012 (13pp)

## Slide 8
Computed Screening Currents: With and Without Rotation (2020)
“Traditional” calculation
Computed, w/ coupling
Measured
Measured
Computed, w/ coupling
“Traditional” calculation
D. Kolb-Bond, M.D. Bird, I.R. Dixon, T.A. Painter, J. Lu, K.L. Kim, K. Kim, R.P. Walsh, F. Grilli, SuST, 34, 095004 (2021).
Y. Yan, P. Song, C. Xin, M. Guan, Y. Li, H. Liu, and T. Qu, SuST. 34 (2021) 085012 (13pp)
Strain in a mini-fatigue coil (~5m REBCO)
Outboard Strain Gauge
Inboard Strain Gauge
Strains due to screening currents now agree with measurements for small coils.
Screening current induced field calculations also now agree.
Uniform j
Tiny solenoid energized in a background field.

## Slide 9
Y. Suetomi, et al., Fri-Mo-Or2-02
Impact of non-Uniform REBCO (Cu) Thickness
J. Park et al., IEEE-TAS 35, 5, Aug. 2025, 4801805, doi: 10.1109/TASC.2025.3538680.
w/o Winding Compressibility
Hoop strain [%]
w/ Winding Compressibility (REBCO, OB)
Exaggerated REBCO tape cross-section.
High radial compressibility.
MagLab end module
Traditional model = low radial compressibility
Electroplated Cu results in a non-uniform cross-section in REBCO tape.
We are adding additional co-wound steel.
4 mm
0.1 mm

## Slide 10: REBCO Challenges

## Slide 11: Unpredictability of REBCO Ic Tape-to-Tape
There is significant variation in Ic at 4 K between tapes.
It cannot be predicted from a 77 K measurement.
(Present variation seems to be less than in the past.)
84 mm
Temp | Ave | St. Dev. | St. Dev. | Min | Max | Ratio
4.2 K | 393 A | 94 A | 24% | 200 A | 725 A | 3.6
77 K | 130 A | 15 A | 12% | 100 A | 175 A | 1.8
32 T Magnet: 2012 – 2014
SuperPower tape
Dima Abraimov
4K
77K

## Slide 12: Iop/Ic of outer REBCO coil of 40 T all-SC
84 mm
60% | 63% | 64% | 63% | 60% | 56% | 53% | 51% |  | 89% | 93% | 94% | 92% | 88% | 82% | 78% | 74%
63% | 65% | 65% | 63% | 61% | 58% | 56% | 54% |  | 87% | 89% | 89% | 87% | 84% | 80% | 77% | 74%
65% | 66% | 65% | 64% | 62% | 60% | 58% | 56% |  | 85% | 86% | 85% | 84% | 81% | 78% | 76% | 73%
66% | 66% | 66% | 64% | 63% | 61% | 59% | 57% |  | 82% | 83% | 82% | 81% | 79% | 76% | 74% | 72%
67% | 67% | 66% | 65% | 64% | 62% | 61% | 59% |  | 80% | 80% | 79% | 78% | 76% | 75% | 73% | 71%
67% | 67% | 67% | 66% | 65% | 63% | 62% | 60% |  | 78% | 78% | 77% | 76% | 75% | 73% | 71% | 70%
67% | 67% | 67% | 66% | 65% | 64% | 62% | 61% |  | 76% | 76% | 75% | 74% | 73% | 72% | 70% | 69%
67% | 67% | 67% | 66% | 65% | 64% | 63% | 61% |  | 74% | 74% | 74% | 73% | 72% | 71% | 69% | 68%
67% | 67% | 67% | 66% | 65% | 64% | 63% | 62% |  | 73% | 73% | 72% | 72% | 71% | 70% | 68% | 67%
66% | 66% | 66% | 66% | 65% | 64% | 63% | 62% |  | 71% | 71% | 71% | 71% | 70% | 69% | 68% | 66%
66% | 66% | 66% | 66% | 66% | 65% | 63% | 61% |  | 70% | 70% | 70% | 70% | 70% | 69% | 67% | 65%
65% | 65% | 66% | 66% | 66% | 65% | 63% | 61% |  | 69% | 69% | 69% | 69% | 69% | 68% | 66% | 65%
66% | 66% | 67% | 67% | 66% | 65% | 64% | 63% |  | 68% | 68% | 68% | 68% | 68% | 67% | 65% | 64%
66% | 66% | 66% | 66% | 66% | 65% | 64% | 63% |  | 67% | 67% | 67% | 67% | 67% | 66% | 65% | 63%
66% | 66% | 66% | 66% | 66% | 65% | 64% | 63% |  | 66% | 66% | 66% | 66% | 66% | 65% | 64% | 63%
66% | 66% | 66% | 66% | 66% | 66% | 65% | 63% |  | 66% | 66% | 66% | 66% | 66% | 65% | 64% | 62%
65% | 65% | 65% | 65% | 66% | 66% | 64% | 62% |  | 65% | 65% | 65% | 65% | 66% | 66% | 64% | 62%
63% | 64% | 64% | 64% | 65% | 66% | 64% | 61% |  | 63% | 64% | 64% | 64% | 65% | 66% | 64% | 61%
63% | 63% | 63% | 63% | 64% | 66% | 64% | 61% |  | 63% | 63% | 63% | 63% | 64% | 66% | 64% | 61%
64% | 64% | 64% | 64% | 65% | 67% | 65% | 62% |  | 62% | 62% | 62% | 62% | 63% | 65% | 63% | 60%
67% | 66% | 66% | 66% | 66% | 67% | 66% | 63% |  | 61% | 60% | 60% | 60% | 61% | 61% | 60% | 58%
68% | 68% | 67% | 68% | 68% | 68% | 67% | 65% |  | 58% | 58% | 58% | 58% | 59% | 58% | 58% | 56%
68% | 68% | 68% | 68% | 68% | 68% | 67% | 65% |  | 56% | 56% | 56% | 56% | 56% | 56% | 56% | 54%
68% | 68% | 68% | 68% | 68% | 68% | 67% | 66% |  | 55% | 55% | 55% | 55% | 55% | 55% | 54% | 53%
68% | 68% | 68% | 68% | 68% | 68% | 67% | 66% |  | 53% | 53% | 53% | 53% | 53% | 53% | 52% | 51%
68% | 68% | 68% | 68% | 68% | 68% | 67% | 66% |  | 52% | 52% | 52% | 52% | 52% | 52% | 51% | 50%
68% | 69% | 69% | 69% | 69% | 69% | 68% | 66% |  | 50% | 51% | 51% | 51% | 51% | 51% | 50% | 48%
68% | 68% | 69% | 69% | 69% | 69% | 68% | 66% |  | 49% | 49% | 49% | 49% | 49% | 49% | 49% | 47%
68% | 68% | 69% | 69% | 69% | 69% | 68% | 66% |  | 47% | 47% | 47% | 47% | 47% | 47% | 47% | 45%
69% | 69% | 69% | 69% | 69% | 69% | 69% | 67% |  | 45% | 45% | 45% | 45% | 45% | 45% | 45% | 44%
69% | 69% | 68% | 68% | 68% | 69% | 69% | 67% |  | 44% | 43% | 43% | 43% | 43% | 43% | 43% | 43%
69% | 68% | 68% | 68% | 68% | 69% | 69% | 68% |  | 42% | 42% | 42% | 42% | 42% | 42% | 42% | 41%
69% | 68% | 68% | 68% | 69% | 69% | 70% | 68% |  | 41% | 40% | 40% | 40% | 40% | 41% | 41% | 40%
68% | 68% | 67% | 67% | 68% | 69% | 70% | 68% |  | 39% | 39% | 39% | 39% | 39% | 40% | 40% | 40%
66% | 66% | 65% | 65% | 66% | 68% | 71% | 68% |  | 38% | 38% | 38% | 38% | 38% | 39% | 41% | 39%
67% | 66% | 65% | 65% | 66% | 67% | 71% | 68% |  | 37% | 37% | 36% | 36% | 37% | 38% | 39% | 38%
68% | 67% | 66% | 65% | 65% | 66% | 67% | 66% |  | 36% | 36% | 35% | 35% | 35% | 35% | 36% | 35%
66% | 65% | 64% | 63% | 62% | 62% | 62% | 62% |  | 36% | 35% | 34% | 34% | 34% | 33% | 33% | 33%
65% | 63% | 62% | 61% | 60% | 59% | 59% | 58% |  | 35% | 34% | 33% | 33% | 32% | 32% | 32% | 31%
65% | 63% | 62% | 60% | 59% | 58% | 57% | 57% |  | 34% | 33% | 32% | 32% | 31% | 31% | 30% | 30%
64% | 62% | 60% | 59% | 57% | 56% | 55% | 54% |  | 34% | 33% | 32% | 31% | 30% | 29% | 29% | 29%
62% | 60% | 58% | 56% | 55% | 53% | 52% | 51% |  | 33% | 32% | 31% | 30% | 29% | 28% | 28% | 27%
61% | 59% | 57% | 55% | 53% | 52% | 50% | 50% |  | 33% | 32% | 31% | 30% | 29% | 28% | 27% | 27%
61% | 59% | 56% | 54% | 52% | 51% | 49% | 48% |  | 33% | 31% | 30% | 29% | 28% | 27% | 26% | 26%
61% | 58% | 56% | 54% | 52% | 50% | 48% | 47% |  | 32% | 31% | 30% | 29% | 28% | 27% | 26% | 25%
59% | 57% | 54% | 52% | 50% | 48% | 46% | 46% |  | 32% | 31% | 30% | 29% | 27% | 26% | 25% | 25%
59% | 56% | 54% | 52% | 50% | 48% | 46% | 45% |  | 32% | 31% | 30% | 28% | 27% | 26% | 25% | 25%
58% | 55% | 53% | 51% | 49% | 47% | 45% | 44% |  | 32% | 31% | 30% | 28% | 27% | 26% | 25% | 25%
Axis of rotation
Mid-Plane
If Iop/Ic varies, quench protection is harder,
SCS can be excessive.
Ic | Min | Max
Uniform | 25% | 94%
graded | 44% | 71%
Ic variation is helpful.

## Slide 13: Unpredictability of REBCO Ic Dropouts
84 mm
77 K measurements with TapeStar & YateStar indicate there can be spots along a conductor with low critical current.
Similar measurements at 4K have not been made.
Variation at 4K is unknown.
YateStar measurement of critical current at 77K of SuperPower tape M4-352-5 0912
Paul Hu
MagLab (& others) are using multiple tapes in parallel.

## Slide 14: Ic Variation at low angles
84 mm
J. Jaroszynski, et al., arXiv:2502.02706v2, 11 Feb. 2025. Sat-Mo-Or3-05
It is difficult to measure Ic of REBCO tape at 4 K at angles close to the ab-plane.
For MagLab 32 T magnet we adopted measurements at 18° from the ab-plane, and extrapolation.
>1/2 of 32T & 40T operate at <5°.
There is typically a variation in microstructure along the edge of REBCO (12 mm) tape as it is fabricated.
Recent measurements suggest that this variation might contribute to significant variation in Ic along the length of 4 mm tapes cut from the edges as opposed to the center.
Transport measurements at appropriate angle, field, temperature are rare.

## Slide 15: No-Insulation REBCO Challenges & Status

## Slide 16: Quench
NI-REBCO shows similar current spikes to those of diode-protection.
Many more degrees of freedom.
Markiewicz, et al., SuST, 29, (2016) 025001. Markiewicz, et al., SuST, 32 (2019) 105010
During quench, current can flow radially between turns.
Quench is inductively driven
  As current drops in a turn, increased current will be induced in nearby turns, driving them normal.
This can result in large current spikes with large forces (concentration of stored energy).
Large NI-REBCO coils require controlled inter-turn resistance (and active protection?)
Each turn of an NI-REBCO coil includes:

## Slide 17: RI-REBCO Test Coils at the MagLab
PTC1
(Apr 2019)
LBC-AR
(Apr 2019)
LBC-CO
(Aug 2019)
PTC2
(Jun 2019)
PTC1R
(Oct 2019)
Mini fatigue coil
– Off centered
(Jun 2020)
PT3
(July 2020)
PTC4
(Dec 2020)
Mini fatigue coil
– centered
(Mar 2021)
PTC5
(Mar 2021)
Axial pressure
test coils
(Mar 2021)
Axial pressure
test coil:
two in hand
(April 2021)
PT6
(June 2021)
Axial pressure
test coil:
two in hand
(April 2021)
RI-NC (Sep. 2022)
158 mm
2022: Were not able to consistently control the quench dynamics

## Slide 18: I-REBCO Technology for 40 T SC

## Slide 19: I-REBCO Coils at the NHMFL
TC0
(October 2020)
Axial pressure
test coils
(Mar 2021)
Axial pressure
test coil:
two in hand
(April 2021)
Axial pressure
test coil:
two in hand
(April 2021)
TC2 (August 2022)
32 T
(2015)
LSC
(Nov 2024 &
Oct 2025)
320 mm
232 mm
TC1
(May 2021)
140 mm
Axial pressure & TC are on same scale

## Slide 20
Large Scale Coil Test Oct. 7 – 17, 2025
Troy Brum
298 mm
161 mm
The coil reached the full design current of 645 A in a background field of 11.4 T, corresponding to 70% of the coil’s critical current.
3: 600 A / BM 11.2 T / PFN 500 V
7 intentional quenches were conducted with various combinations of background field, insert current, and PFN voltage.
Kwangmin Kim, Iain Dixon, Yu Suetomi, and Dharmendra Shukla reviewing test data.
Iain Dixon, Yu Suetomi, et al.
No Degradation was observed!
Measured voltages agreed well with predictions!

## Slide 21
The 40 T SC Magnet at MagLab
Liz Green
Troy Brum
298 mm
Developed software packages to compute screening currents, strains, and field distributions that are being adopted by other labs worldwide.
Developed quench modelling software for both I-REBCO and RI-REBCO coils.
Introduced critical-current graded coils.
Demonstrated coils survive
  50,000 cycles with 125 MPa axial pressure
  >23,000 cycles at 0.4% strain; >200 cycles at 0.5% strain.
  Multiple quenches at high JCu, Iop/Ic, & strain.
Increase in copper current density, Jcu, compared with 32 T.
Introduced Reinforcement grading.
Demonstrated numerous improvements to quench protection
Tested >20 different REBCO coils.
161 mm
Presently have funding for Preliminary & Final Design
Expect to submit proposal for Construction in 2026.
Anticipate Construction funding to arrive 2027.
Start placing orders for tape, etc. at that point.

## Slide 22
Alex Downey,
Thank You!
Alex Downey, Scott Marshall, Robert Stanton, Kathleen Amm, Shannon Griffin, Dharmendra Shukla, Kwangmin Kim, Yu Suetomi, Lee Marks, Tom Painter, Justin Deterding, Erick Arroyo, Iain Dixon, John Rogers, Hongyu Bai, & Jun Lu celebrate the successful LSC test.

## Slide 24: Summary
High Magnetic Fields play an important role in numerous types of basic research in physics, chemistry, biology, materials science, etc.
The HTS materials have (finally) started to enable higher fields.
There are several REBCO-based magnets in operation at 28 – 32 T around the world.
There are still a number of challenges that require further development.

## Slide 25: Some NI- & RI-REBCO High Field Magnets & Test Coils
Design | Site | Field | Bore | Year | Status | Protection
Magnets put into Service |  |  |  |  |  | 
IEE-Beijing | Huairou, China | 30 T | 32 mm, 4 K | 2023 | Operational | Passive
IEE-Beijing | Huairou, China | 27 T, 12 ppm/cm | 50 mm, 4 K 30 mm, 1.4 K 27 mm, 20 mK | 2023 | Operational | Passive
Test Coils (Hahn’s list is more comprehensive) |  |  |  |  |  | 
MIT-SuNam | Daejeon, Korea | 26 T | 35 mm, 4 K | 2015 | Out of Service | Passive
MagLab | Tallahassee, FL, USA | 45.5 T | 14 mm, 4 K | 2018 | Test | Passive
LNCMI-CEA | Grenoble, France | 32.5 T | 38 mm, 4 K | 2019 | Test | Passive
Tsingua U. | Beijing, China | 32.4 T |  | 2024 | Test | Passive
ETH | Zurich, Switzerland | 45 T | 4 mm, 4 K | 2024 | Test | Passive
Projects Underway |  |  |  |  |  | 
CAS-IPP | Hefei, China | 46 T? |  | 2025 |  | Passive?
MIT | Boston, MA, USA | 30.5 T | NMR |  |  | Passive
CERN | CERN | 40 T | 50 mm, 4 K |  |  | Active
LNCMI-CEA | Grenoble, France | ≥30 T |  |  |  | Passive

## Slide 26: I-REBCO Magnets & Test Coils
Design | Site | Field | Bore | Year | Status | Protection | 
Magnets Put Into Service |  |  |  |  |  |  | 
MagLab | Tallahassee | 32 T | 34 mm, 4.2 K | 2017 | Repair Underway | Heaters between modules. | 
Bruker | Nashville | 25.8 T, 10 ppb | 56 mm, 300 K | 2018 | Operational | ? | 
Bruker | Florence + 12 | 28.2 T, 10 ppb | 56 mm, 300 K | 2019 | Operational | ? | 
Test Coils |  |  |  |  |  |  | 
RIKEN | ? | 27.8 T | 68 mm. 4.2 K | 2018? | Test | ? | 
Tohoku U | Sendai, Japan | 25 T |  | 2022 | Test | ? | 
Projects Underway |  |  |  |  |  |  | 
MagLab | Tallahassee | 40 T | 34 mm, 4.2 K | 2031? | Design nearing Completion. |  | Similar to 32 T.
Tohoku U | Sendai, Japan | 33 T |  | 2027 | LTS fabrication complete, HTS underway. |  | External resistor to limit current-rise.
Bruker |  | 32.9 T, 10 ppb |  |  |  |  | ?
Bruker |  | 35.2 T, 1 ppm |  |  |  |  | ?

## Slide 27: Options for 40 T all-SC Magnet