## Slide 1
Setup of a probe station for FCC
and first measurement
Yau King Ying (Harrison)

## Slide 2: Table of Content
Today’s Discussion
Motivation for my project
  What is FCC?
  Why do we need a probe station for FCC?
  Setting up analysis code
Objectives and goals
  Software development
  Setting up fully functioning probe station
  Validation measurements
Research methodology
Results
Conclusions and future prospects

## Slide 3: Motivation for my project
What is FCC?
Future Circular Collider
  Pushing energy and intensity frontiers of particle colliders
  Searching for new physics at higher collision energy (
Design
  90.7 km circumference tunnel is used
  Colliders:
    Electron-positron collider (FCC-ee)
    Hadron collider (FCC-hh)
  Four interaction points for experiments

## Slide 4: Motivation for my project
Why do we need a probe station for FCC?
Performance optimization
Component validation
R&D for new sensor
Explore new technologies for FCC
  LGADs (timing sensors)
  AC-LGADs
For example: Timing at the FCC-ee
Improve event reconstruction
Provide a time-of-flight measurement for particle ID
Improve momentum resolution in tracking system

## Slide 5: Objectives and goals
Objectives and goals
Setting up fully functioning probe station
  Obtaining correct measurement results
  Functioning under low temperature (
Software development
  Improving user interface
  Adding buttons that we may need in the future
  Analysing our data
Validation measurements
  Qualifying our probe station
  Cross checking earlier sensor measurements

## Slide 6: Research methodology
How do we accomplish our goals?
Setting up fully functioning probe station
  Fixing hardware issues
  Implementing a proper temperature control unit
Software development
  Adapting existing code from other institute (very old) to our setup
  Adding buttons that we may need in the future
  Adopting the old analysis script
Validation measurement
  Comparing our results with past measurements
  Measuring at different temperature

## Slide 7: Research methodology
Setting up the probe station
Full setup
Systematic Diagram of the probestation in LHEP Uni Bern

## Slide 8: Research methodology
Setting up probestation
Temperature control
Setting up the probe station
Temperature control
Chiller
Temperature monitor(old)
Arduino setup
Power supply for peltier
Instead of using this…

## Slide 9: Research methodology
Setting up the probe station
At low temperature()
Before
After
tank
A Lid for the chuck

## Slide 10: Research methodology
Software Development
GUI

## Slide 11: Research methodology
Software Development
Make the old analysis script functioning
Adding temperature conversion equation
Adding options for plotting with different scales

## Slide 12: Research methodology
Validation measurements
Probing several pixel sensors using our setup
  Comparing our results with past measurements
PIN

## Slide 13: Research methodology
Sensors at LHEP Uni Bern
PIN

## Slide 14: Research methodology
What is a PIN diode?
p-type and n-type regions
  P-doped: Region with a surplus of holes
  Intrinsic semiconductor layer
  N-doped: region with a surplus of electrons
Applying voltage
  Current increases until full depletion
  If voltage keep increasing breakdown voltage
Pixel structure sensor
Higher spatial resolution
Reduced pile-up effects
PIN

## Slide 15: Research methodology
Validation measurement
Measuring our sensors at different temperatures
  i.e.
Temperature conversion equation
  ,
  Minimising effects due to the temperature dependence of the leakage current
  Comparing results at different temperature
Old result at
Our result at
Old result at

## Slide 16: Results(PIN diodes)
Dependency of IV curves with different parameters
More Overdrive
Less Overdrive
Consistency check
No. of needles
Overdrive
Temperature
I04
2 needles
I04
1 needles

## Slide 17: Results(PIN diodes)
Measurement at 20℃
Pad Current
VS
Total Current
New Data
VS
Old Data
K04 Pad current(old)
K04 Pad current(new)
F10 Pad current(old)
F10 Pad current(new)
K04 Pad current
K04 Total current
F10 Pad current
F10 Total current

## Slide 18: Results(PIN diodes)
Studying IV curves of old data with different temperatures
Direct
Measurement
Converting
to
K04 Pad current(20)
K04 Pad current(−20)
F10 Pad current(20)
F10 Pad current(−20)
K04 Pad current(20)
K04 Pad current(−20)
F10 Pad current(20)
F10 Pad current(−20)

## Slide 19: Results(PIN diodes)
Rescaling the new measurements to match the old measurements
25℃-> 20℃
27℃-> 20℃
F10 Pad current(new)
F10 Pad current(old)
K04 Pad current(new)
K04 Pad current(old)

## Slide 20: Results(PIN diodes)
Studying the effect on IV curves with different parameters
Ramp speed
= 100ms
Ramp speed
= 200ms
Ramp speed
= 400ms
Ramp volt
= 5V
Ramp volt
= 10V
Ramp volt
= 20V
Wait time
= 500ms
Wait time
= 50000ms

## Slide 21: Results(pixel sensors)
S1P1 and S1P2
Comparing old and new measurement at 20℃
S2P1 and S2P2

## Slide 22: Conclusions and future prospects
What have I accomplished?
Contributed to setting up a functional probe station for FCC sensor measurements
Revived a set of functioning analysis scripts
Produced results for analysis and comparison
Developed a better user interface
What needs to be done?
Get cooling automated for the chuck
Understand the disagreement between our results and old results
Clean up the code
Test new sensors

## Slide 23: THE END
Thank you for listening.
Any questions?