## Slide 1: LHC Beam Synchrotron Radiation Telescope for Imaging Status, Performance, and HL-LHC Outlook
F.Roncarolo , D.Butti, S.Burger, J.Pucek
C.Pasquino, W.Andreazza
T.Lefevre , R.Veness
HL-LHC Beam Profile Review – 23-Sep-2025

## Slide 2: LHC Beam Synchrotron Radiation Monitor (this morning)

## Slide 3: Contents
Instrument principles, design, current and expected performance
Towards Run4 Exploitation: Issues, mitigations, risks et al.
Template questions to easy compare to other monitors
Summary
N.B.: todays review (beam profile monitoring) and this presentation focus on operational synchrotron radiation monitor --managed with SY-BI exploitation budget, no HL-LHC costs

## Slide 4: Glossary
BSR — Beam Synchrotron Radiation
BSRTM — BSR mirror-based extraction (legacy/operational line in IR4) In-vacuum mirror system that extracts visible/UV SR from: undulators at 450 GeV and D3 dipoles above ~2 TeV. Light is routed to the optical benches under the beam pipe in RA43/47. Serves BSRT (transverse profile), BSRA, and BSRL. No HL-LHC deliverables
BSRTMB — BSR mirror-based extraction, new “B” line (HL-LHC deliverable) Upgraded extraction from D4 dipoles with light transported to the UA43/47 caverns. Prototype in LS2, full implementation in LS3. Designed to serve BSRH (halo). Potential for transverse and longitudinal profiles redundancy at flat top energy
BSRT — Beam Synchrotron Radiation Telescope Transverse beam size/profile monitor imaging SR onto a gated, intensified CMOS/CCD chain. Provides bunch-resolved emittance/size, optics studies, and routine beam quality monitoring.
BSRH — Beam Synchrotron Radiation Halo Monitor High-dynamic-range SR instrument to measure beam halo and core-to-halo contrast at flat-top, leveraging the BSRTMB extraction and dedicated optics/apertures.
BSRA — Beam Synchrotron Radiation Abort-Gap Monitor Gated SR diagnostic measuring residual charge in the abort gap (and near-by empty buckets). Used to protect the dump system and assess cleaning efficiency.
BSRL — Beam Synchrotron Radiation Longitudinal Monitor SR-based longitudinal diagnostic for ghosts and satellite bunches and overall bunch-length/structure characterization (complements BWS/BPM-based info).
BSR Extraction Systems/Tanks
BSR Monitors

## Slide 5: LHC Synchrotron Radiation (SR) monitors
Non-invasive diagnostics for transverse, longitudinal and intensity beam parameters
 | SR extraction Stations | 
 | BSRTM (Operational) | BSRTMB - HL-LHC (for ref., not for today’s review)
Location | Left and Right of IR4, for beams existing RF zone | Left and Right of IR4, for beams entering RF zone
SR sources | 450 GeV: Undulators (visible SR) >2 TeV: D3 dipoles (visible/UV SR) | >2 TeV: D4 dipoles (visible/UV SR)
Light transport | RA43/47 tunnel: To optical setup & detectors under beam pipe | UA43/47 service caverns: to new optical hutch, easier access, low-rad environment
Operational status | Operational since start of LHC. Will stay in Run4 | Prototype (tank+mirror only) installed LS2, tested Run3, final implementation LS3 (HL-LHC deliverable)
Main diagnostics served | BSRT: transverse beam profile BSRA: abort gap population BSRL: ghosts & satellites | BSRH: beam halo monitor (Transverse beam profile, Abort gap, Longitudinal profiles possible as redundant to BSRTM + R&D, not baseline)

## Slide 6
BSR extraction systems
SR
S
X
SR
To detectors below beam pipe
FIXED MIRROR
SR
BEAM
Y
S
SR
To detectors in service gallery
MOVABLE MIRROR
Operational since LHC start (BSRT, BSRA, BSRL)
HL-LHC for beam halo Run4
(BSRH)
Side View
Top View
For transverse and longitudinal beam profile (BSRTM)
For transverse beam halo (BSRTMB)

## Slide 7: BSR Systems Layout/Synoptics

## Slide 8: BSR Monitors optical table below beam pipe
Transverse beam profile line
Abort Gap and Long Profiles

## Slide 9
BSR monitors optical table

## Slide 10: BSR profile monitor - Current Performance
Absolute calibration vs Wire Scanners (~10% accuracy)
Robust bunch-by-bunch, machine operation & MDs
Downtime rare
  < few events per year resulting in missing monitoring for large part or entire LHC fill(s)
  Radiation induced SEU in digital electronics and cameras
    Automatic watchdog and reset make these have a minor impact on overall system availability
    Occurrence rate increased beginning of Run3 --partially mitigated, see later slide -- more radiation expected with HL-LHC beams
Since Run3: Used for tails characterization (q-Gaussian fits)
  Sofia Kostoglou’s talk, also for more numbers on performance

## Slide 11: BSR profile monitor - Limitations
BSRT operates at different wavelengths at injection and flat top energies
  Visible range @ injection: Offers robustness and long-term stability but at the cost of reduced spatial resolution (not a problem for mm-like beams at injection
  UV range @flat top. Offers higher resolution but: special –often custom made- optics, more expensive than for the visible, more sensitive sensitivity to ageing
    Calibration drifts at flat-top (optical ageing)
    Even with reduced diffraction effects, resolution ~=200um, similar to beam size (1sigma).
Energy ramp: unreliable
Ions: no signal at injection (450 ZGeV)

## Slide 12: BSR profile monitor - calibration drift example
BSR imaging vs LHC emittance scans (from LHC experiments , convoluting B1 and B2)

## Slide 13: Issues – General Summary
Component | Ageing Signature & Observations | Mitigation (in use)
Image Intensifier Photocathode | Local QE loss / non‑uniform response at SR spot; progressive sensitivity drop | “Spot painting” (translate camera on motorized X–Y stage to spread dose); keep MCP gain and gate width as low as practical. Image intensifier is a ‘consumable’ to be regularly replaced.
Image Intensifier MCP gain chain | Reduced gain over time; slower gain dynamics if routinely driven at high voltage | Operate at the lowest gain compatible with SNR; reserve high‑gain modes for special studies
Vacuum viewport / windows | Surface etching / haze → increased scatter, throughput loss | Cleanliness control; scheduled inspections and replacements
Optical filters | Coating burn marks and transmission drop when SR is focused on the filter | Place filters at a less‑focused plane (larger spot, lower W/mm²); share attenuation with upstream ND filters; monitor and replace as needed
In-Vacuum extraction mirror | Heating due to beam-coupling impedance. Major failure before LS2 | From metallic to dielectrics coating, LS2 BSRTM re-design, new BSRTMB concept
Radiation effects on electronics | Digital Cameras and other electronics SEU | New relay boxes with radiation tolerant components. Watchdog / Reset of cameras

## Slide 14: Issues – Specific Examples (I)
 | Issue / Observation | Cause / Mechanism & Mitigation
 | Image‑intensifier non‑uniform response | High SR photon dose concentrated on a fixed region of the photocathode; secondary MCP gain drift over time  Spot painting  MD: freeze pos. to ensure no response drifts
 | Viewport surface etching / haze (increased scatter, throughput loss) | Prolonged UV/SR exposure and thermal load on the vacuum viewport; surface chemical activation/contamination on the air side  Scheduled maintenance
 | 250 nm band‑pass filter damage | Focused SR spot concentrating power on the filter, local coating damage and spectral shift  Relocate filter to defocused spot , optimize ND filter use
G.Trad et al., 2019
From YETS to TS1
From TS1 to TS2
Painting map
SR footprint

## Slide 15: Issues – Specific Examples (II)
 | Issue / Observation | Cause / Mechanism & Mitigation
 | Extraction mirror coating detachment at specific location | X-Ray energy deposition from undulator emission at LHC top energy  Undulators (only used for imaging at injection) switched off during beam energy ramp
SR on mirror
Coating damage
Flatness profile measurement of damaged extraction

## Slide 16: Expected Performance at HL-LHC
No major LS3 upgrades for operational BSR profile monitor
LS2 re-design and dielectric mirrors solved impedance/heating
  On the paper compatible with HL-LHC beam spectra and power (see Carlo Zanini’s talk)
  See also next slides on temperature measurements
BSRTMB extraction system for halo coverage, offers possibility of redundancy at flat top

## Slide 17: Risks – Heating due to beam coupling impedance
General (recap from earlier table): present systems design based on beam coupling impedance simulations and RF lab measurements  Optimized geometry, dielectric coating. Re-iteration on HL-LHC beam parameters effects on-going (see C.Zanini’s talk)
Temperature measurements
Instrumentation: In-vacuum thermocouples mounted on the mirror support and near the mirror base.
Acquisition & Alarms: Continuous online logging since 2023, supervised by an automatic watchdog (thresholds & alerts).
Representativeness: Sensor readings track actual mirror temperature (validated during bake-out).
Operational Result: During physics operation, temperatures remain essentially unchanged—no measurable beam-induced heating with the current low-impedance design
See next slide

## Slide 18: Temperature Measurements
2025
T is correlated to tunnel heating/cooling not to beam power
2 probes , in contact with mirror and mirror support

## Slide 19: ‘Template’ Questions (I)
Accuracy:
10% accuracy on emittance, after calibration to BWS
Can your technique provide:
1D beam profiles? ✅yes
2D beam images? ✅yes
Fastest acquisition rate:
Bunch per Bunch selection: ✅ yes (via the intensifier gate, typically 20–25 ns, synced to the RF).
Single-turn exposure: technically possible (open the gate once during a camera frame which lasts ~260turns), but not how we usually run because of photon statistics/SNR. In practice we integrate multiple identical gates over successive turns for the same bunch and build one image per frame  more infos in next slide

## Slide 20: 📑 BSRT Acquisition Parameters
Component / Item | Parameter | Value | Notes
Image Intensifier (Hamamatsu C14245) | Minimum gate ON | 5 ns | Typical operation 20–25 ns (bunch selection)
 | Maximum repetition rate | 30 kHz (≈33 µs period) | Max gating pulses/s
Camera (Basler acA1920-40gm) | Minimum exposure time | 34 µs (≈29 kHz equiv.) | Shortest possible integration window
 | Maximum frame rate | 42 fps (≈24 ms period) | Full frame (1920×1200); scales with ROI
BSRT acquisition (Run 3) | Injection | ~2 × 8 Hz = 16 Hz (≈62 ms period) | Two frames per bunch, 8 and 16 Hz to have margin on S/N (see also photo budget below) Processing bottleneck (FESA/publishing), not photon statistics
 | Flat-top | ~2 × 13 Hz = 26 Hz (≈38 ms period)  80sec to meas 2000 bunches | 
Photon budget (flat-top, UV) | Neutral density filters | ND1/ND2 | Removing filter increases light → higher acquisition rate possible
 | Boost mode (concept) | Higher rate possible | At expense of photocathode lifetime
LHC revolution period | 89 µs (≈11.2 kHz)

## Slide 21: ‘Template' Questions (II)
What is the minimum number of particles needed per bunch?
Works down to pilot intensity (~5×10^9 p/b). Lower also possible at cot of exposure/integration time
Can it measure both protons and ions?
Protons: ✅ (all along cycle)
Ions: ✅ at top energy, ❌ at injection
Does your technique require calibration?
Yes: vs BWS, at least ~2×year. Recent drifts of beam1: more often, as agreed/requested by OP and experiments
Calibration meant to determine both optical magnification and point-spread-function, requires multi points == special fills with different emittances
If beam size at detector increased, would it gain/lose performance?
Larger beams → improved calibration stability and resolution margin

## Slide 22: Q&A Summary Table
Template Question | Answer | Comments / Limitations
1D beam profiles? | ✅ Yes | Projection of SR image, calibrated vs BWS
2D beam images? | ✅ Yes | Full SR images, resolution-limited at flat-top
Fastest acquisition rate? | Single-turn, 20 ns gated camera | ~10% accuracy after calibration
Bunch-by-bunch measurements? | ✅ Yes | Integrating multi turns
Turns needed? | 260 | Margin for going faster exists
Minimum particles/bunch? | ~5×10⁹ protons | Below pilot intensity SNR deteriorates
Protons and ions? | Protons ✅ (all cycle), Ions ✅ (top energy), ❌ at injection | No visible SR from ions at injection
All along LHC cycle? | Protons: OK except ramp; Ions: not at injection | Ramp: strong SR source variation
Calibration required? | ✅ Yes, vs BWS (~2×/year) | Drift at flat-top due to optical ageing
Effect of larger beam size? | ✅ Gain | Improved calibration stability and resolution margin

## Slide 23: Summary & Outlook (I)
Operational Experience (also profiting of Sofia’s talk)
BSR has proven essential for routine beam profile/emittance and beam quality monitoring since Run 1
High availability and reliability, with only rare interruptions per year.
Successfully used for tail characterization and advanced beam studies (q-Gaussian fits)
Performance
Absolute calibration to BWS within ~10% accuracy
Robust bunch-by-bunch profiles at injection and flat-top
Optical resolution limited at flat-top (~200 µm), but adequate for HL-LHC beams
Ion operation: no signal at injection

## Slide 24: Summary & Outlook (II)
Issues & Risks
Optical ageing and calibration drifts at flat-top → mitigated by regular BWS reference scans
If drifts confirmed to be dominated by UV optical components:
  Is re-designing flat top BSR optics for visible an option ?
Radiation effects (SEUs, optics damage) → watchdogs, rad-tolerant boxes, and spares strategy
Heating risk from beam coupling impedance → ruled out with redesigned dielectric mirrors and continuous temperature monitoring – re-checks ongoing
HL-LHC Outlook
Present BSRTM line remains operational and compatible with HL-LHC beams
BSRTMB line (HL-LHC deliverable) enables beam halo diagnostics and potential for redundancy at flat-top
Regular BWS calibration campaigns will remain crucial for accuracy
Long-term consumables (intensifiers, cameras, optics) require continued replacement strategy (exploitation budget)
System remains a key non-invasive diagnostic for HL-LHC, with expanded role in halo monitoring and advanced profile characterization

## Slide 25
SPARE

## Slide 27: BSRT Calibration

## Slide 28: SR spectra and x-ray absorption length

## Slide 29: BSRTM