Digital Radiation Tolerant Camera 
Aiming At Beam Particle Imaging

BeamCAM

| Document Title: | Test Plan and FDIR for Charm |
| --- | --- |
| Document N°: | MCSE-BCAM-PL-002 |
| Issue: | 1.0 draft |
| Date: | 10.10.2025 |

| Classification | Status |  | Export Control | Status |
| --- | --- | --- | --- | --- |
| Public |  |  | Unclassified | X |
| Internal |  |  | EU / Switzerland | X |
| Proprietary | X |  | US EAR |  |
| Confidential |  |  | US ITAR |  |

PROPRIETARY INFORMATION

|  | Name | Function | Date | Signature |
| --- | --- | --- | --- | --- |
| Prepared by | Quentin Bouille | HW Engineer | 10.10.2025 |  |
| Verified by | Daniel Krähenbühl | Technical Manager | 10.10.2025 |  |
| Released by | Marc Jeanmonod | Project Manager | 10.10.2025 |  |

Document Change Record

| Issue | Date | Changes Description |
| --- | --- | --- |
| 1.0 | 10.10.2025 | Initial issue |

Distribution List

| Name | Organisation | Copies |
| --- | --- | --- |
| S. Burger | CERN | 1x electronic copy |

Content

1	Applicable and reference documents	5

1.1	Applicable documents	5

1.2	Reference documents	5

1.3	Norms, standards	5

2	Acronyms and definitions	6

2.1	Acronyms	6

2.2	Definitions	6

2.2.1	Data sizes	6

2.2.2	Data rates	6

3	Scope	7

4	Irradiation test presentation	8

4.1	Irradiation test objectives	8

4.2	Irradiation characteristics	8

4.3	Test main timeline	8

4.4	Test Sequences	8

4.5	Test SW operations flow	9

4.6	Keys monitoring parameters or data	9

5	Device Under Test (BeamCAM EBB) presentation and set-up	10

5.1	Description	10

5.2	Block Diagrams	10

5.3	Mechanical Information	10

5.4	Proof of concept evolution / Components improvements	11

5.5	Evolution of the power topology	12

5.6	BeamCAM EBB Connectors identification	13

5.7	Interfaces Connections at Camera level	13

5.8	Electrical schematic of specific connectors	14

5.8.1	Barrel HR10 (specific to BeamCAM EBB)	14

5.8.2	Trigger SMA	14

5.8.3	Power 4MM Plug Bananas	14

5.8.4	Ethernet RJ45	14

5.9	BeamCAM EBB HW settings	15

5.10	BeamCAM EBB FW settings	15

5.11	BeamCAM EBB identification	15

6	Test premises	16

6.1	CHARM Test Chamber CERN	16

6.2	Preparation room 157/1-N26	16

6.3	Beam area	17

6.3.1	Conveyor	17

6.3.2	Panels	18

6.3.3	Personnel preparation	19

6.3.4	Installation in the Beam Area	19

6.3.5	Precautions to be taken during the Beam test	19

6.3.6	End of the test	19

7	Test Set-up	20

7.1	Diagram Test Set-up	20

7.2	Test Equipment	21

7.3	Coax cable (item 11) solutions	21

7.3.1	Solution 1: SMA and 4mm to BNC	21

7.3.2	Solution 2: HR Barrel to 4mm then to BNC	21

7.4	As-run Set-up installation	21

7.5	Previous Test-Set-up Installation	21

7.6	Test set-up support for conveyer	22

7.7	External (MCSE) equipment needed in Control Room	23

8	FDIR during Test	24

List of figures

Figure 1: test timeline	8

Figure 20: Test SW operations flow	9

Figure 2: BeamCAM EBB	10

Figure 3: Power Path	12

Figure 4 Connectors IFs of EBBv4	13

Figure 5: J4 Schematic	14

Figure 6: J5 schematic	14

Figure 7: 4mm plugs schematic	14

Figure 8: ETH connector schematic	14

Figure 9: Top: CHARM facility irradiation layout, with the BeamCAM setup in location ’10’. Bottom: Two BeamCAM systems, each with a dedicated black-and-white target to probe visibility.	16

Figure 10: preparation room (general view)	17

Figure 11: preparation room data acquisition patch panel	17

Figure 12: preparation room irradiation panel	17

Figure 13: Overhead conveyor description	18

Figure 14: Data acquisition panel	18

Figure 15: Irradiation panel	18

Figure 16: test set-up in the Beam control room	20

Figure 17: test set-up in the Beam test room	20

Figure 18: Pervious test set-up	22

Figure 19: test support for conveyer	22

List of tables

Table 1: Changes / designs or PNs	11

Table 2: Connectors to be use	13

Table 3: Test Equipment	23

# Applicable and reference documents

## Applicable documents

| Doc. Id. | Document Title | Document Nr, issue and date |
| --- | --- | --- |
|  | CERN Technical Specifications | BTV-RTCAM-TS002 revision 0.4 (22.10.2024) |
|  | MCAMv4 Design Description | MCSE-MCAMv4-DD-001 issue 1.5 84.12.2024) |
|  | MCAMv4 Equipment ICD | MCSE-MCAMv4-ICD-001 issue 1.2 (9.09.2024) |
|  | Acronyms List | MCSE-MCAMv4-LI-004 issue 1.1 (5.02.25) |

## Reference documents

| Doc. Id. | Document Title | Document Nr, issue and date |
| --- | --- | --- |
|  | List of commands | MCSE-BeamCAM-CERN-TN-0001 issue 1.1 (14.06.2021) |
|  | PoC previous test: SN01-SN02-CERN-CHARM-Test-Report : | MCSE-BeamCAM-RP-0014 : BeamCAM_CHARM _Results_2022_12_02.pdf |
|  | Generic Test Setup Installation Guide | MCSE-BCAM-PR-002 issue 1.1 (9.10.2025) |
|  | CHARM Technical Meeting - CHARM team | EDMS 2721313 |
|  | BOM updated | MCSE-BM4E-TN-001 1-0 issue 1.0 (13.11.2024) |
|  | MOSFET commercial plastic components survey and its SEE behavior | ALTER presentation version 3 |

## Norms, standards

| Doc. Id. | Document Title | Document Nr, issue and date |
| --- | --- | --- |
|  | None |  |

# Acronyms and definitions

## Acronyms

The definition of the acronyms used for the BeamCAM are given below. The other ones, related to the MCAMv4, are included in [AD04].

| DUT | Device Under Test |
| --- | --- |
| GPI | General Purpose Input |
| HK | Housekeeping (TM containing DUT status information) |
| IF, I/F | Interface |
| MAIT | Manufacturing, Assembly, Integration and Test |
| TC | Telecommand |
| TM | Telemetry (answer to TC, containing data, as for instance image pixel values) |

## Definitions

### Data sizes

8 bits = 8 b = 1 byte = 1 B;

1000 bits = 1 kilobit = 1 kb; 106 bits = 1 megabit = 1 Mb; 109 bits = 1 gigabit = 1 Gb

1000 bytes = 1 kilobyte = 1 kB; 106 bytes = 1 megabyte = 1 MB; 109 bytes = 1 gigabyte = 1 GB

1024 bits = 1 kibibit = 1 Kib; 10242 bits = 1 mebibit = 1 Mib; 10243 bits = 1 gigibit = 1 Gib

1024 bytes = 1 kibibyte = 1 KiB; 10242 bytes = 1 mebibyte = 1 MiB; 10243 bytes = 1 gigibyte = 1 GiB

### Data rates

1000 b/s = 1 kb/s; 1000 kb/s = 1 Mb/s;

Example: 1 Kib of data to be transferred at 1 kb/s takes 1.024 s.

# Scope

This document is the test plan for the evolution of the proof of concept of the BeamCAM to be tested at CHARM irradiation test chamber at CERN. It has been prepared by Micro-Cameras & Space Exploration SA (MCSE) in the frame of the development of a camera for CERN.

The new BeamCAM will be mainly based on the new generation of hardened, space grade micro-cameras currently developed by MCSE, which is called MCAMv4 (4th generation). Electronic and Electrical Elements on the BeamCAM prototype board are then the same as the ones embedded in the MCAMv4, with some CERN dedicated components, mainly in the power part.

The generic Test setup installation of the MCAMv4 / BeamCAM EBB is described in the dedicated document [RD03]. The present document contains the specific constraints to be followed for the Test Set-up at the CHARM facility in CERN premises.

# Irradiation test presentation

## Irradiation test objectives

The BeamCAM tests in CHARM facility is aimed to demonstrate that the overall HW design of the future BeamCAM is able to sustain the constraints which will be encountered in the different CERN scientific equipment, for example like in the current Large Hadron Collider (LHC) or the Future Circular Collider (FCC), a project currently in preparation.

Even if this test results could not directly be used by MCSE as a formal validation of the MCAMv4, it will also bring useful information in the frame of the MCAMv4 development and test.

To increase the test success chances, bring a certain redundancy and give more information, two BeamCAM prototypes (EBB, see a description in section 5) will be tested in parallel. To simplify the test set-up complexity and to allow an easier interpretation of the possible test failures, the complete test set-up will simply be completely duplicated : there will be two DUTs, two computers, two power supplies, two set of cables, etc. This document must be read in having this philosophy in mind.

## Irradiation characteristics

The test will be performed in the CERN High energy AcceleRator Mixed field (CHARM) facility, which is a unique testing facility that completes CERN's radiation testing installations. It is located in CERN’s East Area where high-flux radiation testing on smaller experimental equipment takes place.

During experiments, the CHARM area receives Proton Synchrotron (PS) beam spills of up to 5e11 protons, 400ms long, at 24 GeV/c. The beam strikes a metallic target, producing a mixed radiation field—neutrons, protons, photons, electrons, and muons—via spallation. Resulting dose rates range from 10 µGy/h to 1 Gy/h, with fluxes up to 1e9 particles/cm2/s, enabling electronics testing in radiation conditions representative of accelerators, space, and nuclear facilities. More information can be found in [ref: https://charm.web.cern.ch].

The irradiation conditions for the BeamCam testing requested over 3 weeks at CHARM are the following:

- Total Ionizing Dose TID [Gy]: >400

- High-energy Hadron fluence, 𝐸 > 20MeV: Θ𝐻𝑒𝐻 [cm−2]: 2e12

## Test main timeline

The CHARM test policy is to prepare each test within a dry run session, one week before starting the irradiation test. The CHARM provides facilities to test the equipment during the dry-run. The timeline is then as this :

Figure 1: test timeline

Specifically, for the BeamCAM EBB test the foreseen dates are :

- Dry run start : Wednesday the 15th of October 2025

- Irradiation test start : Wednesday the 22nd of October 2025

- The irradiation test should last for three weeks. It could be interrupted sooner if the DUT is not able to work properly before having be submitted to the total irradiation dose foreseen, or temporarily in case of problems which could be fixed.

## Test Sequences

Below agreed test sequences principles:

- Acquisition during the spill; ‘irradiated’ image acquired at start-up of Beam spill. That will be activated by a trigger provided by the CHARM equipment.

- Download and storage of the ‘irradiated’ image on the control computer. The image will be identified and time tagged.

- Acquisition after the Beam spill:  ‘non-irradiated’ image acquisition is triggered also by the available  trigger provided by the CHARM equipment. Neither the DUT FW or the test SW will perform any timing control, which will be given by the CHARM equipment.

- Download and storage of the non-irradiated image on the control computer. The image will be identified and time tagged.

- Loop to point 1.

- In parallel, check of the HK which will be send periodically from the DUT. If the HK flow is lost, a power cycle will be automatically generated after 10 minutes (TBC).

The images processing will be performed by an operator accessing remotely to the control computer, on a day-by-day basis.

## Test SW operations flow

A dedicated test SW is needed to operate and control the test. It will run on a dedicated test computer, remotely accessible. The SW main functions will be according to the test sequence described in the previous section. Here below a description of the SW test flow :

TBW

Figure 2: Test SW operations flow

## Keys monitoring parameters or data

Parameters and Data to monitor:

- Images acquisition when the beam is present

- Imagers Integrity

- Housekeeping data from Camera, As over-current protection counters on powerlines of sensor: Detection of Single event or latch-up / In case of OCP event, the FPGA power cycle the sensor.

- Voltage consummations

- Current consummations.

# Device Under Test (BeamCAM EBB) presentation and set-up

## Description

A development board (MCAMv4 Elegant Breadboard – EBB) was built by MCSE in the frame of the development of the 4th generation of the MCSE Micro-Cameras.

The BeamCAM, monitoring camera for CERN applications, will be developed in the future on the basis of the MCAMv4.

The EBB of the BeamCAM is a specific configuration of the MCAMv4 EBB. This configuration set all the needs for the BeamCAM design as components and functions. The generic test setup is described in [RD03] and sometimes in the present document when needed. All topics specific to the tests at CHARM are part of this document.

The components variations from the previous Proof of Concept board (see test report in [RD02]) are described in section 0.

Here below a picture of the BeamCAM EBB – actually the same as the MCAMv4 EBB.

Figure 3: BeamCAM EBB

For the BeamCAM, EBB, the components adaptations are the followed:

- Rad-tolerant Clock

- 12v to 5v module: CERN DC/DC.

## Block Diagrams

See in [RD03]

## Mechanical Information

See in [RD03]

## Proof of concept evolution / Components improvements

In the following table are listed the components changes between the prototype used during the first test in the CHARM facility (see [RD02]) and the new BeamCAM EBB. These changes are coming from the MCAMv4 design, which will be the basis for the future BeamCAM. The below information is also part of [RD05].

Table 1: Changes / designs or PNs

| Elements or functions | Descriptions / Remarks |
| --- | --- |
| Transistor PMOS | - PN change for automotive grade from same techno and same manufacturer. - Reduced package size. - Radiation impact: No gate rupture expected (less than 4V operations). Ref doc: see [RD06]. |
| FPGA PN | - For integration purpose the FPGA is taking with smaller size (484 BGA with pitches of 0.8mm instead of 1mm previously). - Power consideration to reduce SEE sensitivity: NO BANK powered at 3V3 -> Power topology design changes. - Same die techno between COTS and RadTol |
| Power On Reset (IC and passives elements) | Removed: done by FPGA. Present on EBB but as back-up. |
| Power topology Design | Major changes: - All components were changed for having COTS/Automotive grade and Radtol equivalents counterpart that are packages and sizes similar. To allow easy swap for screen quality for any camera’s types. - Following “new space plastic opportunity” - Remove one FEAST for size purpose -> 5V to 1V - DC/DC for FPGA power supplies to reduce power loss for conversion. - LDO changed - Sensors power topology optimization - DC/DC changed - Optional: added a Radtol Power E-FUSE IC from TI in power path (that is more for the space version of the camera) / Watch out: price by camera will be more expensive. |
| Trigger Input | Add a reverse diode on negative input line, to avoid damage when a wrong polarity connection is made. Generic 1N4148. (see JAN level for more screening or databases for RT) |
| Clock | POC PN BeamCAM Radtol Clock will be use. Another Radtol clock can be also used on the EBB as alternative. |

## Evolution of the power topology

Below Power Topology for Ethernet camera type Beamcam:

Figure 4: Power Path

Main changes:

- PN LDO converter changed

- Power converters for the sensors are reduces to 3x LDO instead of 4x.

- Power converters for the FPGA are changed for DC/DC’s types. NEW PN. Dual output ICs.

- No FPGA banks at 3V3 to reduce radiation sensitivity.

#### Latch-up protections

Some sensitives parts have their power supplies protect against latch-up events:

- All sensors supply

- SRAM supply

In Case of a detection of an over-current on the sensor or on the SRAM latch-up detection, the FPGA will process a restart and Reset of the sensor and SRAM.

## BeamCAM EBB Connectors identification

J4

Here below a description of the BeamCAM EBB connectors and interfaces:

Switch left on ON position

J01E

J9

J7

J8

J40

J39

J5

J10

J01

J02

Figure 5 Connectors IFs of EBBv4

## Interfaces Connections at Camera level

Table 2: Connectors to be use

| Designators | Type | Description | PN | Remarks / Notes |
| --- | --- | --- | --- | --- |
| J01 | MDM9-Femal | SpW IF V4 | NA | NOT USED FOR BEAMCAM |
| J02 | MDM15-Femals | PWR IF V4 | NA | NOT USED FOR BEAMCAM |
| J01E | RJ45 | Ethernet IFs | 615008137421 | Note Ground connection to chassis: 2x 100pf 2kv. |
| J4 | Barrel Connector | BeamCAM IFs | HR10-7R-6PA(73) | All Beamcam IF integrated Power and trigger.Adaptor cable to banana 4mm shall be mounted before closing top panel of the mechanic |
| J5 | SMA Connector | Isolated Trigger Ifs | SMA-RA-E | Second option for Trigger IF for Beamcam |
| J7 | Black 4mm Plug Banana IF | GND Interface | FCR7350B | Ground connection for Power the camera |
| J8 | Red 4mm Plug Banana IF | +5vdc Interface | FCR7350R | NOT USED FOR BEAMCAM |
| J9 | Yellow 4mm Plug Banana IF | Chassis Interface | FCR7350Y | NOT USED FOR BEAMCAM |
| J10 | Green 4mm Plug Banana IF | +12vdc Interface | FCR7350G | +12V connection for Power the camera |
| J39 | SMA Connector | FPGA Live probe B | SMA-RA-E | NOT USED FOR BEAMCAM |
| J40 | SMA Connector | FPGA Live probe A | SMA-RA-E | NOT USED FOR BEAMCAM |

## Electrical schematic of specific connectors

### Barrel HR10 (specific to BeamCAM EBB)

Figure 6: J4 Schematic

### Trigger SMA

Figure 7: J5 schematic

### Power 4MM Plug Bananas

Figure 8: 4mm plugs schematic

### Ethernet RJ45

Figure 9: ETH connector schematic

## BeamCAM EBB HW settings

The BeamCAM EBB will be tested according to the last HW definition coming from the MCAMv4 project (currently on-going), meaning all HW changes identified in the frame of this development project, as improvements or bugs fixing are implemented.

The tested HW is then in the last MCAMv4 Ethernet evolution and contains also the specific DC/DC and trigger IFs which are foreseen to be implemented in the future BeamCAM.

The BeamCAM EBB jumpers settings are according to the Ethernet configuration and specific power supply and trigger IFs. This is under MCSE responsibility.

The Power On switch (see in section 5.6) shall be always in ON position and shall never be switched OFF when the CERN DCDC is used.

## BeamCAM EBB FW settings

Each MCAMv4 / BeamCAM image sensor shall be individually calibrated. Four parameters must be set-up at the beginning of a video or images capture session. The image sensor dedicated values are measured during camera manufacturing and provided to the customer.

In the frame of the CHARM test the dedicated MCSE test SW will be in charge to set-up correctly these values.

Concerning the image acquisition, it is foreseen to set a standard acquisition time between 10 and 20ms. It is possible that the internal gain must be adapted during the test, to consider image degradation.

## BeamCAM EBB identification

As described in the test presentation, there will be two identical test set-up (meaning the list provided in section 7 must be prepared twice and the generic tests setup procedure followed two times (two reports to be provided).

The two DUT (BeamCAM EBB) will be :

- MCSE S/N T112 with a monochrome image sensor

- MCSE S/N T133 with a color image sensor

Note that, due to tight schedule in the FW development, the embedded FW of both DUT will be exactly the same. The HK and TM will indicate in a false manner that the image sensor is a b/w one, but that will not prevent to make images interpretation by a dedicated SW in a second step.

The optics will be provided by the CERN. The opto-mechanical I/F shall be C-Mount.

# Test premises

## CHARM Test Chamber CERN

The CHARM facility, located in CERN’s East Area, receives Proton Synchrotron (PS) spills of up to
5 × 1011 protons, 400 ms long, at 24 GeV/c. The beam strikes a metallic target, producing a mixed radiation field—neutrons, protons, photons, electrons, and muons—via spallation. Resulting dose rates range from 10 µGy/h to 1 Gy/h, with fluxes up to 1e9 particles/cm2/s, enabling electronics testing in radiation conditions representative of accelerators, space, and nuclear facilities.

Figure 10: Top: CHARM facility irradiation layout, with the BeamCAM setup in location ’10’. Bottom: Two BeamCAM systems, each with a dedicated black-and-white target to probe visibility.

## Preparation room 157/1-N26

The preparation room contains equipment (mainly the interfaces) which simulate the ones which are located in the irradiation test control room and the irradiation test room. It allows to prepare the irradiation test in an advance manner, in testing and validating the DUT and equipment provided by the customer (i.e. MCSE) in the same conditions than within the irradiation test control and test rooms.

Here below some pictures of the preparation room :

Figure 11: preparation room (general view)

| Figure 12: preparation room data acquisition patch panel | Figure 13: preparation room irradiation panel |
| --- | --- |

Some recommendations from the CHARM Team :

- Cables must be properly labelled and as easy and fast as possible to connect /disconnect (no terminal screw)

- Cables used to connect the patch panels should be the cables that will be used for the irradiation test (already installed or not in the cable holder)

IT network (GPN and TN) is available in preparation and control rooms, allowing remote control of computers.

The length of the cables between the patch panels have been checked by MCSE and are compatible with the foreseen test set-up.

## Beam area

### Conveyor

The DUTs will be installed in a conveyor which is used to move them under the beam location. There are several of them of different sizes. For the BeamCAM test the Overhead conveyer (position G16) will be used, as during the previous test. The mechanical test set-up will also be the same, all needed mechanical parts are available.

Figure 14: Overhead conveyor description

### Panels

Figure 15: Data acquisition panel

Figure 16: Irradiation panel

### Personnel preparation

The people who will work in the CHARM facilities shall respect the following points :

- External users must be preregistered in order to be able to access to their CERN web account (in which most of the formalities could be done)

- Access rights required for the different rooms and locations

- Safety trainings to be followed

### Installation in the Beam Area

The following operations are foreseen on Wednesday the 22nd of October :

- Access to beam area on Wednesday morning from 9h

- As soon as previous users removed their rack from control room 157-1-N28: installation of the new rack (not powered)

- Access in beam area on agreement with CHARM team

- Connection of the cables : patch panel + rack (1 team at a time)

- Powering and test of the setup (after confirmation of CHARM team)

- Transport of the rack to final test position

- End of access

- Beam expected usually in the afternoon

### Precautions to be taken during the Beam test

The following precautions must be followed during the whole Beam Test :

- No access possible into the beam area for the whole week

- No restriction to access in preparation and control rooms

- Personal dosimeter mandatory. Preferably take active dosimeters (DMC) from bld 55 (some pool dosimeters available at CHARM)

- Shoes and helmets mandatory

- No drink or food

- Close the doors as often as possible (air conditioning)

- Inform CHARM team in case your setup fails or if you reached your target dose. We could anticipate the target removal in order to reduce the radiation doses during the next weekly access

### End of the test

Activities to be performed at the end of the test :

- Access on Wednesday morning

- Disconnection patch panel + rack

- Protection of cables connectors for storage

- Material stored 1-2 weeks in internal buffer zone for cooldown

- User will be contacted few weeks later to retrieve their materials (TREC registration from RP buffer zone 157/R-O21)

- Material from the control and preparation rooms can be retrieved without any request.

- The facility can store the cables for later reuse

- Keep the control room tidy and clean

# Test Set-up

In a general manner the instructions given in the  Generic Test Setup Installation Guide [RD03] must be strictly followed. In this section 7 specific information for the test of the BeamCAM at CHARM are given.

## Diagram Test Set-up

The following test diagram is the same as in the Generic test set-up [RD03] section 5.4, but consider the two rooms in CHARM :

Figure 17: test set-up in the Beam control room

Figure 18: test set-up in the Beam test room

## Test Equipment

Generally, the test Equipment listed in [RD03] must be used (see section 6.1.3 List of equipment).

However, the following list of equipment must be discussed with the CHARM Team before the dry run in order to be fully compatible with the CHARM test equipment:

## Coax cable (item 11) solutions

### Solution 1: SMA and 4mm to BNC

Solution presented in sketch. TBD

### Solution 2: HR Barrel to 4mm then to BNC

Back-up solution: MCSE to provide adaptor cable from HR barrel interface to 4mm plug adaptor or Coax interface. TBD.

## As-run Set-up installation

The Generic Test Setup Installation Guide [RD03] contains an as-run test set-up installation in section 6.5, to be performed twice since there will be two identical test set-ups. Here below some notes about the specific installation in the CHARM facility :

- Step 4 : the remote access SW will be TeamViewer. Settings and access :

- TBD

- TBD

- Networks requests to be done ASAP (Dry run + control room)

- Step 5 : the dedicated list of equipment (number 11 in section 6.1.3 List of equipment in [RD03]) : the coax connection shall still be discussed between the CHARM team and MCSE, if possible before the dry test.

- Step14 (optional) : will not be done. It seems the final voltage in the accelerators facilities will be 10V. The DUT input DC/DC is able to accept as low as 5V.

- Step 15 (optional) : will not be done, the trigger is provided by the CHARM facility and the trigger I/F is already compatible.

## Previous Test-Set-up Installation

Figure 19: Pervious test set-up

## Test set-up support for conveyer

Test support shall be the same as the one from the previous one. A sight shall be mounted in front of both DUT optics. .

Figure 20: test support for conveyer

The CHARM team is responsible for the correct lightning of the two sights, which must be as uniform as possible.

## External (MCSE) equipment needed in Control Room

The following MCSE equipment is needed in Control Room :

Table 3: Test Equipment

| Type | Description | Remarks / Notes |
| --- | --- | --- |
| PSU Min: 12V 2 Amp. | Power supply for the camera | 4wires interfaces for the voltage level to guaranty the 12V at the DUT level is a plus! (Cables for power shall be the all the same length and gauge). Current protection / limitation at 1.8A. TBD Data to be monitor and log to a computer |
| Computer | Computer for safety | to log all data from PSU, Power cycling counter of the camera and trigger etc. Manage Failures and event by power cycling unit. |
|  | Computer to operate the camera | software to control the camera; Test sequences; etc. Log images and Housekeeping data |
|  | Computer storage | The computer shall have sufficient memory place to store all the data |

Note : the trigger signal will come from the CHARM equipment.

# FDIR during Test

Below descriptions of potential faults that may occurred during the irradiation test, with the potential causes and the recovery actions to apply if encountered

The software which controls the test will only check that the DUT is still functioning (see in section x). The below error’s conditions detection and recovery action need to regularly (ideal once a day) connect to the remote PC, download a few images from the saved ones and assess them. Power cycles will be done remotely as needed and new checks performed.

A decision to continue or stop the whole test session will be taken on a case-by-case basis, individually for each DUT.

| ID | Name | Description of fault | Potential Causes | Recovery |
| --- | --- | --- | --- | --- |
| 1 | Image corrupted | Only noises in images are sent back from the camera | Sensor connection with FPGA is having an issue. Single events on sensor-FPGA LVDs interfaces; FCC cable connection between sensor-FPGA is not properly connected | Power cycles the camera and resend a ping and wait for HKs and next images. If persistent, the sensor is lost. Continuing the test with the Image test acquisition. |
| 2 | Image harshly under / over exposed toward ref image | Too much exposition or too lower exposition that cannot be related to the degradation from sensor because of irradiation accumulation dose. | Sensor signal conditioning of pixels may have suffered from single event (Sensor ADC, PGA, etc.). Or corruption of sensor parameters | Send back to the camera, the sensor parameters. If the fault is persistent power cycle the camera. If it is still present the sensor is damaged |
| 3 | Image pixels are mixed / lines, and columns, Splits of images | Lines and/or columns are mixed in the images received. | Issue on the ethernet link or in the ethernet driver. FPGA lost a loop in Transmission. | Sent a HKs request or ping, and wait for next image, if persistent power cycle the camera. If it is still present the camera FPGA and or Ethernet driver have a degraded mode in ethernet sub-function. |
| 4 | Housekeeping Data sensor counter OCP increased | HKs are sending back OCP counters with increments | The power supplies of the sensor have been in limiter mode for the current: latch up events on sensor powers supplies. ------ FPGA is protected the sensor by a quick switch off of the sensor powers supplies. | The FPGA is power cycle the sensor, log the event, the image shall be corrupted. If events of counter happen a lot the sensor is sensible of the intensity of the radiation cocktail. If persistent power cycle the camera. If still present and, no image can be acquired the sensor is probably dead |
| 5 | No Ping feedback on ethernet | The camera doesn’t answer a ping from the control computer on ethernet | Loss of ethernet communication link on FPGA or Eth driver | Power cycles the camera and resend a ping and wait for HKs. If persistent, the camera communication link is lost. |
| 6 | Over current on PSU | Hight non-expected current happen on the PSU. Current drawn maybe to Hight | Latch-up events in camera somewhere not expected in power supply path. | Power cycles the camera and resend a ping and wait for HKs. If persistent, the camera is lost. |
| 7 | Voltage drop on PSU | Hight non-expected current happen on the PSU. Voltage drops with current limiter | Latch-up events in camera somewhere not expected in power supply path. | Power cycles the camera and resend a ping and wait for HKs. If persistent, the camera is lost. |

-------------------------------   End of document   -------------------------------