## Slide 1: Magnetic measurements at 1.9 K on the LMBRD01 (D2 cold mass with MBRD2)
L. Fiscarelli, M. Pentella, G. Dima, G. Deferne, R. Beltron
Colleagues from the sections MQA, LMF, CMI
B. Caiffi, A. Pampaloni S. Farinon, A. Bersani from INFN Genova
19.11.2025

## Slide 2: MBRD
  In industry at warm
  Rotating coil 750 mm – 11 positions
    Single apertures
    Two apertures in the sleeve
    Full magnet
  At CERN at warm
  Rotating coil 600 mm – 14 positions
    Magnet upon arrival
    Cold-mass
  At CERN at 1.9 K (with anticryostats)
  Rotating coil 1200 mm – 7 segments
    Machine cycles
    Stair-step
    Ramp-rate
  Stretched wire
    Nominal current (12330 A)
AP1 AP2

## Slide 3: Setup with rotating coils in anticryostats
New anticryostats
  Acknowledgement to CMI
New rotating-coil shafts
  66-mm diameter
  7x1200-mm segments
  Parallel measurement of both apertures

## Slide 4: Magnetic measurements at 1.9 K
TF
  Stretched wire
    At nominal (table)
  Rotating coils
    During ramp (plot)
Field quality
  Rotating coils
    At injection (table)
    At nominal (table)
    During ramp (plot)

## Slide 5: Transfer function
  Stretched wire
    At nominal
  Rotating coils
    Ramp from inj. to nom.
12330 A | B1 (Tm) | Field direction* (mrad)
AP1 | 35.7709 | -0.85
AP2 | 35.7713 | 0.33
*Angle wrt gravity when cryo-assembly was tested

## Slide 6: Field quality tables
AP1 |  |  |  |  |  | AP2 |  |  |  | 
 | Injection |  | Nominal |  |  |  | Injection |  | Nominal | 
n | bn | an | bn | an |  | n | bn | an | bn | an
2 | 5.85 | 0.40 | -0.08 | 1.83 |  | 2 | -3.71 | 1.58 | 2.11 | 3.89
3 | -14.19 | -1.86 | 3.47 | -3.15 |  | 3 | -14.70 | -2.04 | 2.86 | -2.73
4 | -1.81 | -0.08 | -0.28 | -0.40 |  | 4 | 1.89 | 0.00 | 0.23 | 0.61
5 | 3.47 | -0.61 | 2.83 | -1.16 |  | 5 | 2.43 | -0.91 | 2.31 | -1.39
6 | 0.66 | -0.02 | 0.13 | -0.28 |  | 6 | -0.33 | 0.07 | -0.25 | -0.21
7 | 1.44 | -0.65 | 1.96 | -0.54 |  | 7 | 1.36 | -0.46 | 1.91 | -0.41
8 | 0.83 | -0.16 | 0.99 | -0.19 |  | 8 | -1.11 | -0.22 | -1.09 | -0.26
9 | 0.94 | -0.20 | 1.29 | -0.20 |  | 9 | 1.03 | -0.06 | 1.25 | -0.17
10 | 0.46 | 0.56 | 0.44 | 0.44 |  | 10 | -0.30 | 0.38 | -0.28 | 0.38
11 | -2.08 | -0.52 | -2.14 | -0.57 |  | 11 | -2.00 | 0.31 | -1.98 | 0.33
12 | 2.02 | 0.55 | 2.04 | 0.54 |  | 12 | -1.69 | 0.56 | -1.77 | 0.53
13 | -1.87 | -0.37 | -2.05 | -0.37 |  | 13 | -1.74 | 0.24 | -1.88 | 0.27
14 | 0.95 | 0.19 | 1.10 | 0.24 |  | 14 | -0.87 | 0.18 | -0.93 | 0.24
15 | -0.52 | 0.30 | -0.60 | 0.31 |  | 15 | -0.49 | 0.07 | -0.67 | 0.11
  Reference radius 35 mm

## Slide 7: Field quality vs acceptance criteria
Acceptance criteria MBRD https://edms.cern.ch/document/2051868/1.39
  Reference radius 35 mm
 |  | Acceptance criteria |  |  | AP1 |  | AP2 |  |  | AP1 |  | AP2 | 
n |  | ¦bn¦ | ¦an¦ |  | bn | an | bn | an |  | ¦∆bn¦ | ¦∆an¦ | ¦∆bn¦ | ¦∆an¦
2 |  | 4.00 | 2.72 |  | -0.08 | 1.83 | 2.11 | 3.89 |  |  |  |  | 1.17
3 |  | 6.67 | 1.13 |  | 3.47 | -3.15 | 2.86 | -2.73 |  |  | 2.02 |  | 1.61
4 |  | 2.40 | 1.78 |  | -0.28 | -0.40 | 0.23 | 0.61 |  |  |  |  | 
5 |  | 4.00 | 0.61 |  | 2.83 | -1.16 | 2.31 | -1.39 |  |  | 0.55 |  | 0.78
6 |  | 0.24 | 0.70 |  | 0.13 | -0.28 | -0.25 | -0.21 |  |  |  | 0.01 | 
7 |  | 0.66 | 0.23 |  | 1.96 | -0.54 | 1.91 | -0.41 |  | 1.30 | 0.32 | 1.25 | 0.18
8 |  | 0.11 | 0.24 |  | 0.99 | -0.19 | -1.09 | -0.26 |  | 0.88 |  | 0.98 | 0.01
9 |  | 0.26 | 0.08 |  | 1.29 | -0.20 | 1.25 | -0.17 |  | 1.03 | 0.12 | 0.99 | 0.09
10 |  | 0.03 | 0.10 |  | 0.44 | 0.44 | -0.28 | 0.38 |  | 0.41 | 0.34 | 0.25 | 0.28
11 |  | 0.08 | 0.03 |  | -2.14 | -0.57 | -1.98 | 0.33 |  | 2.07 | 0.54 | 1.91 | 0.31

## Slide 8: Multipoles vs current
  Reference radius 35 mm

## Slide 9: MCBRD’s
 | 393 A | B1 (Tm) | A1 (Tm)
MCBRDP4 | AP2_B_H | 5.0185 | -0.0134
 | AP1_B_V | 0.0286 | 5.0306
MCBRDP3b | AP1_A_H | 5.0202 | -0.0297
 | AP2_A_V | -0.0182 | -5.0032
V
H
  Stretched wire
    At nominal 393 A

## Slide 10: Summary
Complete magnetic characterization available for the first series magnet
MBRD
  TF is larger (+2%) than the nominal 35.00 Tm and almost identical on the two apertures (current can be set lower)
  Correction applied on b3 and b5 is effective
  Field quality marginally outside the acceptance criteria
MCBRD’s
  Integral field at nominal with stretched wire for calibration of TF