Minutes of the 193rd EATM Meeting held on 18 November 2025

Minutes and slides available at: https://indico.cern.ch/event/1612297/

Present: A. Devienne (HSE-RP), J. Bernhard (BE-EA), N. Kahn (BE-EA), B. Maksiak (EP-SME), P. Martinengo (EP-DT), L. Nevay (BE-EA, chair), R. Piandiani (EP-UFT), B. M. Veit (EP-UFT).

Online: N. Charitonidis (BE-EA), H. Danielsson (EP-DT), A. Ebn Rahmoun (BE-EA), R. Folch (BE-EA), F. Gautheron (BE-EA), X. Genillon (SY-EPC), S. Girod (BE-EA), M. Jaekel (EP-DT), D. Jaillet (BE-EA), I. Krautsztrung (BE-EA), D. Lazic (EP-UCM), J. Lehtinen (EN-CV), M. Mentink (EP-ADO), I. Ortega Ruiz (SY-BI), X. Palle Lopez (BE-EA), B. Rae (BE-EA), P. Schwarz (TE-MSC), I. Shahzadi (HSE-OHS), M. Van Dijk (BE-EA).

The minutes from the 191st EATM were approved and the minutes for the 192nd EATM will circulate shortly.

Actions for 188th EATM

- S. Girod to confirm the detailed plan offline for GIF++ extension work at the end of YETS 25/26. → Meeting to take place. New safety-related works for exhaust line to be added.

- G. Romagnoli to clarify general XTAX FS vs P4. → Done.

Action items (D. Banerjee) – Slides

The procedure for the galleries access document will circulate shortly. → J. Bernhard.

Key information from Different Meetings

SBA Highlights (D. Banerjee) – Slides

The treatment of electrical non-conformities in the North Area is 90% complete. For the various electrical incidents, a meeting was held, and solutions were found that will be implemented in YETS 25/26. An extra fence is required for the NA61 hydrogen buffer zone outside EHN1, which was installed. Additional electrical sockets and support for gas bottles were installed in the EHN1 grey room. The flashing lights in EHN1 near the gallery access points are caused by the failure of the ventilation fans. Parts are ordered and will be exchanged during the YETS 25/26. The shipment of the WCTE tank has been scheduled for 19 November.

ECRs (N. Kahn) – Slides

The status of the ECRs for information and future approval was summarised and the full list can be found in the slides. Eight documents were presented for approval:

- Production/Installation of a New Access Gangway to TELMAX/ASACUSA Platforms. Approved.

- NA-CONS Fire Resistant Partitions 3 Partitions #09, #10, #12, #13, #14, #17 and #33. Approved.

- PPE118/128 Shielding Modification for Better Crane Access to TT82. Approved.

Planning and Important Dates (B. Rae)

The list of MD plans and access are available online: https://be-dep-ea.web.cern.ch/content/md-planning-north-area. No change to the YETS planning and for the start-up dates for next year.

Strategy Update for North Area Profile Monitors (I. Ortega Ruiz) – Slides

I. Ortega presented the consolidation strategy for beam profile monitors in the North Area. The plan foresees the progressive replacement of existing devices (XWCAs, XWCMs, DWCs and FISCs) with XBPF monitors under NA-CONS, including the integration of the XSCI directly into the XBPF tank to improve beam quality and reduce long-term maintenance. A limited number of XBPFs (16 planes) are being produced already in Phase 1 for the primary beamlines, while the full deployment will follow in Phase 2. In parallel, an R&D programme is ongoing for a radiation-hard version of the monitor, mechanically compatible and motorised for use in high-radiation locations. Installation planning was performed in 2022 and coordinated with the beamline physicists, with an updated EDMS document currently under approval. The replacement strategy currently foresees 16 XBPF planes in Phase 1 and 66 in Phase 2.

LS3 Constraints and Planning (B. Rae) – Slides

B. Rae presented the current draft LS3 master schedule and constraints affecting activities in the Experimental Areas. The planning remains under development and the restart phases (HWC, BC, etc.) will be further defined at a later stage. In the East Area, CLOUD will be impacted by EN-CV dismantling works for the new ventilation system, currently foreseen for October 2026 but potentially starting already in September if asbestos is detected during checks planned in YETS 25/26. In the North Area, availability will be affected by several infrastructure interventions at different times: for MADMAX and RADES, EN-EL work on the BA81 false floor is scheduled in October - November 2026 and March - July 2027, and EN-CV work will not be possible during CT2 activities planned from September 2026 to October 2027, while TE-CRG availability will depend on coordination with other services; the CMS GEM Test is subject to the same BA81 and CT2 constraints; for the Neutrino Platform, EN-CV operation is possible with a local chiller, whereas EN-EL and TE-CRG implications remain to be clarified; and for AMBER, EN-EL work under the BA82 false floor in December 2026 - January 2027 and May - October 2027, together with the CT2 schedule, similarly constrains EN-CV activities, while TE-CRG planning will depend on overall resource availability. No long EN-EL interruptions are foreseen for GIF++, and CV operation is assured using a local chiller. All dates shared remain preliminary and subject to validation.

Preparation for North Area Ion Run (S. Girod) – Slides

S. Girod presented the planned configuration of H2, H4 and H8 for the 2025 ion run, detailing the required access restrictions, shielding installations, vacuum changes, interlock updates (PPE172 for H2A, PPE144 for H4A) and bend limitations. The activities will be executed following the injector schedule, with further instrumentation removals in H8 to be carried out during the run as requested. The configuration steps are tracked via the standard CERN checklist tool.

News From Experiments – Mid-Year Update

IRRAD / CHARM: F. Ravotti (offline) – IRRAD completed 53 experiments with over 1300 samples, and CHARM carried out 42 system-level tests plus 120 component tests. The high-intensity, stable beam since 2024 enabled smooth operation throughout 2025, including accommodating additional late requests related to Phase II detector electronics. Remaining requests are planned for early 2026, and the call for CHARM experiments in 2026 is being reopened. The proton run ends tonight; a two-week Pb run for HEARTS will follow.

AMBER: M. Veit – All detectors for PRM have been installed and the final safety clearance was given. The TPC was filled with hydrogen. Access for EHN2 and adjacent workshop is now part of ATEX zoning, i.e., no works allowed.

NA62: R. Piandani – NA62 reports frequent issues with missing timing, leading to unwanted p0survey triggers. In addition, the interlock is occasionally activated because the vacuum valve status becomes undefined for short periods.

NA64e: V. Poliakov – No update during this meeting.

NA61: B. Maksiak – NA61 has been running stably with the liquid-hydrogen target since 2 November and will continue until the switch to ions with a Pb target on Monday morning. The main issue remains the significantly reduced efficiency of the delay wire chambers compared to the summer run, with improvements from 40% to 80-95% achieved through HV and threshold adjustments, chamber positioning and a modified gas supply. One chamber (DWC3) continues to show particularly poor performance due to a still-unknown source of excess noise on one channel.

CLOUD: A. Onnela (offline) – CLOUD18 is progressing well and will conclude on Monday morning, 1 December, followed by a week of calibrations and instrument removal/shipment, which will make the entrance zone next to CLOUD busy; an IMPACT will be issued for these activities. An unintended evacuation of the East Hall occurred on 15 October when an alarm button was accidentally pressed during equipment handling; the fire brigade responded, and the alarm was promptly cleared.

GIF++:  P. Martinengo – The gas exhaust line needs to be brought up to ATEX standards for operation with flammable gases.

HiRadMat: N. Charitonidis – No news.

AoB

- M. Jäkel: Draft for East Area scheduling close to completion.