## Slide 1: Beam Diagnostics Lecture 2
Measuring Complex Accelerator Parameters
Uli Raich
CERN AB-BI

## Slide 2: Contents of lecture 2
Some examples of measurements done with the instruments explained during the last lecture
  Spectroscopy
  Trajectory and Orbit measurements
  Tune measurements
    Traditional method
    BBQ method
  Multi-turn extraction
  Transverse and longitudinal emittance measurements
  Longitudinal phase space tomography

## Slide 3: Faraday Cup application Testing the decelerating RFQ
Antiproton decelerator
Accelerate protons to 24 GeV and eject them onto a target
Produce antiprotons at 2 GeV
Collect the antiprotons and cool them
Decelerate them and cool them
Output energy: 100 MeV
In order to get even lower energies:
Pass them through a moderator
  High losses
  Large energy distribution
=> Build a decelerating RFQ

## Slide 4: Waiting for Godot
RFQ
Proton beam
FC
FC
Spectrometer
magnet
Testing the
decelerating RFQ

## Slide 5: Setup for charge state measurement
The spectrometer magnet is swept and the current passing the slit is measured
beam
Ion
source
Faraday
Cup
Spectrometer
Magnet

## Slide 6: Measuring charge state distribution
Faraday Cup
Slit
Spectrometer
magnets

## Slide 7: Charge state distribution measured with a Faraday Cup on a heavy ion source

## Slide 8: Trajectory and Orbit measurements
Definitions:
Trajectory: The mean positions of the beam during 1 turn
Orbit: The mean positions over many turns for each of the BPMs
The trajectories must be controlled at injection, ejection, transition
Closed orbits may change during acceleration or RF “gymnastics”

## Slide 9: Trajectory Measurement at LHC
Knowing the optics one can deduce the orbit correction from the measurement

## Slide 10: The PUs

## Slide 11: The PS, a universal machine
All beams pass through the PS
Different particle types
Different beam characteristics
Concept of a super cycle

## Slide 12: The super cycle

## Slide 13: Position Measurements
Red: The sum signal
Green: The difference signal
Procedure:
Produce integration gates and
Baseline signals
Baseline correct both signals
Integrate sum and difference signals
and store results in memory
Take external timing events into account e.g. harmonic number change, γ-transition etc.

## Slide 14: Trajectory readout electronics

## Slide 15: Trajectory measurements in circular machines
Needs integration gate
Can be rather tricky
Distance between bunches
changes with acceleration
Number of bunches
may change
Raw data from pick-ups
double batch injection

## Slide 16: Beams in the PS

## Slide 17: Baseline restoration
Low pass filter the signal to get an estimate of the base line
Add this to the original signal

## Slide 18: RF Gymnastics

## Slide 19: Changing bunch frequency
Bunch splitting or recombination
One RF frequency is gradually decreased while the other one is increased
Batch compression
For all these cases the gate generator must be synchronized

## Slide 20: Batch compression

## Slide 21: Tune measurements
When the beam is displaced (e.g. at injection or with a deliberate kick, it starts to oscillate around its nominal orbit (betatron oscillations)
Measure the trajectory
Fit a sine curve to it
Follow it during one revolution
kicker

## Slide 22: The Sensors
The kicker
Shoebox pick-up
with linear cut

## Slide 23: Tune measurements with a single PU
Design by P. Forck

## Slide 24: Kicker + 1 pick-up
Measures only non-integral part of Q
Measure a beam position at each revolution
Fourier transform of pick-up signal

## Slide 25: Windowing
The Discrete Fourier assumes one cycle of a repetitive signal.
Blackman-Harris Window is used
Each sample is multiplied with a coefficient
Coefficients are pre-calculated and stored in a table

## Slide 26: Periodic extension of the signal and Windowing
discontinuity

## Slide 27: Peak search algorithm
Power value is bigger than its predecessor
Power value is bigger than its successor
Power value is biggest in the whole spectrum
The power value is at least 3 times bigger than the arithmetic mean of all power bins.

## Slide 28: Q interpolation
Betatron signal is not a pure
Harmonic but includes rev. freq
Harmonics, noise …
The windowing process is not
Perfect
Coherent betatron signal is
Damped in the time domain

## Slide 29: Q-Measurement Results

## Slide 30: Direct Diode Detection Base Band Q measurement
Diode Detectors convert spikes to saw-tooth waveform
Signal is connected to differential amplifier to cut out DC level
Filter eliminates most of the revolution frequency content
Output amplifier brings the signal level to amplitudes suitable for long distance transmission
Curtesy M. Gasior

## Slide 31: BBQ Results from CERN SPS
Results from Sampling
After Fourier Transform

## Slide 32: Tune feedback at the LHC

## Slide 33: Multi-Turn Extraction
Losses on the
septum blade

## Slide 34: Islands in transverse phase space
Create stable island
in phase space through excitation with hexapoles and
octopoles and capture
the beam in them
Can this be measured?
Poor man‘s phase space
meter using 2 BPMs at
90° phase advance

## Slide 35: Profile of Islands

## Slide 36: Stable Islands in Phase Space

## Slide 37: Projection in Phase Space

## Slide 38: Emittance measurements
A beam is made of many many particles,
each one of these particles is moving with
a given velocity. Most of the velocity
vector of a single particle is parallel to the
direction of the beam as a whole (s).
There is however a smaller component of
the particles velocity which is
perpendicular to it (x or y).
Design by E. Bravin

## Slide 39: Emittance measurements
If for each beam particle we plot its position and its transverse angle we get a particle distribution who’s boundary is an usually ellipse.
The projection onto the x axis is the beam size
x’
x
Beam size

## Slide 40: The slit method
If we place a slit into the beam we cut out a small vertical slice of phase space
Converting the angles into position through a drift space allows to reconstruct the angular distribution at the position defined by the slit
x’
x
slit

## Slide 41: Transforming angular distribution to profile
When moving through a drift space the angles don’t change (horizontal move in phase space)
When moving through a quadrupole the position does not change but the angle does (vertical move in phase space)
x’
x
slit
x’
x
slit
x’
x
slit
Influence of a drift space
Influence of a quadrupole

## Slide 42: The Slit Method

## Slide 43: Phase Space Scanner

## Slide 44: Emittance plot Solenoid
0
The solenoid splits
the trajectories according to
particle type.
The source produces
protons
H0
H2+
H3+

## Slide 45: Moving slit emittance measurement
Position resolution given by slit size and displacement
Angle resolution depends on resolution of profile measurement device and drift distance
High position resolution → many slit positions → slow
Shot to shot differences result in measurement errors

## Slide 46: Transformation in Phase Space
X’
X
X’
X
X’
X
X’
X
X’
X
Quadrupole
Quadrupole
SEMGrid
Kicker
Kicker
at slit
first drift space
first quadrupole
second drift space
Second quadrupol

## Slide 47: Result of single pulse emittance measurement

## Slide 48: Transverse Emittance measurement
E
φ
E
E
φ
Kicker
Buncher
Kicker
SEMGrid
Transformer
Spectrometer magnet
E
φ
at slit
first drift space
buncher
second drift space
φ
Spectrometer produces image of slit on second slit
second slit selects energy slice
first kicker sweep phase space over all energies
buncher rotates energy slice in phase space
at second spectrometer the phase distribution is transformed into an energy distribution analyzed by the second spectrometer
second kicker corrects for first kick

## Slide 49: Single Shot Emittance Measurement
Advantage:
  Full scan takes 20 μs
  Shot by shot comparison possible
Disadvantage:
  Very costly
  Needs dedicated measurement line
  Needs a fast sampling ADC + memory for each wire
Cheaper alternative:
  Multi-slit or pepperpot measurement

## Slide 50: Multi-slit measurement
Needs high resolution profile detector
Must make sure that profiles dont overlap
beam
Scintillator + TV + frame grabber
often used as profile detector
Very old idea, was used with photographic plates

## Slide 51: Pepperpot
Uses small holes instead of slits
Measures horizontal and vertical emittance in a single shot
Photo P. Forck

## Slide 52: Computed Tomography (CT)
Principle of Tomography:
Take many 2-dimensional Images at different angles
Reconstruct a 3-dimensional picture using mathematical techniques (Algebraic Reconstruction Technique, ART)

## Slide 53: The reconstruction
Produce many projections of the object to be reconstructed
Back project and overlay the “projection rays”
Project the back-projected object and calculate the difference
Iteratively back-project the differences to re-construct the original object

## Slide 54: Some CT resuluts

## Slide 55: Computed Tomography and Accelerators
RF voltage
Restoring force for non-synchronous particle
Longitudinal phase space
Projection onto Φ axis corresponds to bunch profile

## Slide 56: Reconstructed Longitudinal Phase Space

## Slide 57: Bunch Splitting