## Slide 1
Contents
1 – Introduction History and current status of the PS B-train system
2 – Magnetic measurements of reference U101 Summary of main results 2010-2011 First results of 2012 campaign on field harmonics and eddy currents
3 – The new B-train system Concept and functional specifications Hardware and software components
4 – Summary
Magnetic Measurements with the PS B-train Recent results and planned developments
M Buzio, D Giloteaux, D Oberson, L Walckiers

## Slide 2
Acknowledgment
Thanks to A. Beaumont, F. Caspers, G. Golluccio, O. Dunkel, L. Gaborit, P. Galbraith, R. Beltron Mercadillo for their technical and scientific contribution
Thanks to F. Boattini, S. Gilardoni, O. Michels, X. Genillon and the PS operation team for the organization and execution of recent tests
Thanks to H. Damerau, J. L. Gomez Costa, Q. King, P, Odier, J. Serrano, L. Soby for their collaboration towards the definition of the new serial distribution link

## Slide 3
Introduction

## Slide 4
local measurement 
What is a B-train ?
I(t)
real-time dipole field measurement for accelerator operation
B(t)
qualitative feedback to operation
beam diagnostics (beam current DCCT)
power supply control (B or dB/dt)
RF control (mandatory)
sometimes B(t) ≠ f(I,dI/dt,I(≤t))
Hysteresis (“Pulse to Pulse Modulation”)
Temperature
Eddy currents
Ageing of iron yoke
I(t) measurement accuracy
Timing jitter
reference dipole (in the ring or external + series excitation)

## Slide 5
B-train working principle – dynamic marker
“PS pulses” (24 V, 1 ms)
0.02, 0.1, 1 or 10 G/pulse
digital tech from the ‘60s !
Bflat top
time
B(t)
Bflat bottom
t1
marker 1
coil measurement Ḃ=-Vcoil/Acoil
UP-train
DOWN-train
preset B(t1)
preset
marker trigger pulse
strategy currently implemented at CERN in five machines: PSB, PS, SPS, LEIR, AD

## Slide 6
Timeline of PS B-train systems
3 systems in the lifetime of PS:
Mark I (1960-1980) - included in the original magnet design - only one side of U101 (block 4) equipped with peaking strips @ 50 G + flux coils - 3 channels @ 0.1, 1 and 10 G;
Mark II (1980-2000) - motivation: adapt to multi-batch filling + new F8L - different F/D fields  sensors in both halves - new peaking strips designed for up to 400 G + new flux coils - extra logic to interpolate B(t) between triggers
Mark III (2000 ‒ ) - motivation: improved precision and reliability - peaking strips @ 50 G (D only) + flux coils (D + F) - peaking strip F removed from B-train, after a few tests, in order to be compatible with the newly installed PSB/AD systems
Mark I
R. Gouiran, 1979
Mark II & III

## Slide 7
CERN PS B-train: current configuration
1D
2D
3D
4D
5D
1F
2F
3F
4F
5F
peaking strip (in use)
flux coils
reference magnet U101, bldg. 355
peaking strip (spare)
A0(1-k)
A0(1+k)
VD
VF
B-train = weighted average of the local field in F/D halves
In-built assumptions: BD(t0)=BF(t0)=4.98 mT, k=0.091†
Only one field marker in use (the second one is foreseen only for diagnostics)
When these assumptions are not satisfied  B-train error  beam may become unstable
Integrator is reset at the end of every magnet cycle (synch to control system timing)
Integrator internal auto-calibration every 15 cycles (not synch to timing ….)
† k takes into account an oscillating dipole component generated by the PFW in 3-current mode. See: R. Gouiran, Enroulements polaires de l'aimant du CPS, Note Interne PS/SM 76-1

## Slide 8
Reference magnet U101
Longitudinal field distribution varies with IF8L, IPFW, saturation
Calibration vs. beam depends on sensor position
Sensors may not work properly in end blocks (inhomogeneous/3D field)
Source: A. Asklöv, Magnetic measurement on the CERN proton synchrotron, Master’s thesis, EDMS 609417, 2005

## Slide 9
Existing field marker: peaking strips
Developed at CERN in 1956 specifically for combined-function PS magnets.
Based on a pre-stressed bi-stable magnetic needle: magnetization flips over at a preset level and generates a pulse detected by a pick-up coil
Two coils powered in series for bias and screening field, pulsed to avoid overheating
Main constraints: bias coil heating at high field (> 5 mT); does not work at too high or too low dB/dt
Experience shows that this sensor is very stable (drift <50 T in 20 years),
Very few spares available, all different from each other (making new ones is difficult)
pick-up coil
permalloy needle
needle
bias and screening coils
pick-up voltage
power

## Slide 10
Peaking strips in side U101

## Slide 11
Diagnostic off-line “spy” system (2010)
Based on National Instruments hardware: PXIe chassis, RT controller, 2 6366 DAQ cards (82 MHz parallel 16-bit channels + 3210MHz digital I/O channels)
Acquisition of all sensors and B-train signals
Introduced in 2010 to track the magnetic state in F/D halves
Used to check the consistency of sensor and B-train signals, timing and currents
Testbed for the design of new sensors and acquisition algorithms
NB: high sample rate in the 200 kHz range is necessary for:
accurate integration of flux coil output
correlation of current/field to few 0.01 G

## Slide 12
Harmonic coil array (2012)
New 5-coil, 102 mm fluxmeter for the dynamic measurement of normal harmonics B1  B5
First test carried out 16/17 feb
Support can be quickly repositioned in all free F or D blocks (optimal position to be defined)
Second symmetrically placed unit foreseen
Notes: (Re)establish clear responsibility boundaries with BE/CO: - local control racks - synthetic B-trains - closed-loop field control regulation (e.g. ISOLDE)
“piquet” service: so far a best-effort approach with three people on-call has been sufficient, what about the long-term ?
standardization of synthetic trains with FIDEL-like algorithms: magnetic measurements+ beam measurements + numerical simulations  mathematical model of magnet behaviour

## Slide 13
Dynamic measurement of harmonics
Open issues:
integrator drift is very sensitive to calibration errors  only B1 and B2 OK for the moment (more accurate calibration in preparation)
truncation error due to neglected n6 harmonics to be evaluated
major improvements planned: - analog bucking of main components - circular array of tangential coils (well conditioned calibration + skew harmonics + possibility of rotation to measure accurately remanent harmonic components )
5 measured integrated fluxes
wanted harmonics (normal only)
weight matrix (NB: problem intrinsically ill-conditioned due to its polynomial-fitting nature)
Notes: (Re)establish clear responsibility boundaries with BE/CO: - local control racks - synthetic B-trains - closed-loop field control regulation (e.g. ISOLDE)
“piquet” service: so far a best-effort approach with three people on-call has been sufficient, what about the long-term ?
standardization of synthetic trains with FIDEL-like algorithms: magnetic measurements+ beam measurements + numerical simulations  mathematical model of magnet behaviour

## Slide 14
Test results highlights

## Slide 15
Dipole transfer function
Pre-cycling at high field + higher minimum current  better field stability
large fluctuations due to history-dependent residual field
20 A  4.98 mT
240 A  60 mT
going into saturation erases the previous magnetic history

## Slide 16
PS beam stability test – April 2010
Hysteresis curves BF(I) and BF(I) – Difference w.r.t. straight line fit
MRP 4 mm
MRP 14 mm
BF remanent shift  -10 G
IF8L pre-cycle
“good” cycle
“bad” cycle
grossly asymmetric powering imbalances BF and BD current B train system is blind to initial BF Bavg measurement error  beam MRP error

## Slide 17
Field imbalance
dipole and quadrupole component are always proportional  tune trim circuit operation (IF8L) at low field causes field imbalance in the two magnet halves
Independent field marking in the two halves is necessary for correct B(t) integration
negative quadrupole increment in the defocusing half the field changes along the downwards limit cycle
positive quadrupole increment in the focusing half the field changes towards the upwards limit cycle

## Slide 18
PS beam stability test @ injection – June 2011
Imain
IF8L
systematic correlation between initial field imbalance (B-train error) and large RP error

## Slide 19
PS beam stability test @ injection – June 2011
  CNGS
  CNGS
  CNGS
  2nd CNGS cycle unstable
  2nd CNGS cycle OK
  2nd CNGS cycle unstable
specific powering cycle sequences lead to reproducible MRP errors now we can observe correlation with reproducible magnetic errors

## Slide 20
Possible mechanical side-effects
Laser tracker measurements done on U17 to test stability of NMR marker support in 2008
Courtesy A. Beaumont
Notes: (Re)establish clear responsibility boundaries with BE/CO: - local control racks - synthetic B-trains - closed-loop field control regulation (e.g. ISOLDE)
“piquet” service: so far a best-effort approach with three people on-call has been sufficient, what about the long-term ?
standardization of synthetic trains with FIDEL-like algorithms: magnetic measurements+ beam measurements + numerical simulations  mathematical model of magnet behaviour

## Slide 21
Possible mechanical side-effects
Vertical displacement of bottom pole = 0.16 mm @ 5400 A (gap closes due due magnetic forces)
All B-train sensors fixed to bottom pole
Top/bottom asymmetry  possible vertical displacement of sensors w.r.t. to ideal axis
Feed-down  skew spurious harmonics e.g. 0.1 mm  5G (systematic effect as a function of main current)
Dz
A1
Notes: (Re)establish clear responsibility boundaries with BE/CO: - local control racks - synthetic B-trains - closed-loop field control regulation (e.g. ISOLDE)
“piquet” service: so far a best-effort approach with three people on-call has been sufficient, what about the long-term ?
standardization of synthetic trains with FIDEL-like algorithms: magnetic measurements+ beam measurements + numerical simulations  mathematical model of magnet behaviour

## Slide 22
Bdot signal noise issues (2011)
Abnormally high noise on Bdot reported by Ops in Feb 2011 (concern for RF).
B(t) noise at f > 1 kHz found to be up to 40 dB higher in 2011 than 2010
Observations consistent with actual (small) B fluctuations (see EDMS 1168725)
Noise levels dropped back well below 2010 levels in 2012
POPS with filters 2012
Generatrice 2010
POPS 2011
Field spectrum inferred from flux coil measurements in U101:
(spurious resonances > 104 Hz possible)
Spectrum of magnetic field oscillations
Notes: (Re)establish clear responsibility boundaries with BE/CO: - local control racks - synthetic B-trains - closed-loop field control regulation (e.g. ISOLDE)
“piquet” service: so far a best-effort approach with three people on-call has been sufficient, what about the long-term ?
standardization of synthetic trains with FIDEL-like algorithms: magnetic measurements+ beam measurements + numerical simulations  mathematical model of magnet behaviour

## Slide 23
Dynamic effects (eddy currents)

## Slide 24
B (obtained from flux coil integration)
I
I
t
t
t
with saturation
without saturation (ideal)
Scale so that the curves coincide at the end of the flat-top
apparent field advance due to the scaling procedure
No eddy currents  I* I
ts
ts
Eddy current measurement method
equivalent current:
Eddy currents show up as a transient with exponential decay at end of ramp-up (NB: direction depends on current overshoot!)

relative field error
Ideal model:
Transient amplitude A  dB/dt
Decay time  = constant

A

## Slide 25
Current-controlled plateau @ 3000 A, 9000 A/s
I(t)
I(t) (zoom-in)
I*2(t) (quadrupole)
I*1(t) (dipole)
t (s)
I (A)
relative field error
3.510-3 (quadrupole)
1.010-3 (dipole)
  50 ms (quadrupole)
  25 ms (dipole)
At low field (linear range) eddy current effects are relatively small and fast
Quadrupole effects significatively larger than on dipole
Notes: (Re)establish clear responsibility boundaries with BE/CO: - local control racks - synthetic B-trains - closed-loop field control regulation (e.g. ISOLDE)
“piquet” service: so far a best-effort approach with three people on-call has been sufficient, what about the long-term ?
standardization of synthetic trains with FIDEL-like algorithms: magnetic measurements+ beam measurements + numerical simulations  mathematical model of magnet behaviour

## Slide 26
Current-controlled plateau @ 5380 A, 8150 A/s
B1(t)
B2(t)
Respect to intermediate field, the time constant and effect amplitude on the quadrupole are ~2 x
No appreciable change on the dipole
Dimensional considerations based on magnetic field diffusion equations  time constant at high field expected to drop by 20% due to saturation (low ) only  dominant effect appears to be increase of magnetic circuit length due to flux leakage
I(t) (zoom-in)
I*2(t) (quadrupole)
I*1(t) (dipole)
quadrupole relative field error  510-3
  100 ms
Notes: (Re)establish clear responsibility boundaries with BE/CO: - local control racks - synthetic B-trains - closed-loop field control regulation (e.g. ISOLDE)
“piquet” service: so far a best-effort approach with three people on-call has been sufficient, what about the long-term ?
standardization of synthetic trains with FIDEL-like algorithms: magnetic measurements+ beam measurements + numerical simulations  mathematical model of magnet behaviour

## Slide 27
Current-controlled plateau @ 5380 A, 3300 A/s
Results at high field with reduced ramp rate show smaller relative field error  behaviour consistent with eddy currents
Equivalent current plot indicates much higher saturation effects on quadrupole than on dipole
I(t) (zoom-in)
I*2(t) (quadrupole)
I*1(t) (dipole)
quadrupole relative field error  210-3
  100 ms
B1(t)
B2(t)
Notes: (Re)establish clear responsibility boundaries with BE/CO: - local control racks - synthetic B-trains - closed-loop field control regulation (e.g. ISOLDE)
“piquet” service: so far a best-effort approach with three people on-call has been sufficient, what about the long-term ?
standardization of synthetic trains with FIDEL-like algorithms: magnetic measurements+ beam measurements + numerical simulations  mathematical model of magnet behaviour

## Slide 28
Mitigation of eddy current effects
Compensation of eddy currents by controlled current overshoot:
Ideally, a linear overshoot of length 2 can cancel out completely eddy currents on the flat-top
I(t)
I*(t)-I(t)
2
 = Le/Re=time constant of the eddy currents
Can this be done with Imain or IF8L ?
Notes: (Re)establish clear responsibility boundaries with BE/CO: - local control racks - synthetic B-trains - closed-loop field control regulation (e.g. ISOLDE)
“piquet” service: so far a best-effort approach with three people on-call has been sufficient, what about the long-term ?
standardization of synthetic trains with FIDEL-like algorithms: magnetic measurements+ beam measurements + numerical simulations  mathematical model of magnet behaviour

## Slide 29
The new B-train system
functional specifications

## Slide 30
Why a new B-train
Idea conceived in 2007 and green-lighted by IEFC in 2011.
Main goals evolved over time:
Ensure long-term reliability concerns about spares: peaking strips (PS), old electronics (SPS)
Improve performance and flexibility as requested by users higher resolution 0.05 G, higher dB/dt for future machines remove operation constraints and downtime: <50 G flat-bottom due to peaking strips, constraints on cycle sequencing, recurrent timing and synch errors …
Correct/extend measurements to counter beam instability (e.g. instrument fully F half to cure injection instabilities observed in since 2010)
Applies to PS in priority; progressively to be extended to SPS, AD and ELENA, PSB (upgraded) and LEIR
Notes: (Re)establish clear responsibility boundaries with BE/CO: - local control racks - synthetic B-trains - closed-loop field control regulation (e.g. ISOLDE)
“piquet” service: so far a best-effort approach with three people on-call has been sufficient, what about the long-term ?
standardization of synthetic trains with FIDEL-like algorithms: magnetic measurements+ beam measurements + numerical simulations  mathematical model of magnet behaviour

## Slide 31
B-train upgrade project: aim
All technical solutions discussed in detail with users: operation/power converters/RF/beam diagnostics + control
Progressive upgrade in parallel with existing system (for cross-check and as a fallback solution )
Phase | Objectives | Time frame
I – Component test | Test flux integration and correction algorithms Test individual electronic cards and components Compute offline B(t) and validate vs. beam behaviour | 2012
II – System test | Minimal viability test: broadcast simplified B(t) to POPS and/or RF with beam | Nov-Dec 2012
Deployment | Replace current system | LS1

## Slide 32
Functional specifications – B(t) integration
N different marker levels  up to 4 N corrections per cycle
Corrections smeared over a certain t to avoid sudden jumps
ADC offset ( integrator drift) updated every time the same marker level is reached
Integrator gain updated every time a different marked level is reached
New concept: continuous digital distribution of B(t) on a fast serial line to replace old pulse train
Output stream continuously updated with current B(t) values  no need for complex triggering and reset logic; no B0 initialization burst
Adaptive auto-calibration of gain and offset using field markers on flat-top and flat-bottom  fully transparent correction of measurement error
fast and robust transmission, no more timing conflicts but: drift correction must be applied continuously on-the-fly

## Slide 33
Functional specifications: integration F and D halves
1D
2D
3D
4D
5D
1F
2F
3F
4F
5F
flux coils
reference magnet U101, bldg. 355
A0(1-k)
A0(1+k)
VD
VF
FMR
FMR
Independent high-field markers and field integration in F/D halves
No complicated marker interpolation required: Bavg is constantly updated as t marker triggers transparent internal correction in F/D haves independently
Higher ~600 G flat-bottom level to improve magnetic reproducibility and allow the utilization of new FMR markers (NB: a stopgap solution using existing peaking strips was discussed and found to be technically as complex as switching over to FMR, minus all the advantages)

## Slide 34
Component tests (2012)
peaking strips
flux loops
peaking strips
flux loops
FMR
NMR / high field FMR
Prototype single-crystal FMR (due in March)
stability vs. peaking strips/beam
precision of trigger generation
test 600 G flat-bottom mode
current B-train sensors (stay in place)
Prototype harmonic coil array (in operation)
Bn(Ik) relationships
dynamic effects on multipoles
provide multipoles to operation
harmonic array
test statistically real-world reproducibility of the FMR vs. peaking strips and beam
test NMR as a high field marker (the effectiveness of different combinations of FMR/NMR probes for offset and gain correction has to be proven)
measure time evolution of field harmonics
FMR
harmonic array
Second harmonic coil array (symmetric)
Final FMR (single or poly-crystal)
symmetric placement
generate full integral
High field marker for on-the-fly gain correction

## Slide 35
Magnetic stability test: 500 G flat-bottom cycling mode
descent below 50 G no longer needed when FMR replaces peaking strip
10× flat bottom=500 G  better magnetic stability (should be tested shortly !)
flat bottom 500 G (200 A)
low field marker 600 G (240 A)
injection level p 1013 G (404 A) ions 687 G (275 A)
zero cycle flat-top 200 A

## Slide 36
Integral fluxmeter
Single-layer coils with <10 turns sufficient for good signal levels
candidate technology: MedAustron-style multi-PCB
near-ideal sensor: integral coil array for Bnd
removes one major uncertainty source i.e. extrapolation from local to integral measurement  future-proof against more surprises
no design available to date, but unlikely to be mechanically compatible with existing sensors can we dare to give it a try with beam before LS1 ?
central coil is larger to accommodate FMR markers

## Slide 37
The new B-train system
hardware & software

## Slide 38
Hardware platform
Standard CERN HW/SW platform, mostly based on the Open Hardware concept strongly endorsed and fully supported by BE-CO in view of integration in the control infrastructure:
PICMG1.3 – (Open Modular Computing Specification) bus standard (industrial PC)
SPEC (Simple PCIe FMC carrier) + FMC (FPGA mezzanine cards) for modular custom electronics, easily portable to other bus standards e.g. VME64x
Wishbone serializer bridge FPGA core (used e.g. for DMA access)
White Rabbit serial link
Scientific Linux with C/C++, optional real-time kernel

## Slide 39
White Rabbit B(t) distribution
New CERN fieldbus standard = Ethernet + ns-accurate timing
inexpensive, robust, extremely fast (Gbit/s) and accurate optical link
Ethernet standard guarantees maintainability and interoperability with future control systems
64 byte standard data frames allow to pack in extra data for free (e.g. dB/dt)
New electronic cards needed for compatibility with POPS (with support from BE/CO)
The impact of switches on latency and jitter needs to be assessed
Proposed data format:
32-bit B(t) stream  ppb-level resolution (e.g. 50 nT over a range of 20 T)
PS only needs 4 ppm (0.05 G), however the extra range is useful for: - upcoming applications e.g. ELENA (very low field) will use same electronics and software - reducing quantization noise in the computed derivative
1 MHz frame rate  max. 0.02 G differences between updates (even if the spectral content of the field is practically negligible above 10 kHz, high update rate comes at no cost and guarantees consistency for all users, which might be otherwise forced to use different interpolation methods)
No UTC referencing (1 s latency target, to be tested)
System based on industry standard BUT largely overspecified respect to actual needs leaves headroom to scale down in case 1 MHz frame rate cannot be met in practical tests
NB: users wishing to maintain old system (i.e. Beam Current DCCT) will retrofit their HW with our help but under their responsibility

## Slide 40
New hardware: PCIe FCM cards
Two cards designed and prototyped by MSC-MM (no commercial equivalent exists):
FMR peak detection card: based on 16-bit, 10 MHz ADC Function: produce a trigger corresponding to the FMR resonance peak via differentiation/threshold comparison. Compatible with NMR demodulated signal. Trigger validation (today done via error-prone synchronization with the timing system) could be accomplished downstream in software (e.g. by comparison with expected values B(I))
Dual integrator card: based on a 18-bit, 2 MHz ADC carries out the integration according to the model (page 3). Fully floating input for better CM rejection. Calibration coefficients are expected to be computed in software by the host controller (much less stringent timing needed, more flexibility to adapt software algorithms). General purpose ADC functions, card to be made available CERN-wide and beyond (OHCR community).
Overall system block diagram

## Slide 41
FMR field marker

## Slide 42
FMR resonator
Æ0.3 mm YIG sphere
semi-circular RF loops
Yittrium Iron Garnet (YIG) sphere coupling two orthogonal semi-circular RF loops
Widespread commercial component used as bandpass RF filter
Insensitive to field gradient due to small diameter
Units tested at CERN:
  old poly-crystalline YIG, installed in the PS since the ‘90s as a manually operated diagnostic tool low filter Q, but insensitive to temperature
  new single-crystal YIG units. Higher Q 500-1000, critical alignment and T dependence tested in 2010/2011: one standard commercial + one customized unit
  new poly-crystalline unit: awaited May
RF in
RF out

## Slide 43
Metrological characterization of FMR field marker
Calibration of a factory-optimized Noryl-case resonator in a reference dipole vs. NMR:
DC reproducibility: 0.04 G over the range 550 to 1200 G
Ramp reproducibility: 0.1 G @ 600 G, 2.4 T/s
Temperature dependence: 0.04 G/C (2C in U101  close to tolerance !)
Roll angle dependence: 0.04 G/mrad @ 600 G; non-linearity vs. B disappears at certain angles  very difficult to use in teslameter mode; strict mechanical tolerance in marker mode
Resonance spread due to B: still very good Q1000 at B/B=4.6 m-1 (as in U101)
ramp-rate dependence  reproducibility
non. linearity error vs. B, roll angle
nominal horizontal mounting

## Slide 44
NMR vs FMR
100 T
FMR
Sensor output during a field ramp in “marker mode” (fixed frequency excitation)
 10 T
NMR
Marker |  | 
NMR | absolute reference (metrological standard) commercially available instrument | Requires B compensation limited ramp rate: 20 to 50 mT/s B43 mT
FMR single-crystal | works up to several T/s simple direct acquisition of the resonance commercially available sensor larger dynamic range for a given sensor (current unit: 60-300 mT) | B60 mT broader resonance peak needs complex calibration temperature-dependent (must be optimized for a target field)

## Slide 45
Conclusions & outlook

## Slide 46
Summary
Diagnostic Spy System:
The system can be used to explain beam observations in terms of the magnetic response of U101. Indication obtained: the tune drift on the flat top can be mitigated by controlling current overshoot
Systematic correlations will continue to be investigated at the demand of Operation and Beam Physics teams. Other tests will proceed in parallel when possible (e.g. impact of eddy currents in the vacuum chambers).
Upgraded B-train:
Prototypes of sensors and electronic components and are being built and shall be tested in the next months. When the existing B-train appears to be inconsistent with the beam, we shall provide a simulation of the upgraded B(t) to validate the measurement.
Detailed schedule of tests and requests of MD for 2012 is in preparation. Online tests involving generation and distribution of B(t) with both proton and ion beams should be carried out to validate the design before the 2013 shutdown. An extended period in which the old and new B-train work in parallel is foreseen in 2014.
Notes: (Re)establish clear responsibility boundaries with BE/CO: - local control racks - synthetic B-trains - closed-loop field control regulation (e.g. ISOLDE)
“piquet” service: so far a best-effort approach with three people on-call has been sufficient, what about the long-term ?
standardization of synthetic trains with FIDEL-like algorithms: magnetic measurements+ beam measurements + numerical simulations  mathematical model of magnet behaviour