## Slide 1: The ALICE TPC upgrade
C. Garabatos, GSI
for the TPC group

## Slide 2: The largest TPC ever
5 m x 5 m
High mechanical precision: 200 mm
100 kV in central electrode
Cooling everywhere
Wire chambers with pad readout
Ne-CO2(-N2)
Designed for dN/dy = 8000 (now 1600) and 200 Hz readout rate Pb-Pb

## Slide 4
Challenging requirements for the gas
High
Multiplicities
Momentum
Resolution
High Rate
(200 Hz)
dE/dx
Resolution
Occupancy
2 track res.
Field
Distortions
Multiple
Scattering
Drift speed
(max. HV)
Signal/
Noise
Argon
Neon
Hydrocarbons
CO2
CF4
Ageing
Flammability
Neutrons
Concentration:
90-10
Gain:
2 ´ 104
?
???
?
Small pads:
4 ´ 7.5 mm2

## Slide 5: Gas dependencies: calibration challenge
Drift velocity | 90-10 | 90-10-5
Temperature | +0.37 % / K | +0.34 % / K
Pressure | -0.15 % / mbar | -0.15 % / mbar
CO2 concentration | -7.6 % / %CO2 | -6.4 % / %CO2
N2 contamination | -1 % / %N2 | -1 % / %N2
Gain | 90-10 | 90-10-5
Temperature | +0.9 % / K | ? % / K
Pressure | -0.34 % / mbar | ? % / mbar
CO2 concentration | +67, -20 % / %CO2 | +17, -14 % / %CO2
N2 contamination | +34 % / %N2 | +6.3 % / %N2

## Slide 6: Baseline scenario: benefit from the excellent rate capability of GEMs
Continuous –untriggered- readout of Pb-Pb collisions at 50 kHz  replace wires by triple GEMs
  to limit space-charge in drift volume
  cope with rate
Spatial resolution somewhat limited by space-charge, but ITS+TRD provide enough momentum resolution (3% at 80 GeV/c)
Maintain excellent PID
Perhaps a fast gas
New FE electronics
  x2 channles
Existing FC at ≤ 100 kV

## Slide 7: About GEMs
50 mm metalized polyimide foils with tiny holes at 150 mm spacing where amplification takes place
Signal entirely produced by electrons
Ions absorbed at the electrodes, some 0.1 % leak back into drift
Rate >106 counts/mm2
Triple assembly
Segmentation to avoid
damaging discharges
Produced at CERN and others
Widely used now

## Slide 8: Gain dependence of GEMs
Gain increases at switch-on (30%)
Effect depends on foil properties and rate (mins-hour)
Slow decrease with time (hours/days)
Slow decrease at switch-off (days)
Gain depends on rate
Calibration issue
F. Sauli et al, NIM A409 (2002)
B. Azmoun et al,
2006 IEEE NSS
F. Simon et al,
IEEE TNS 54-6 (2007)

## Slide 9: Gain dependence issue
Charge-up times at Stable Beams in ALICE probably or order of many minutes
Careful monitoring and calibration (offline, in baseline scenario) is needed
R&D on minimization of polyimide charge-up and polarization/charge migration effects is needed
  field configuration
  water? (Brrr…, not good with CF4)

## Slide 10: Fast gas: Ne-CF4?
Up to factor 3 faster than CO2 at 400 V/cm
Provides flexibility to choose drift field
Unfortunately, the transverse diffusion is smaller, and the longitudinal is larger, than in CO2 (220 mm/√cm)  wrong direction

## Slide 11: Fast gas: Ne-CF4?
In addition, diffusion in CF4 shrinks with magnetic field  more channels
Mobility of ions in Ne-CF4 not known, but taking CF4+ on CF4, distortions would be 3-4 cm max., comparable with Ne-CO2 (50 kHz Pb-Pb with 4 ions/electron leaking out)  not critical
A fast gas only buys us less overlapping events (5  2), but TPC designed for dN/dy = 8000
mm
Ar-CF4-iC4H8
B = 0 T
B = 1 T
Kobayashi et al,
arxiv.org/pdf/1008.5068.pdf

## Slide 12: Fast gas: Ne-CF4?
CF4 exhibits electron attachment near and in the amplification region, leading to further gain fluctuations
Kobayashi et al,
arxiv.org/pdf/1008.5068.pdf
In addition, the compatibility of CF4 with all materials must be validated, and H2O must be avoided

## Slide 13: Estimate of number of readout channels
At > 1 m drift distance, diffusion of order of ~2-4 mm
Probably can live with 4 mm wide pads for IROCs
Reduce the OROC pads (6 mm) to 2-3 mm
This results in roughly a factor 2 more pads: 1-1.5 M
Anyway, position resolution can be somewhat relaxed. PID is important

## Slide 14
CMS High Eta Upgrade Electronics System

## Slide 15
The SALTRO Demonstrator Chip
Charge amplifiers, ADC, digital processing all in one chip

## Slide 16
Further evolution for CMS/LCTPC
(could support continuos readout)

## Slide 17: Short-term plans
Prepare a small set-up for laboratory measurements (with RD51)
  ion mobility
  ion back-flow
  gain stability
  gas mixtures
  cluster size
  field configuration
  ageing: 0.1 C/cm2/month Pb-Pb (100 times more!)
Prepare an IROC with GEMs to be exposed to p-Pb in the cavern
  main goal is to prove stability of GEMs under LHC conditions (unfortunately no Pb-Pb), including gain behavior
  exact running conditions to be worked out now

## Slide 18: Fit a triple GEM into a spare IROC
GEM foils being designed by the Munich group
No problem to produce single-mask full surface GEMs
  perhaps thin frames in the active area are needed to prevent sag
Enough feed-throughs in chamber support to power 6 electrodes
Segmentation (≤100 cm2) results into 19 segments
FC box needs a ‘last resistor’ to match the drift field to the potential of 1st GEM
1-2 chambers will be prepared

## Slide 19: Schedule
2012: prepare, and carry out, test in the ALICE cavern and basic R&D
  decision to go ahead
  define the project: chambers and readout
2012: build Collaboration
2012-2017: design, produce, test
2018: installation

## Slide 20: Conclusions
Use existing Field Cage with GEM readout and new electronics for continuous readout in Pb-Pb at 50 kHz
Start tests now; put a prototype in the experiment
Staying with the current, slow gas would save short-term R&D on
  ion mobility
  ‘PRF’
  compatibility with materials and water
  electron attachment
Start search for collaborators now

## Slide 21: BACKUP SLIDES

## Slide 22: GEM with Ne-CO2 – Space-charge distortions
50 kHz Pb-Pb, 4 ions leakage per primary electron
S. Rossegger

## Slide 23: GEM with Ne-CF4 – Space-charge distortions
50 kHz Pb-Pb, 4 ions leakage per primary electron
S. Rossegger

## Slide 24: Ion mobility in Ar and Ne
No information found on ion mobility of Ne-CF4 mixtures. It must be measured