## Slide 1: The design and preliminary tests of microstrip-microgap RPC (M-M-RPC)
P. Fonte1, R. Oliveira2,G. Paic2,3, V. Peskov2,3
F. Pietropaolo4, P. Pichhi5
1LIP, Coimbra, Portugal
2CERN, Geneva, Switzerland
3UNAM, Mexico
4 INFN Padova, Padova, Italy
5INFN Frascati, Frascati, Italy

## Slide 2: There are some experimental conditions which require detection of particles or gammas with high position and time resolutions
Typical example: upgrade ATLAS wheel, TOF PET, X-ray scanners
Today we will report about the first step in this direction:
development high position and time resolutions detectors

## Slide 3: 1. Detector design

## Slide 4
a)
Multilayer PCB with a Cu
layer on the top and
one layer of readout strips
on the bottom, 0.5 pitch
Upper Cu layer etching
The grooves were then
filled with resistive paste
(ELECTRA Polymers
Removal of the Cu
v
If necessary, filling with
Coverlay (an option)
b)
c)
d)
e)
M-M- RPC manufacturing steps:
Resistive strips
Readout strips
0.5 mm
0.2mm
v
0.035mm
0.1mm

## Slide 5
A
B
C
Contact
pad
Contact
pad
Resistive strips
Total resistivity of
the zone B 500MΩ
(adjustable)
Resistivity of
zones A and C
500MΩ (adjustable)
Surface resistivity
100kΩ/□ (can be
adjusted to exper.
needs)
Top view:

## Slide 6: This plate is in fact a reproduction of the resistive MICROMEGAS anode board
The idea is to assemble from these plates a parallel- plate detector (M-M-RPC), so that mesh is not used

## Slide 7
Orthogonal
resistive strips
Current
Inner signal strips
Artistic view of the M-M RPC
PCB sheet
From these plates RPC were assembled with gaps ether 0.5 or 0.18mm

## Slide 8
Magnified photograph of the inner surface of the M-M- RPC

## Slide 9: Resistive strips
Readout strips
located below the resistive strips

## Slide 10: An option with pillars

## Slide 11: A fundamental different between “classical “ RPC and M-M- RPC
Film resistor
M-M-RPC offers high2D position resolutions (with orthogonal strip or various stereo
strip arrangements to avoid ambiguity) and good timing properties
Usual
RPC
M-M-RPC
“Signal”
electrodes
Current
Orthogonal
resistive strips
Current
500MΩ
Inner signal strips

## Slide 12: 2. Experimental setup

## Slide 13
Va
Inner strips
Gas chamber
Window
Charge-sensitive amplifiers (a possibility)
UV lamp or X-ray gun
Vc
Collimator
M-M-RPC
Signal pickup
strips
Sr source
Betas

## Slide 14: 3. First results
(Preliminary measuremenst of basic characteristics: gas gain, induced charge profile in Ar+ethane and Ar+CO2 in order to compare with results obtained with mesh RPC and with resistive MICROMEGAS)

## Slide 15: 3.1. Gain measurements

## Slide 16: Gain estimation in detectors with a cathode mesh:
X-rays
P. Fonte et al., arXiv:physics/9803021, 1998
Drift mesh
Cathode mesh
Anode

## Slide 17: Gain estimation in an RPC geometry:
CsI layer
UV
X-rays
Fe anode
0.5mm
0.1-0.2mm
Photoelectron tracks
Due to the time constrains the CsI coating was done by a spray technique

## Slide 18
P. Fonte et al., NIM A431,1999, 154

## Slide 19: R-R-RPC with spacers in corners
Preliminary estimations: no good plateau, charging up(?), space charge( ?)…

## Slide 20: The highest gains were obtained among all resistive micropattern detectors
(to be discussed)

## Slide 21: 3.2.Charge profile measurements
The current prototype have not connectors allowing to readout all strips independently

## Slide 22
Induced signal profile measured with 0.5mm gap M-M-RPC in Ar+25%CO2
Preliminary (step-strip scan)!
Anode strip
Cathode strip
I. Crotty et al., NIM A505, 2003, 203

## Slide 23: Preliminary: R-R-RPC with pillars
More studies should be done
(Changes in CsI? Problem with pillars?)
Region of current instability

## Slide 24: Conclusions:
•Preliminary it looks like M-M-RPC can be an interesting alternative to R-MICROMEGAS
• Certainly more work should be done to prove this.
• Potential advantages: good position and time (much below ns) resolutions, possibility to apply automatic procedure to build large area M-M-RPC (no mesh)
Nearest plans:
More test with pillars and tests with Miranda CsI
Test of new designs (currently in manufacturing process)
Large- area M-M-RPC
Tests in “standard” RPC gases ensuring high time resolution
( in these studies Ar based mixtures were chosen
to compare to PPAC and resistive MICROMEGAS )
Study CsI coating as a sec. electr. emitter
Work more closely to J. Wotschack and P. Fonte group on applications for muon detection and TOF PET