## Slide 1: Test pulse calibration in Laboratory
Kenji YOSHIDA

## Slide 2: Calibration for electronics with test charge input
Y. Ueyama, S. Ozawa

## Slide 3: Test charge calibration
Test charge input
Calibration of ..
  Linearity
  Pedestal for PH-ADC(TASC 1st layer & CHDjp & IMC Dynode)
Charge injector
Pulsar
9.283pF
51Ω
Attenuator
Charge

## Slide 4: TASC-TOP readout SPS2011
Layer 1
PMT
P.A.
PWO
Slow S.A. H.G.
Slow S.A. L.G.
PH-ADC
PH-ADC
PMT
P.A.
PWO
Slow S.A. H.G.
Slow S.A. L.G.
PH-ADC
PH-ADC
PMT
P.A.
PWO
Slow S.A. H.G.
Slow S.A. L.G.
PH-ADC
PH-ADC
Layer 1
Layer 1
Slow S.A. H.G.
Slow S.A. L.G.
PH-ADC
PH-ADC
APD/PD
PWO
Layer 2-12
TASC-FEC
F.S. + Discri.
Input Reg.
P.A.
CSA(calet)
P.A.
OP-AMP(calet)
S.A.
Clear pulse
S.A.
calet
CAMAC
VME
SUM-AMP
IMC Anode readout SPS2011
SciFi
PMT
SciFi
PMT
SciFi
PMT
SciFi
PMT
FEC For bCALET-2
S/N 073195

## Slide 5: IMC Dynode and CHD readout for SPS2011
PMT
P.A.
CHD(JPN)
Slow S.A.
PH-ADC
F.S. + Discri.
Input Reg.
PMT
P.A.
CHD(JPN)
Slow S.A.
PH-ADC
F.S. + Discri.
Input Reg.
SciFi
PMT
P.A.
dynode
Slow S.A.
PH-ADC
F.S. + Discri.
Input Reg.
SciFi
PMT
P.A.
Slow S.A.
PH-ADC
F.S. + Discri.
Input Reg.
SciFi
PMT
P.A.
Slow S.A.
PH-ADC
F.S. + Discri.
Input Reg.
SciFi
PMT
P.A.
Slow S.A.
PH-ADC
F.S. + Discri.
Input Reg.
IMC-FEC
anode
64
P.A.
P.A.
CSA(as calet)
OP-AMP(as calet)
S.A.
Clear pulse
CAMAC
VME
64
64
64

## Slide 6
Qin: 2.1fC～1MIP
Qin = V ×10pF
Linearity ~ 3 ×103 MIPs
(1MIP 2.1 [fC]@APD G=50)
Linearity ~ 60MIPs
(1MIP 200 [fC]@PMT500[V])
Linearity ~ 1200 MIPs
(1MIP 100 [fC]@PMT450[V])
In this graph,
MIPs was defined with
APD 1mip charge.
Feedback capacitor was 6 pF.
CHD
TTOP
Cf = 5 pF＋1pF(internal)
Cf = 120 pF
Cf = 120 pF
Cf = 6 pF
Checking the Linearity of CSA

## Slide 7
Linearity ~ 33 MIPs
(1MIP 160 [fC]@PMT 800 [V])
Linearity ~ 85 MIPs
(1MIP 64 [fC]@PMT 700 [V])
Linearity ~ 245 MIPs
(1MIP 22 [fC]@PMT 600 [V])
1 MIP(dynode)
Qin = V ×10pF
Dy 1,2
Dy 3
Dy 4
Linearity of Pre AMP for IMC Dynode

## Slide 8: The calibration of PH-ADC

## Slide 9: PH-ADC pedestal and linearity calibration
CHD 1
Dynode 1
Dynode 4
TASC-TOP SUM HighGain

## Slide 10: PH-ADC slope and pedestal calibration (preliminary)
PHADC channel | target | (slope) | pedestal [ch]
1 | CHD #1 | 1.68035 | 154.094
2 | CHD #2 | 0.574692 | 197.157
3 | Dynode #1 | 3.54453 | 146.255
4 | Dynode #2 | 3.44006 | 40.6197
5 | Dynode #3 | 4.39587 | 138.69
6 | Dynode #4 | 25.4713 | 133.277
7 | TASC TOP SUM High | 39.8917 | 67.2193
8 | TASC TOP SUM Low | 1.0813 | 37.4566
9 | TASC TOP #1 High | 3.93418 | 155.461
10 | TASC TOP #1 Low | 0.124272 | 209.485
11 | TASC TOP #2 High | 3.84957 | 121.782
12 | TASC TOP #2 Low | 0.117459 | 191.29
13 | TASC TOP #3 High | 3.73582 | 159.043
14 | TASC TOP #3 Low | 0.120923 | 24.533
15 | --- no use --- | --- | ---
16 | --- no use --- | --- | ---
ADC [ch] = input charge [fC] × (slope) + pedestal

## Slide 11: Issues on the PH-ADC calibration
The PH-ADC slopes and pedestals in the table is used in the present data analysis.
As shown in the PH-ADC calibration figures, some PH-ADC calibration data points are not enough.
The present PH-ADC calibration table is a preliminary version, and later we will upgrade and distribute the table with enough calibration data points.

## Slide 12: The calibration of IMC FEC

## Slide 13: IMC FEC calibration
Ex. CH461
Cross talk (input -24.956pC)
Input charge [pC]
ADC [count]
(ADC) = - 1379.1 (Input [pC]) + 33011.9
→ 8.78 [fC/count]

## Slide 14: Backup

## Slide 15: Readout equipment for SPS2011
For IMC MaPMT Anode
  CP 80067 S/N 073195 (for bCALET-2)
    Va32HDR14.3
    PCB is same as bCALET-2
For TASC-TOP PMT
  Charge sensitive amp
    Similar as CALET read out
  Shaping AMP
    Similar as CALET (Made by D. Ito)
    Dual range
  CP 1116 Peak hold ADC module
    16bit
    Gate width = 6μs
For TASC-TOP Sum and CHD and IMC MaPMT Dynode
  Almost same as TASC-TOP PMT
  Pre AMP of IMC Dynode are OP-AMP(same as CALET)
  Shaping amp is CP 8679/M(Unipolar, gain adjustable )

## Slide 16: The gain of amplifiers
 | 1MIP ADC [ch] | 1MIP Charge out[fC] | Pre amp | Shaping Amp gain
TASC-TOP High-gain | 400 | 100@-480 | 120pF | (600)
TASC-TOP Low Gain | - | - |  | (20)
CHD | 1000 | 200@-600V | 120pF | 
IMC(Dynode 1st,2nd) | 1000 | 200@-800V |  | 
IMC(Dynode 3rd) | 1000 | 200@-700V |  | 
IMC(Dynode 4th) | 1000 | 200@-600V |  | 
TASC TOP SUM | 1000 | 200@-450V | 120pF | 
Pre AMP gain = Q/Cf

## Slide 17: PH-ADC pedestal and linearity calibration backup
CHD 1
CHD 2

## Slide 18: PH-ADC pedestal and linearity calibration backup
Dynode 3
Dynode 4
Dynode 1
Dynode 2

## Slide 19: PH-ADC pedestal and linearity calibration backup
TASC TOP #1 HighGain
TASC TOP SUM HighGain
TASC TOP SUM LowGain
TASC TOP #1 LowGain

## Slide 20: PH-ADC pedestal and linearity calibration backup
TASC TOP #3 HighGain
TASC TOP #2 HighGain
TASC TOP #2 LowGain
TASC TOP #3 LowGain

## Slide 21: Linearity of Pre AMP for IMC Dynode

## Slide 22: Checking the Linearity of CSA
Required
（0.5～3×103 MIPs）
Dynamic Range
Qin: 2.1fC = 1MIP
400ch
～2.1fC
Qin = V ×10pF
9.283pF
51Ω